# A Compact Low-Noise Broadband Digital Picoammeter Architecture

M. Crescentini1, M. Tartagni1,2, H. Morgan3, and P.A. Traverso4

1Department of Electrical, Electronic and Information Engineering (DEI) “Guglielmo Marconi”, Cesena Campus, University of Bologna, Cesena IT-47521, IT.

2Advanced Research Center on Electronic Systems (ARCES), Cesena Campus, University of Bologna, Cesena IT-47521, IT.

3Electronics and Computer Science & Institute for Life Sciences, University of Southampton, Southampton SO17 1BJ, U.K.

4Department of Electrical, Electronic and Information Engineering (DEI) “Guglielmo Marconi”, Bologna Campus, University of Bologna, Bologna IT-40136, IT.

Corresponding author is M. Crescentini: tel: 0039 0547-339236; fax: 0039 0547-339208; e-mail: marco.crescentini3@ unibo.it; address: Via Venezia 52, Cesena 47521 (FC), Italy.

Abstract

A low-noise (≤ 4 fA/√Hz), broadband (≥ 100 kHz) compact architecture and related operation solutions are proposed for portable and low-cost time-domain acquisition of currents with effective resolution in the order of 1 pA and below. The front-end architecture is based on an integrating-differentiating scheme to achieve the optimal performance in terms of input-referred equivalent noise, but it overcomes the typical noise/bandwidth trade-off by making the sampling frequency of the A/D conversion independent from the rate at which the analog front-end is reset. In order to strongly mitigate the main drawback, i.e., the introduction in the system of an inherent time-variance, a Track-and-Hold circuit synchronized with the reset is exploited.

For validation purposes, a dual-channel prototype was implemented in a low-cost CMOS technology. The prototype is characterized by standard figures of merit and is experimentally validated by two simple case studies, which are typical of practical applications.

Index Terms

Ammeter, low-current measurement, current, leakage currents, charge-sensitive amplifier, low-noise amplifier, femto-ampere.

# Introduction

Low-current measurement is a critical task in applications dealing with either physical phenomenon which intrinsically generates a flow of a few carriers or high-valued impedances that cannot be biased at high voltages. Typical examples are the measurement of dark current in novel high-performance photodetectors [1], the measurement of trans-membrane current in living cells [2,3], electrochemical impedance spectroscopy [4], nanotechnology and nanoelectronics investigation [5,6], noise characterization of electron devices [7] and many others. Since there exists a wide spread of applications, each with its own specific requirements, the state-of-the-art low-current ammeter architecture should be very flexible, offering both DC and AC measurements, different input ranges, resolution as low as a few fA, and adjustable acquisition bandwidth. Reduced size and cost of the system are additional features that allow to exploit it in portable instruments or as part of a more complex set-up, such as probe-stations. The common requirement of most of the above applications is the capability to accurately measure very low currents (below 10 nA) with associated output impedances in the order of megaohm and above.

There is not a commercial instrument with all these features in the market. For instance, Keysight B2980 [8] offers a minimum resolution of 1 aA with an acquisition bandwidth lower than 1 Hz, while Axon Axopatch 200B [9] acquires currents as low as 100 fA with 10-kHz bandwidth; they both are bulky and expensive instruments.

This paper proposes a compact, flexible, low-noise (< 10 fA/√Hz) and wide-bandwidth (up to 1 MHz and above) digital picoammeter architecture. The analog front-end, which is the key element of the architecture, is based on an integrating-differentiating (int-diff) solution in order to achieve the low-noise performance. In this block, the critical issue of avoiding the saturation of the integrator stage is performed by periodical reset at frequency fR, while the following Analog-to-Digital Converter (ADC) samples at frequency fS much higher than fR. Such a relationship between sampling frequency and reset frequency is essential to achieve the MHz-bandwidth capability while preserving the noise performance, although its implementation modifies the statistical properties of the noise that becomes a cyclostationary stochastic process, and introduces non-linear distortion in the output signal since the overall architecture results to be, strictly speaking, time-variant. Non-linear distortion due to time-variance is strongly mitigated by a Track-and-Hold (T&H) circuit placed before the ADC. The fast sampling of the analog signal at the output of the front-end is guaranteed by a delta-sigma (ΔΣ) ADC. The data stream at ADC output requires digital processing, i.e., low-pass filtering and decimation, so as to recover the multi-bit data at the equivalent Shannon sampling frequency.

The compactness features and performance capabilities of the proposed architecture have been experimentally assessed by means of the design and implementation of a prototype. This proof-of-concept integrates on a single chip both the overall analog circuits and the A/D conversion stage (fS = 10 MSa/s) in a standard low-cost 0.35-µm CMOS technology, while digital data processing is carried out externally on a low-cost FPGA.

This paper is organized as follows. The general architecture is described in section II, with particular insight on the description of the analog front-end and the solutions proposed to simultaneously achieve high sensitivity and wide bandwidth. The experimental characterization and validation of the fabricated prototype are reported and discussed in section III, in which the proposed architecture demonstrates its potentialities by also considering simple case studies that implement the operating conditions typical of real applications, i.e., the measurement of broadband thermal noise generated by a 1-MΩ precision metal film resistor, and the characterization of flicker noise in a commercial low-dropout (LDO) regulator. Conclusions are drawn in section IV.

# Architecture

The architecture can be subdivided into three main blocks (Fig. 1): i) analog front-end, ii) ΔΣ ADC and iii) digital data elaboration. The proposed analog front-end topology and related operation are the key elements to allow for both very low-noise and wide-bandwidth current measurement. This block is usually the bottleneck of the entire system from the bandwidth standpoint, when considering conventional architectures. This section will discuss the hardware and operational solutions adopted to overcome the bandwidth limitations while preserving the noise performance.

Fig. 1 Block diagram describing the proposed architecture in terms of its main sub-systems.

## ANALOG FRONT-END

The analog front-end acquires the input current i(t) and converts it into a voltage vO(t). It is based on the integrating-differentiating approach, which proved to be the best scheme for very low-noise current interfaces [10]. In fact, by considering the conventional resistive-feedback transimpedance amplifier (TIA) of Fig. 2-B, the equivalent, input-referred, noise power density (PSD) , i.e. the unilateral spectral density of the random process  that perturbs , is given by

(1)

where k is the Boltzmann constant, T is the temperature in Kelvin degrees, Rf is the feedback resistor, Gin and Gen are the input-referred noise PSDs of the operational amplifier (OPA), Cf is the parasitic capacitance in parallel to the feedback resistance, Cs is the source capacitance of the DUT (see Fig. 2-A)  and Cp is the parasitic input capacitance of the OPA. Instead, by considering the int-diff scheme of Fig. 2-C, the equivalent input-referred noise PSD can theoretically be as low as

(2)

with Cin = Cs + Cp,. In Eq. (2) the noise due to the differentiator is neglected since a high gain of the integrator stage is assumed. The better noise performance of the int-diff scheme is apparent from the comparison of (1) and (2), since an important source of broadband noise (i.e., Rf ) has been eliminated in the solution of Fig. 2-C.

Fig. 2 A) Equivalent circuit showing the DUT current source, the input-referred equivalent noise current δi that perturbs the measurand i(t) and the picoammeter assumed noiseless. Equivalent circuit of a resistive feedback TIA (B) and an int-diff scheme (C), highlighting the main noise sources.

However, from a practical standpoint, the critical limitation of the int-diff approach is the saturation of the integrator stage, mainly in response to any small DC current contribution. Two possible solutions offered by the literature are:

- Add a negative active feedback that cancels out the DC current at the input [11].

- Periodically reset the output node [2,12].

In both cases, an additional contribution Gadd is introduced in the input-referred noise PSD:

(3)

is either the noise due to the electron devices within the active feedback (case 1) or the folding noise due to reset and sampling (case 2) [10]. The detailed comparison of the two solutions can be found in [10], where all their pros and cons are listed. In brief, cancellation of the DC current through active feedback provides wide acquisition bandwidth but it involves worse noise performance and stability issues as well [11]; while periodical reset is straightforward to be implemented and shows better noise performance at low frequencies but it leads to a limited acquisition bandwidth, since sampling frequency fS usually equals the reset frequency fR, so as to optimize the system in terms of area and power consumption. In fact, when fR = fS, the memory of the integrator is canceled every cycle; hence the output of the integrator can be treated as an estimation of the mean value of the input, and a practical implementation of the differentiator stage is not required any longer, resulting in less area and power consumption [12].

By considering the periodical reset of the integrator (case 2) with the basic, conventional constraint , the additional noise  can be approximated as

(4)

where GBW is the gain-bandwidth product of the OPA, while

(5)

is the equivalent transresistance of the ammeter [10]. Equation (4) has been obtained by computing the noise voltage PSD at the output node and referring it back to the input dividing by Req squared [10].

The input noise PSD given by (4) is explicitly proportional to the sampling frequency fS, thus a noise-performance/bandwidth trade-off is manifest: higher sampling frequency allows for wider acquisition bandwidth, but at the critical detriment of input-referred noise performance and current sensitivity. In other words, to preserve the better noise performance potentialities of solution in Fig. 2-C with respect to the conventional resistive feedback, bandwidth must be strongly limited in the practice.  It is worth noticing that (4) is not in contrast with the theory of folding noise in sampling processes [10,13]. In fact, from a general standpoint, the amount of folded noise at the output node lowers with fS, as expected, but the input-referred noise PSD in the architecture of Fig. 2-C actually rises when  due to the decrease of Req, which acts squared in the expression (4) of .

To overcome such a restriction, an integrator-differentiator scheme is proposed for the analog front-end of the picoammeter (Fig. 3), which is operated with sampling frequency fS much higher than reset frequency fR. More precisely, the output voltage  of the analog front-end is sampled at frequency fS regardless of the rate of saturation of the integrator, while the reset of the latter is performed at frequency fR, which is k times lower than fS. In this way the reset of the analog front-end becomes a sporadic, independent event with respect to the sampling process. To avoid sampling of corrupted signal during reset phase, a T&H placed between differentiator and sampler works at frequency fR, as shown in Fig. 3.

Fig. 3 Block diagram of the proposed architecture for the analog front-end. The input current i(t) is integrated, amplified and then differentiated to obtain a voltage x(t) proportional to it. All these three blocks periodically reset with frequency fR, while the T&H stage that follows holds the output signal vO(t) to the last value before reset. The sampling is carried out at frequency fS >> fR, independently of the reset process.

Thanks to the separation of fS from fR (i.e., the implementation of a sampling strategy independent from the reset phase), the equivalent transresistance Req is no longer related to the sampling frequency but it is now given by

(6)

hence, the input-referred folding noise  nominally (i.e., by neglecting for the moment the perturbation effects due to the reset phase) becomes inversely proportional to the sampling frequency (see the second member of Eq. (4)), as in a standard sampling process [13]. Thus, with the proposed solution, the acquisition bandwidth can be increased without degrading the noise performance.

Unfortunately, the reset phase, although short and carried out at , clearly introduces a perturbation in the operation of the system. To avoid sampling of the corrupted signal  during the reset phase, a T&H working at frequency fR is placed between differentiator and sampler. This solution strongly reduces the distortion on the useful signal, but it does not eliminate the shaping effect on the noise in the time domain; more precisely, a periodicity is created in the main statistical moments (i.e., expected value and autocorrelation function) of the noise, which becomes a cyclostationary stochastic process at T&H output [14,15]. This mechanism creates an additional folding of the noise in the frequency domain. By assuming duration of the reset phase τR<<TR, then the input-referred noise PSD for the proposed architecture can be written as:

(7)

where  is the theoretical minimum PSD of Eq. (2), the folding noise due to sampling is neglected (since fS can be arbitrarily increased), and the second term of the last member characterizes the new noise folding process due to reset [16]. The formulation of  expressed in (7) is no longer proportional to the sampling frequency, hence the acquisition bandwidth of this architecture can be increased without any loss of current sensitivity. Maximum acquisition bandwidth is now constrained only by implementation issues, such as the value of GBW of the OPAs and the design of the ADC.

The summation in (7) cannot be easily solved as done in [10] in order to get a more compact and explicit expression for  , since the power density under the summation is not white (see Eq. (2)). However, it is possible to notice that the additional noise  becomes comparable to the theoretical minimum PSD given by (2) due to the presence of the factor . This consideration is even more valid at high frequencies, because  rises with f2. The minimum values for fR and τR are set by passive and active devices used in the implementation.

Rigorously speaking, the differentiator stage, which has been considered not relevant during the above noise analysis, actually amplifies its own noise at high frequencies, potentially worsening the performance expressed by (7); hence a LPF can be added after the T&H block, as shown in Fig. 3.

Another important trade-off in the design of integrator-based schemes involves the value of the feedback capacitance Cf. A small Cf increases the sensitivity of the integrator to the input current, since A1 = 1/Cf, and lowers the minimum input-referred noise , but introduces stability issues at low frequencies. To deal with this drawback, we propose to choose a feedback capacitance comparable to the DUT capacitance Cs of the target applications, preserving in this way the stability of the design, and to balance the gain loss with an additional voltage amplifier placed between integrator and differentiator (Fig. 3). This solution introduces a further noise source but helps to keep the additional noise  low, since this gain stage contributes with the factor A2 in Eq. (7). Moreover, the ratio  directly affects the bandwidth BWINT of the integrator, which follows the relation

(8)

Thus, it is important to implement a variable Cf to match with the source capacitance and preserving the acquisition bandwidth.

## EFFECTS OF TIME-VARIANCE ON USEFUL SIGNAL

The implementation of a sampling strategy that is independent from the reset phase (i.e., the choice fS >> fR ) is the key point that allows the proposed architecture to overcome the trade-off noise-performance/bandwidth. However, this solution introduces time-variance in the operation of the system. From a practical standpoint, the effect of time-variance on the deterministic operation of the system is intermodulation distortion. More precisely, the time-variance generated by a periodical reset can be modeled as the product of the useful signal x(t) (Fig. 3) by a rectangular wave p(t) of period TR, duty cycle  and unitary amplitude. The time-variant signal vO(t) can be therefore written as

(9)

where pn are the coefficients of the Fourier series expansion of p(t). Eq. (9) clearly shows that time-variance creates side-bands in the spectrum of  around frequencies that are multiples of : in the significant case of a periodic signal  with period , the n-th side-band is characterized by spectral tones at with . The amount of distortion can be quantified by considering the Fourier coefficients pn:

(10)

Similarly to what is usually done in the case of conventional non-linear distortion, we can define a signal-to-distortion figure-of-merit SIDTV as the ratio between the power of the signal and the sum of the powers of the side-bands created by time-variance:

(11)

It is worth noticing that the SIDTV coefficient does not depend on the input signal but solely on the time-variance properties, specifically on period and duration of the reset.

Fig. 4 Timing of signals i(t), x(t) and vO(t) in Fig. 3. The periodical reset performed at fR<<fS introduces a strong distortion in the signal x(t). The T&H operation mitigates the distortion on signal vO(t).

The above discussion should be considered as a worst-case, since it clearly does not take into account the beneficial effect of the T&H placed before the A/D section in the architecture of Fig. 3. In fact, the T&H mitigates the distortion but a residual of the latter still affects the acquisition of the useful signal, as clearly shown in Fig. 4, in which the time evolution of signals x(t) (output of the differentiator) and vO(t) (the signal actually sampled) are reported. The two signals are nominally the same, apart from the short reset phase. With the use of the T&H at the start of reset,  does not return to zero as in Eq. (9), but it is hold at the last valid value of x. The weighting function  in (9) should be thus redefined (now depending on the evolution of x) and the analytical computation in (11) modified, with the resulting SIDTV increased. However, the effectiveness of the T&H depends on the properties of the useful signal, specifically on the ratio Δi/τR computed throughout the reset phase, which approximates the derivative of the input current for a very short reset. In case of DC or slow current inputs, the T&H fully counteracts the periodical reset and the output signal is a good reconstruction of the input signal; whereas for input signals with fast vertical dynamics, i.e., either high amplitude or high frequency or both, the T&H progressively loses effectiveness in the capability to reject the intermodulation distortion, and the signal-to-distortion coefficient reduces to values close to the worst-case given by Eq. (11).

## A/D SECTION AND DIGITAL DATA ELABORATION

Given the wide family of possible applications for the digital picoammeter architecture proposed, the main requirement on the A/D conversion section is flexibility. The ΔΣ modulator is a fair architecture for high-resolution, medium-bandwidth integrated A/D conversion: it is rather small, simple to be implemented and it can easily trade-off resolution for bandwidth by changing the oversampling factor (OSR) used in digital decimators [17]. Moreover, the proposed analog front-end requires a sampling frequency fS >> fR, which is inherently offered by a ΔΣ modulator that oversamples the input signal. For these reasons, we chose the ΔΣ modulator for the proposed architecture. The specific architecture of the ADC is not a main concern as far as it is designed to have total noise [22] below the thermal noise of the analog front-end. In this way, the final effective resolution of the picoammeter is determined by the limits of the analog interface. The inherently serialized nature of ΔΣ output is another advantage when concurrently acquisition of multiple signals is required, since it strongly simplifies the design of interconnections [2,18].

Digital data elaboration mainly consists of a FIR LPF and a digital decimator to get equivalent output data at M-bit with the proper rate (fM = fS/OSR). A common architecture for digital filtering and decimation of ΔΣ modulators is the sinc3 filter, since it does not require digital multipliers and relies only on sum and differences, simplifying the hardware implementation. The sinc3 filter architecture consists of three summing stages working at frequency fS and three differentiator stages working at frequency fM, as sketched in Fig. 5 [17].

Fig. 5 Block diagram of a sinc3 digital filter performing low-pass filtering and decimation by a factor OSR = fS/fM.

# Characterization and Validation

A prototype has been designed and implemented to validate the architecture (Fig. 6). The analog front-end and the ΔΣ modulator are implemented in a standard AustriaMicroSystems CMOS 0.35-µm technology (through Europractice Alliance) to limit interference, noise sources and input capacitance. The monolithic integrated circuit (IC) embeds two independent acquisition channels to improve parallelization of the acquisitions and allow for differential measurement: for instance, one channel can be used to estimate the noise and take it into consideration, as discussed in section III.G.

The analog front-end is the OPA-based implementation of the architecture described in Fig. 3. The implemented ADC is a second-order switched-cap 1-bit ΔΣ modulator sampling at fS = 10 MSa/s. FIR filters and data decimators are implemented on a flexible field-programmable-gate-array (FPGA) unit. Data is then sent to an external interface that could be a PC, for real-time measurement, a memory or a RF transceiver for wireless-sensor-node applications.

Fig. 6 A) Block scheme of the prototype. B) Photograph of the test board hosting the picoammeter CMOS IC.

The FPGA also embeds a command unit module for programming the CMOS IC via SPI communication. The entire PCB platform is supplied at 5 V but the IC needs a regulated 3.3-V supply or, alternatively, a dual ±1.65-V supply. Each channel of the ammeter prototype features two input ranges, i.e., ±20 nA and ±200 pA (implemented by setting different values of gains A1, A2, A3), while the acquisition bandwidth can be changed by the user simply varying the OSR used in decimators (the available bandwidths implemented in the prototype are 625 Hz, 5 kHz, 10 kHz and 100 kHz). The maximum achievable bandwidth is bounded to 700 kHz by the LPF in the analog front-end (Fig. 3).

The prototype was characterized by considering the circuit-based equivalent model depicted in Fig. 7. The picoammeter presents a single input terminal (IN) per channel and applies a user-selectable bias voltage Vc on this node. This is an important feature when the ammeter is employed in the characterization of high-impedance devices, such as insulators or cell membranes. In fact, the user can change the bias VC, or even apply a time-variant stimulus, and characterize the current response. However, the fact that the integrator-based architecture inherently applies a bias at the measurement node can represent an issue in those applications in which an external biasing is undesired, or such a measurand-independent DC burden voltage cannot be adequately compensated by means of an external DC matching network.

The random process δi discussed in section II is modeled by the equivalent current generator in parallel to the noiseless ammeter. Input impedance can be represented by a parallel capacitor between input and ground and a series resistance. It is worth noting that pico-current sources have typical output impedance higher than 1 MΩ, hence an input impedance of the ammeter in the kΩ-range is fine, leading to minimal burden voltage.

In the following, static response and transfer function are characterized firstly, since they describe the ideal input-output relationship; then equivalent input noise, input impedance and non-linearity effects are analyzed. Finally, the architecture is validated under typical operating conditions by considering two different, simple case studies.

Fig. 7 Circuit-based equivalent model of the proposed picoammeter.

## STATIC RESPONSE

The investigation on the static response of the prototype was carried out by means of the set-up of Fig. 8. The circuit is composed of a variable DC voltage generator and a high-valued precision resistor RI, which are aimed at implementing a pico-current generator, and the 8-1/2 digit multimeter (DMM) HP 3458A exploited for the indirect characterization of the current. A Keithley 6514 Electrometer was used to estimate RI = 995 MΩ with a relative uncertainty of 0.1%. The entire set-up is placed inside a Warner Faraday cage to limit AC interferences. The ammeter is set to the range 200 pA and the minimum implemented bandwidth (625 Hz). By varying the applied DC voltage it is possible to explore the static response of the ammeter throughout the entire input range. Fig. 9 compares the estimates at the output of the prototype with the reference input current, indirectly measured by the DMM with a relative uncertainty of 0.2%. A good linear relationship can be observed, with a root mean square deviation of 1.5 pA from the ideal response, that is 0.4% of the full range. The prototype shows a 30-pA input-referred offset mainly because of intrinsic offset of the first-stage OPA. This offset could be easily reduced in future implementations using conventional software techniques, such as auto-zero. However, in the characterization of Fig. 9 we opted to not implement software corrections, in order to focus on the intrinsic performance of the hardware architecture only.

Fig. 8 Measurement set-up implemented for the characterization of the DC response of the picoammeter. The system composed of a voltage generator, DMM and high-valued resistor RI realizes a pico-current reference source with relative uncertainty of 0.2%.

Fig. 9 DC response of the picoammeter prototype in the 200-pA range, compared to reference values (r.m.s. deviation = 1.5 pA). Squared points denote measurement results while dashed line denotes the linear regression.

## FREQUENCY RESPONSE

In the following the equivalent transfer function is defined as the ratio between the estimated spectrum and the input spectrum. To this end, a reference white noise current is generated using a 475 kΩ, 0.1% metal film resistor connected to the input node of the prototype. Results in the range 20 nA are reported in Fig. 10, where the output has been normalized over the thermodynamic spectral density 4kT/R. The frequency response is white up to 100 kHz, demonstrating the wide-bandwidth capability of the architecture, which is limited only by OPA implementation. Obviously, the choice of FIR filter and sample decimator affects the actual bandwidth: in Fig. 10 the response obtained with the 150-taps sinc3 filter directly implemented on the FPGA (OSR = 50) is compared to that employing a 800th-order FIR filter with triangular window, with (OSR = 50) and without decimator. The sinc3 filter starts to attenuate the signal around 20 kHz, while the 800th-order FIR filter gives a flat transfer function up to 80 kHz, four times higher than the sinc3. Also decimation process affects the bandwidth, as clearly visible comparing red and black lines in Fig. 10. Spikes at 10 kHz and multiples are related to the time-variance of the architecture.

Fig. 10 Transfer function of the ammeter obtained by exploiting different digital filters. The choice of the filter strongly affects the final acquisition bandwidth. Spikes at 10 kHz and multiple frequencies are related to the time-variance of the architecture.

## EQUIVALENT INPUT-REFERRED NOISE

Fig. 11 reports the input-referred noise PSD  experimentally estimated by i) setting the prototype at the range 200 pA and by limiting the bandwidth to 10 kHz (blue line) using the 150-tap sinc3 filter, and ii) setting the range to 20 nA and by limiting the bandwidth to 100 kHz (black line) using the 800th-order FIR filter. The minimum achievable r.m.s. noise is 4 fA/√Hz for the 200-pA range and 30 fA/√Hz for the 20-nA range, that are state-of-the-art noise performances for compact picoammeter [3,19,20]. Again, spikes shown in the spectrum above 10 kHz are the result of the condition. In correspondence with the experimental results for the 20-nA range, Fig. 11 also reports, for comparison purposes, both the theoretical folding noise contribution and the noise lower bound  in Eq. (7), which have been numerically evaluated by using the parameters reported in figure caption.

Fig. 11 Input referred noise density measured (open input condition, CS = 0 pF) at 200-pA range and 10-kHz bandwidth (blue line), and 20-nA range and 100-kHz bandwidth (black line). Dashed lines report the components of Eq. (7) numerically predicted with Gen = 9x10-18 V2/Hz (white shaped), Cf = 10 pF, Cin = 3 pF, τR = 4.8 µs, fR = 9.76 kHz and Req = 22.5 MΩ.

## NON LINEARITY

Since the architecture can be regarded as a current-input analog-to-digital converter operated at fS, then standard figures of merit typically exploited for ADC characterization [21] are reported in the following. A reference 950-Hz voltage sinewave was generated using an arbitrary waveform generator and filtered by means of a series of RC LPFs and HPFs to achieve 86 dB of spectral purity. The voltage was then converted into a 17-nA amplitude current sinewave through a 10-MΩ resistor, as described in Fig. 8. The averaged Discrete Fourier Transform (DFT) of the measured data is shown in Fig. 12. The second harmonic level is at -78.8 dBc and the third harmonic is at -81.2 dBc. THD was computed by taking into account the first ten harmonic components [21]. The test was performed at 20-nA range while the acquisition bandwidth was limited to 10 kHz, since spurious components due to time-variance arise at frequencies around nfR = 9.76 kHz and above. This distortion effect will be characterized in the next section.

Fig. 12 Averaged DFT of the current samples when a 17-nA amplitude, 950-Hz sinewave is applied at the input of the prototype.

Fig. 12 shows a noise floor ten times higher than that reported in Fig. 11 (20-nA range) due to the noise produced by the signal generator. For this reason, the SINAD parameter was estimated by summing the power of the harmonic components shown in Fig. 12 with the noise power shown in Fig. 11 (i.e., 30 fA/√Hz or 3 pArms in 10-kHz bandwidth for the 20-nA range). The result was SINAD = 69.3 dB, which corresponds to ENOB = 11.5. Values of THD, SINAD and ENOB measured at different test signal conditions and 10-kHz bandwidth are reported in Tab. 1. At the more sensitive range of 200 pA the effective resolution peaks to ENOB = 7.6 (averagely varying around ENOB = 7.3), which is still a good result and corresponds to an effective quantization step of 2 pA at 10-kHz bandwidth.

TABLE I – Standard figures of merit

| Current sinewave and system parameters |  |  | Figures of merit |  |  |
| --- | --- | --- | --- | --- | --- |
| Range | Frequency [Hz] | Amplitude [dBFS] | THD [dB] | SINAD [dB] | ENOB [bit] |
| 20 nA | 950 | -1 | -75.6 | 69.3 | 11.5 |
| 20 nA | 950 | -6 | -77.9 | 65.6 | 11.8 |
| 20 nA | 100 | -1 | -76.9 | 70 | 11.5 |
| 20 nA | 100 | -6 | -75.8 | 65.3 | 11.8 |
| 200 pA | 950 | -1 | -44.3 | 40.6 | 6.8 |
| 200 pA | 950 | -6 | -51 | 43.5 | 7.6 |
| 200 pA | 100 | -1 | -47.9 | 43.5 | 7.2 |
| 200 pA | 100 | -6 | -48.4 | 42.5 | 7.6 |

## EFFECTS OF TIME-VARIANCE

The intermodulation distortion caused by the time-variance of the architecture (see Section II-B) was quantified in simulation applying a 100-pA (Fig. 13-top) and a 2.5-nA (Fig. 13-bottom) current sine wave at 2 kHz (Fig.13-left) and 50 kHz (Fig.13-right). In the case of an input signal with low incremental ratio (i.e., 100-pA sine wave at 2 kHz) the benefits of the T&H are fully exploited and side-bands are below the input-referred noise level, which is not modeled in these simulations. Spikes shown at frequencies  are still related to the time-variant functionality of the system; specifically they are the result of capacitive coupling between the clock signal controlling the reset operation and the output node, hence it is a clock-feedthrough effect. Increasing the mean incremental ratio Δi/τR of the input signal, by means of either higher amplitude (2.5 nA at 2 kHz) or higher frequency (100 pA at 50 kHz), reduces the effectiveness of the T&H, and non-negligible side-bands appear in the spectra (Fig. 14 top-right and bottom-left). The amplitudes of these side-bands do not depend on the amplitude of the input signal but only on Δi/τR, which is the same for Fig. 13-top-right and Fig. 13-bottom-left. If Δi/τR is further increased (2.5 nA at 50 kHz) the T&H fails to counteract the time-variant nature of the system, and the side-bands achieve their maximum value, which is around -25 dBc in the realized prototype. Spectra shown in Fig. 13 are the result of quasi-periodic-steady-state (qpss) simulations, which is a standard simulation tool for the analysis of RF mixers.

Time-variance is the major reason of spectral distortion. Tones due to clock feedthrough are as high as the second harmonic when a 17 nA sine wave is applied to the ammeter (Fig. 12). Intermodulation side-bands can be even worse than clock feedthrough, as shown in Fig. 13 (bottom-right). Although intermodulation distortion is the main drawback of the proposed architecture, in some applications, such as the monitoring of trans-membrane current, the effect of time-variance is not of main concern, since the typical measurand is a train of pulses and significant information lies on the time response and amplitude step value. In addition, the effect can be attenuated by reducing the duration of the reset phase or by increasing the reset period, as discussed in Sec. II-B, so as to make the reset event more sporadic. Alternatively, if the duration of the reset is reduced to be less than TS,eq/4, where TS,eq is the equivalent Shannon sampling period, than it is possible to use non-uniform sampling strategies that sample always out of the reset phase and allow for complete signal reconstruction [22]. Another technique for attenuating time-variance effects is the ping-pong architecture, where two identical analog front-ends work in time-complementary mode [23]. However, unavoidable mismatches between the channels must be properly managed, otherwise time-variance effects can be even worse than reported by our architecture.

Fig. 13 Simulated distortion effects due to time-variant nature of the proposed architecture when a sine wave current is applied to the input with small time derivative (top-left), medium time derivative due to higher frequency (top-right), the same medium time derivative achieved by increasing amplitude (bottom-left), and  high time derivative (bottom-right).

## INPUT IMPEDANCE

From a general standpoint, a fundamental property of an ammeter is the equivalent input impedance, which introduces a parasitic AC burden voltage that affects the measurement. Since the proposed architecture is based on a transimpedance amplifier topology, a very low value of  cannot be expected. In addition, frequency-dispersion is typical of these solutions. The ammeter-voltmeter measurement process of the input impedance of a picoammeter is a challenging task since it requires the r.m.s. estimation of voltages in the µV-range. The measurement setup exploited (Fig. 15-a) is very similar to the one of Fig. 8 and employs the same instrumentation, but now the value of the stray capacitance CI is critical. This capacitance was estimated using the setup of Fig. 14-b, where a reference 2-kΩ 0.1% metal film resistor replaces the ammeter. Fitting DMM measurements with SPICE simulations based on the set-up electrical model led to a parasitic capacitance of CI = 340 fF with a relative uncertainty of 1%. To increase the voltage drop measured by the DMM, a 10-kΩ 0.1% metal film resistor is placed in series to the ammeter in Fig. 14-a. All noise sources, including equivalent noise generated by the ammeter were taken into account when measuring voltages at its input in order to improve the quality of the measurement process.

Fig. 15 reports the magnitude of the input impedance measured from 2 Hz up to 10 kHz, showing the typical behavior of a transimpedance amplifier. The input impedance varies from 1 kΩ at 2 Hz to 10 kΩ at 10 kHz, which is good for a picoammeter, since pico-current generators have typical output impedance higher than 1 MΩ. From this standpoint, the frequency dispersion shown by  is nearly negligible. The relative uncertainty of the measurements shown in Fig. 15 is not constant with frequency but varies from 2% at 2 Hz to 25 % at 10 kHz, because the impedance ZI, that converts the reference voltage into a reference current, changes its behavior from purely resistive, with uncertainty of 0.1% over the nominal value, to purely capacitive, with uncertainty of 1% over the estimated value.

Fig. 14 a) Measurement setup for the estimation of the input impedance of the ammeter. b) A reference resistor is exploited for the estimation of the parasitic capacitance C1, by fitting DMM measurements with the predictions obtained by means of the setup electrical model implemented in SPICE.

Fig. 15 Input impedance magnitude of the ammeter measured from 2 Hz up to 10 kHz.

## VALIDATION

Two different tests were carried out on the prototype in order to experimentally validate the proposed architecture and assess its capabilities to measure pA currents.

In the first test, the target is measuring the short-circuit equivalent noise current generated by a reference 1-MΩ 0.1% metal film resistor and verify the thermodynamic expression for its spectral density. By considering a bandwidth of 100 kHz, such expression predicts a value of about 40 pA of r.m.s. noise.  The resistor is connected to channel 1, while channel 2 is used to run-time evaluate and de-embed the noise generated by the ammeter. Measurement results, compared to the theoretical thermal noise, are reported in Fig. 16 showing a very good match, with a deviation of less than 1E-26 A2/Hz corresponding to 300 fA (r.m.s.) in 100 kHz bandwidth. To achieve such a result, all the parameters characterized in previous sections are taken into account.

Fig. 16 Measurement of the short-circuit noise current PSD of a 1-MΩ 0.1% metal film resistor by means of the picoammeter prototype. Comparison with the theoretical level (red line).

In the second test, the low-frequency (LF) voltage noise of the ADP123 LDO regulator is measured by means of the prototype and compared to data given by the manufacturer. The ADP123 is regulated at VLDO,OUT = 3.3 V: for this operating condition, the manufacturer fully reports device characterization data in terms of LF equivalent noise spectral density up to 4 kHz. The circuit shown in Fig. 17 is used to scale down the DC voltage to a more suitable level and convert the noise voltage into a current. The resistor Rconv used for V/I conversion is a 1-MΩ 0.1% metal film resistor. The thermal noise generated by R1 and R2 is negligible since they are low-value resistors (i.e., 1 kΩ); on the contrary, the thermal noise transferred by RCONV is measured and subtracted from the final result, as well as the intrinsic noise of the ammeter. The ammeter was set to the 200 pA range that provides the maximum sensitivity, while the bandwidth is limited to 5 kHz by acting on the FIR filter. Fig. 18 compares the result of the measurement with manufacturer data given for LDO output voltage of 3.3 V and output current of 1 mA, demonstrating the capability of the architecture to accurately measure very low signals.

Fig. 17 Circuit implemented to scale down the output voltage of the LDO and convert the voltage noise into a current noise.

In order to provide a comparison of the performance achieved in this case study by the proposed architecture with that of a state-of-the-art commercial ammeter, the test was carried out also by using the Axon Axopatch 200B [9]. This instrument is based on classical int-diff architecture and achieves state-of-the-art noise performance with 100-kHz bandwidth. It allows for a maximum DC bias voltage at the input node of a few hundreds of mV, and thus required marginal changes on the measurement circuit of Fig. 17, namely: i) the ADP123 was regulated at the minimum allowed output voltage VLDO,OUT = 0.8 V; ii) R1  was changed into a 3 kΩ-resistor in the same technology, so as to obtain a voltage scale factor of 1/4. Measurement results are included in Fig. 18 and compared to manufacturer noise data, which were extrapolated down to VLDO,OUT = 0.8 V assuming a linear relationship between flicker-like LF rms noise and bias level, which is widely recognized in the literature. Even though the Axon Axopatch 200B (bandwidth limited to 10 kHz) is able to accurately characterize the LF noise of the LDO at very low frequencies, it suffers from strong power line interference (see spikes at 50 Hz and multiple frequencies), and presents a high noise floor, which perturbs the measurement at frequencies higher than 200 Hz.

Fig. 18 Measurements of the LF noise rms spectral density of the ADP123 LDO compared with manufacturer data.

# Conclusion

This paper presented the architecture and the related operative solutions of a broadband, portable, low-noise picoammeter with effective resolution in the order of 1 pA and below. It is based on an integrating-differentiating scheme to achieve the optimal performance in terms of input-referred equivalent noise but overcomes the typical noise-performance/bandwidth trade-off of similar topologies by decorrelating the A/D sampling frequency from the periodical reset of the analog front-end. More precisely, the int-diff scheme operates with periodical reset at frequency fR so as to avoid saturation of the integrator stage, while the A/D converter samples the output of the int-diff front-end at a frequency fS >> fR. Thanks to this solution, the very low-noise performance is preserved while the input bandwidth is now limited only by the frequency response of the operational amplifiers employed and the speed of the implemented ADC.

The main drawback is an inherent time-variance introduced in the system, which leads to a potentially critical distortion effect on the current measurement. Strictly speaking, time-variance leads to intermodulation distortion around frequencies . In order to strongly mitigate this issue, a T&H circuit synchronized with the reset phase is exploited in the architecture. This is a good and simple solution for input signals with moderate incremental ratio. Other solutions, such as non-uniform sampling, have been proposed to ease the intermodulation distortion.

In order to validate the proposed architecture, a dual-channel prototype was implemented in a low-cost CMOS 0.35 µm technology to limit interference and noise sources. A delta-sigma modulator was used as A/D converter ensuring sampling at high frequencies and flexibility on the acquisition bandwidth. Finally, the digital data elaboration was implemented on a flexible and low-cost FPGA. The prototype provides two acquisition ranges, i.e. ±200 pA and ± 20 nA, while the bandwidth can be easily adjusted acting on the digital filters up to 100 kHz. The maximum achievable bandwidth of the prototype is limited to 700 kHz by analog LPF implemented in hardware. The prototype showed a linear static response, with a residual non-linear error of 2 pA, and a flat frequency response, apart from spikes at frequencies multiple of fR due to clock-feedthrough.

The prototype demonstrated low noise capabilities (minimum noise floor of 4 fA/√Hz), wide acquisition bandwidth (maximum bandwidth >100 kHz), good input impedance (<10 kΩ), and maximum measured effective number of bits ENOB=11.8 at fIN = 950 Hz, -6 dBFS (±20-nA range). Further experimental validation was provided by considering two simple case studies, in which the architecture operates under conditions typical of practical applications: the measurement of broadband thermal noise generated by 1-MΩ resistor, where the prototype exhibited a very good accuracy of 300 fA with all the non-idealities taken into account, and the characterization of flicker noise generated by a commercial voltage regulator.

Acknowledgment

This work was funded by the Engineering and Physical Sciences Research Council (UK - EP/H044795/1) and by EU and National funds (IT) under the ENIAC-JTI project LAB4MEMs. “The data from this paper can be obtained from University of Southampton repository http://doi.org/10.5258/SOTON/404864 “

[1]	K.-W. Ang, S. Zhu, M. Yu, G.-Q. Lo, D.-L. Kwong, High-Performance Waveguided Ge-on-SOI Metal–Semiconductor–Metal Photodetectors With Novel Silicon–Carbon (Si : C) Schottky Barrier Enhancement Layer, IEEE Photonics Technology Lett. 20 (2008) 754–756. doi:10.1109/LPT.2008.921092.

[2]	M. Crescentini, F. Thei, M. Bennati, S. Saha, M.R.R. de Planque, H. Morgan, et al., A Distributed Amplifier System for Bilayer Lipid Membrane (BLM) Arrays With Noise and Individual Offset Cancellation, IEEE Trans. on Biomed. Circuits and Syst. 9 (2015) 334–344. doi:10.1109/TBCAS.2014.2346402.

[3]	J. Rosenstein, M. Wanunu, C.A. Merchant, M. Drndic, K.L. Shepard, Integrated nanopore sensing platform with sub-microsecond temporal resolution, Nat. Methods. 9 (2012) 487–92. doi:10.1038/nmeth.1932.

[4]	C. Yang, S.R. Jadhav, R.M. Worden, A.J. Mason, Compact low-power impedance-to-digital converter for sensor array microsystems, IEEE J. Solid-State Circuits. 44 (2009) 2844–2855. doi:10.1109/JSSC.2009.2028054.

[5]	M. Carminati, G. Ferrari, A.P. Ivanov, T. Albrecht, M. Sampietro, Design and characterization of a current sensing platform for silicon-based nanopores with integrated tunneling nanoelectrodes, Analog Integr. Circuits and Signal Process. 77 (2013) 333–343. doi:10.1007/s10470-013-0193-9.

[6]	S.R. Kodigala, S. Chattopadhyay, C. Overton, I. Ardoin, Fowler–Nordheim electron tunneling mechanism in Ni/SiO2/n-4H SiC MOS devices, Solid-State Electron. 114 (2015) 104–110. doi:10.1016/j.sse.2015.08.017.

[7]	K. Lau, C. Tan, B. Ng, K. Li, R. Tozer, J. David, et al., Excess noise measurement in avalanche photodiodes using a transimpedance amplifier front-end, Meas. Sci. and Technol. 17 (2006) 1941. doi:10.1088/0957-0233/17/7/036.

[8]	Keysight B2980A Series Femto/Picoammeter Electrometer/High Resistance Meter datasheet, 2014.

[9]	I. Axon Instruments, The Axon guide for electrophysiology & biophysics laboratory techniques, Axon Instruments, 1993.

[10]	M. Crescentini, M. Bennati, M. Carminati, M. Tartagni, Noise Limits of CMOS Current Interfaces for Biosensors: A Review, IEEE Trans. on Biomed. Circuits and Syst. 8 (2013) 278–292. doi:10.1109/TBCAS.2013.2262998.

[11]	G. Ferrari, F. Gozzini, A. Molari, M. Sampietro, Transimpedance amplifier for high sensitivity current measurements on nanodevices, IEEE J. Solid-State Circuits. 44 (2009) 1609–1616.

[12]	M. Bennati, F. Thei, M. Rossi, M. Crescentini, G. D’Avino, A. Baschirotto, et al., 20.5 A Sub-pA ΔΣ Current Amplifier for Single-Molecule Nanosensors, in: 2009 IEEE International Solid-State Circuits Conference - Digest of Technical Papers, IEEE, 2009: pp. 348–349,349a. doi:10.1109/ISSCC.2009.4977451.

[13]	C. Enz, G. Temes, Circuit techniques for reducing the effects of op-amp imperfections: autozeroing, correlated double sampling, and chopper stabilization, Proc. of the IEEE. 84 (1996) 1584–1614. doi:10.1109/5.542410.

[14]	J. Phillips, K. Kundert, An Introduction to Cyclostationary Noise: Noise in mixers, oscillators, samplers, and logic, in: Proceedings of the IEEE 2000 Custom Integrated Circuits Conference (CICC), IEEE, 2000: pp. 431–438. doi:10.1109/CICC.2000.852702.

[15]	W.A. Gardner, Introduction to Random Processes with Applications to Signal & Systems, 2nd ed., McGraw-Hill, New York, NY, USA, 1990.

[16]	M. Crescentini, M. Bennati, S. Saha, J. Ivica, M. de Planque, H. Morgan, et al., A Low-Noise Transimpedance Amplifier for BLM-Based Ion Channel Recording, Sensors. 16 (2016) 709. doi:10.3390/s16050709.

[17]	R. Schreier, G.C. Temes, Understanding Delta-Sigma Data Converters, Wiley-IEEE press, Piscataway, NJ, USA, 2004.

[18]	M. Crescentini, M. Rossi, M. Bennati, F. Thei, A. Baschirotto, M. Tartagni, A nanosensor interface based on delta-sigma Arrays, in: Proc. of the 3rd International Workshop on Advances in Sensors and Interfaces (IWASI), IEEE, 2009: pp. 163–167. doi:10.1109/IWASI.2009.5184788.

[19]	H.M. Jafari, R. Genov, Chopper-Stabilized Bidirectional Current Acquisition Circuits for Electrochemical Amperometric Biosensors, IEEE Transactions on Circuits and Systems I: Regular Papers. 60 (2013) 1149–1157. doi:10.1109/TCSI.2013.2248771.

[20]	C.-L. Hsu, A.G. Venkatesh, H. Jiang, D.A. Hall, A hybrid semi-digital transimpedance amplifier for nanopore-based DNA sequencing, in: Proc. of 2014 IEEE Biomedical Circuits and Systems Conference (BioCAS), IEEE, Lausanne (CH), 2014: pp. 452–455. doi:10.1109/BioCAS.2014.6981760.

[21]	IEEE Standard for Terminology and Test Methods for Analog-to-Digital Converters, 2000.

[22]	J. Thomas, On Some Stability and Interpolatory Properties of Nonuniform Sampling Expansions, IEEE Transactions on Circuit Theory. 14 (1967) 404–408. doi:10.1109/TCT.1967.1082745.

[23]	S. Dai, R.T. Perera, Z. Yang, J.K. Rosenstein, A 155-dB Dynamic Range Current Measurement Front End for Electrochemical Biosensing, IEEE Transactions on Biomedical Circuits and Systems. 10 (2016) 935–944. doi:10.1109/TBCAS.2016.2612581.

A non-stationary noise is said to be cyclostationary in the wide sense when its statistical moments are periodic in time [15].