Improved sensitivity and limit-of-detection of lateral flow devices using spatial constrictions of the flow-path

Ioannis N. Katis*, Peijun J. W. He, Robert W. Eason and Collin L. Sones

Optoelectronics Research Centre, University of Southampton, Highfield Campus, SO17 1BJ, Southampton, UK

*Corresponding author. E-mail address: I.Katis@soton.ac.uk

Abstract

We report on the use of a laser-direct write (LDW) technique that allows the fabrication of lateral flow devices with enhanced sensitivity and limit of detection. This manufacturing technique comprises the dispensing of a liquid photopolymer at specific regions of a nitrocellulose membrane and its subsequent photopolymerisation to create impermeable walls inside the volume of the membrane. These polymerised structures are intentionally designed to create fluidic channels which are constricted over a specific length that spans the test zone within which the sample interacts with pre-deposited reagents. Experiments were conducted to show how these constrictions alter the fluid flow rate and the test zone area within the constricted channel geometries. The slower flow rate and smaller test zone area result in the increased sensitivity and lowered limit of detection for these devices. We have quantified these via the improved performance of a C-Reactive Protein (CRP) sandwich assay on our lateral flow devices with constricted flow paths which demonstrate an improvement in its sensitivity by 62x and in its limit of detection by 30x when compared to a standard lateral flow CRP device.

Keywords: Lateral flow devices; Photo-polymerisation; Increased sensitivity

- Introduction

Point-of-care (POC) rapid diagnostics have seen tremendous development in recent years, as the emphasis of care is shifting towards prevention and early detection of diseases. Such testing gives immediate results in non-laboratory settings, does not require any technical skills from the user and is relatively inexpensive, making it ideal for use not only in poorly-resourced, remote areas of developing countries, but also in hospitals and at the bedside within well-developed nations. One example of such POC diagnostics is paper-based fluidic devices, that were first proposed by the Whiteside’s group in 2007(Martinez et al. 2007), after which the field has been progressing rapidly(Cate et al. 2015). A subcategory of POC diagnostics, lateral flow devices (LFDs), are currently the primary means for the detection of a variety of analytes and the most common example of such devices is the pregnancy test that can be easily bought over the counter at a pharmacy or local shop. These are POC devices in their simplest form and even though their range of applications has been widely expanded to detection of other diseases or conditions, there has not been any other major advancement in their capabilities, which therefore limits their wider clinical uptake. LFDs possess many intrinsic advantages such as being low-cost, mass-producible, disposable, equipment free, and requiring no external power to operate.

LFDs are constructed using porous media such as paper and other types of materials, namely nitrocellulose, that facilitate various tasks, one being the immobilisation of the capture antibodies to constitute the test and control lines of the LFD. There is no need for external equipment as the sample and other reagents are transported by capillary action within these porous materials. LFDs enable implementation of two commonly used assay formats, namely the sandwich and the competitive assay.(Koczula and Gallotta 2016) In the sandwich format, which is used in this report, the analyte in the sample binds to a labelled detection antibody, usually conjugated with gold nanoparticles, and the complex is transported along the strip until it encounters a capture antibody which is specific to the analyte and is locally immobilised on the nitrocellulose membrane in the form of a test line or test spot. The capture-analyte-detection antibody complex is then visible as a coloured test line (or spot) that can be seen by eye or can be measured using a detector such as a scanner or camera. After the test line there is the control line which is used to confirm that the assay is working properly and has immobilised antibodies that are specific instead to the labelled detection antibody. Capture of these labelled detection antibodies at the control line again produces a colour change that verifies the correct completion of the assay. Surplus reagents, buffers, and assay fluids are further transported toward the absorbent pad that collects unwanted/unused waste.

One of the main disadvantages of LFDs is their relatively low sensitivities and poor limit of detection, which primarily restricts their use to detection of samples with highly abundant analytes.(Sharma et al. 2015){Sharma, 2015 #50},(Posthuma-Trumpie et al. 2009) A critical factor that determines the limit of detection and sensitivity of a lateral flow assay is the immunoreaction time between the capture antibodies pre-deposited at a test/control line and the analyte, which therefore depends on the migration time of the sample over the test line.(Bahadır and Sezgintürk 2016)

There has been significant research aimed at developing better LFDs, both with increased sensitivity and lowered limit of detection, and several different strategies have been followed. To enhance the signal output of LFDs, many new labels have been trialled recently such as silver enhancement technology,(Anfossi et al. 2013) enzymatic amplification(Parolo et al. 2013a) and magnetic nanoparticles.(Tang et al. 2009) Other proposed strategies include changing the architecture of an LFD to improve its sensitivity such as increasing the size of the sample and conjugate pad(Parolo et al. 2013b) or introducing wax pillars in the flow path in order to reduce the flow rate.(Rivas et al. 2014) In this current work, we present a novel strategy of using spatial constrictions in the flow path as a route to increasing the sensitivity and lowering the limit of detection of LFDs. The novelty of the technique presented here is that it allows the enhancement of the lateral flow devices regardless of the assay used, while maintaining the same materials and the same size of the device, therefore making the technique ideal for adaptation in a manufacturing process.

To demonstrate the advantages of specifically designed flow-paths for an improved LFD performance, in this report we have chosen to implement the detection of C-reactive protein (CRP). A schematic of our proposed constricted flow LFD used to demonstrate enhanced performance can be seen in Figure 1, which provides a visual comparison of its design with that of a standard LFD.

Figure 1. Schematic of (a) the constricted flow device and (b) a standard lateral flow device

Figure 1. Schematic of (a) the constricted flow device and (b) a standard lateral flow device

CRP was chosen as the target analyte as it is a well understood inflammation marker,(Ahn et al. 2003) heart attack marker(Leung et al. 2005) and is also an excellent indicator for whether an infection is caused by a viral or a bacterial pathogen.(Putto et al. 1985) CRP is normally present in very low concentrations in healthy individuals (<5 μg/ml) but rises significantly for unwell patients to values greater than 500 μg/ml.(Black et al. 2004) High concentrations of CRP in the bloodstream are not indicative of a specific disease, but quantitative measurements can provide valuable information for the diagnosis and treatment of a patient, and this is why CRP detection is broadly used in the primary care setting in conjunction with testing for other biomarkers.

- Materials and methods

  - LDW manufacturing technique

In our previous reports, we have demonstrated the usefulness and versatility of a laser direct-write (LDW) approach in the manufacturing of diagnostic devices in porous materials such as cellulose(He et al. 2015a; Sones et al. 2014) and nitrocellulose(He et al. 2015b), as well as in creating 3d structures in such membranes(He et al. 2016). Several other approaches for patterning membranes that have also been reported are photolithography,(Martinez et al. 2008) wax printing,(Carrilho et al. 2009) inkjet printing,(Li et al. 2010) laser cutting,(Fu et al. 2010) plasma treatment(Li et al. 2008) and flexographic printing.(Olkkonen et al. 2010)

Compared to these alternative techniques, LDW has some intrinsic advantages: it is a non-lithographic approach with high flexibility, and furthermore has small upfront equipment costs and no special laboratory and material requirements and hence has the potential to be up-scaled for mass-production of paper-based devices at affordable costs.

Our previously reported LDW manufacturing technique relied on the immersion of the membrane in a photopolymer solution, the subsequent photopolymerisation with a laser beam, followed by the immersion of the membrane in a solvent to remove any unpolymerised residue. A decrease in the total number of fabrication-steps would however not only make this LDW process more efficient as a consequence of reduced fabrication times, but would also make it more cost-effective because of the reduced usage of expensive reagent, translating it into a truly mature technique appropriate for commercial manufacture. To optimise our original technique, we have included a deposition tool that allows localised deposition of the photopolymer only at specific locations on the paper where the fluid containing wall/structures need to be formed within the porous substrates to define the microfluidic device. This selective photopolymer deposition eliminates the (global) soaking step required to impregnate the photopolymer within the paper, prior to the laser illumination step, and furthermore also makes redundant the subsequent solvent developing step inherent in our original technique.

The LDW setup that allows the implementation of this improvised methodology is illustrated in figure 2. As shown in the schematic, the liquid photopolymer is locally deposited onto the paper substrate with a deposition nozzle at locations pre-defined by the user’s device design. A laser beam subsequently follows the deposition head and illuminates the deposited patterns thereby inducing photo-polymerisation of the polymer, therefore transforming it into a solid state that is impermeable to liquids. The polymerised patterns define the fluidic walls that serve as demarcation barriers that confine and transport the liquids within the paper device.

Figure 2. Schematic of the modified laser-based direct-write procedure, which shows the inclusion of a deposition head for localised deposition of a photopolymer for creating polymerised hydrophobic structures in the porous substrate

Figure 2. Schematic of the modified laser-based direct-write procedure, which shows the inclusion of a deposition head for localised deposition of a photopolymer for creating polymerised hydrophobic structures in the porous substrate

In this report, this LDW technique was utilised to pattern only one component of the LFDs, the nitrocellulose membrane, with variable spatial constrictions to influence the flow of the samples within the device. The constricted flow path is shaped as follows: an inlet that gradually reduces in width over a length of a few mm, followed by a section of reduced constant-width and finally an outlet where the flow path width returns to its original size, as seen in figure 1. For the purpose of measurements of the time a fluid takes to travel through the devices with constrictions, we made and tested LFDs with constriction widths ranging from 1 mm to 5 mm in increments of 0.5 mm. For evaluating the implementation of the CRP assay only devices with constriction widths ranging from 1 mm to 5 mm with a 1 mm increment were used.

  - Reagents, materials and apparatus

The fluidic channels were fabricated within backed nitrocellulose membrane purchased from Sartorius (UniSart CN95) and the absorbent pads were cellulose-based filter papers from GE Healthcare (CF1). The capture and detection antibodies used were from a DuoSet ELISA development system (R&D Systems DY1707). The analyte was recombinant human C-reactive protein expressed in E. Coli (Sigma Aldrich C1617). The control line antibody was an anti-mouse Goat IgG from R&D Systems (AF007). The streptavidin-conjugated gold nanoparticles were obtained from BBI solutions (BA.STP40). The BSA and PBS used in the solutions were obtained from Sigma Aldrich (A2058, P3813). The photopolymer used was Desolite 3471-3-14 from Desotech which is an acrylate-based photopolymer with a viscosity of 10,000 mPa·s at 25°C.

The laser used for the LDW process was a 405 nm continuous wave diode laser (MLDTM 405 nm, Cobolt AB) with a maximum output power of ~110 mW. The basic LDW setup is the same as described in our previous publications relating to the fabrication of 2D microfluidic devices and the implementation of flow-control within these devices, and has been fully optimised via a series of systematic studies (He et al. 2015a; He et al. 2015b; Sones et al. 2014). The results we have achieved and report in the following sections are therefore based on the same patterning procedure with appropriate adjustment of the patterning conditions such as laser power and scan speed. The specialised polymer dispenser was a PICO Pulse jet valve on a 3-Axis EV series automated dispensing system purchased from Nordson. The antibody dispenser was a XYZ3210 platform with a Biojet HR solenoid dispenser from Biodot. The software used for processing all images obtained from the experiments was ImageJ (NIH, USA).

- Results and discussion

3.1. Flow time measurements within constricted membranes

The initial set of experiments involved studying the implications of the constricted flow designs on the flow time within the LFDs. Two sets of measurements were taken for devices with different constriction widths ranging from 5mm (a standard LFD strip width) to 1mm. The first set of measurements determined the time required for 30 μl of the red dye used for our initial tests to be transported, after being completely soaked at the source pad, through the constriction to the absorbent pad. In the second set of experiments, we measured the time required for the red dye to flow past the constriction. All measurements were performed by adding 30 μl of red dye in wells of a microtiter plate and then immersing the inlet of the nitrocellulose membrane in the wells. All measurements were repeated three times. Red dye was used instead of water for better visualisation of the flow and hence providing accuracy in the measurement of the times.

As can be seen in Figure 3a, the smaller the constriction, the narrower is the flow path, and hence the longer is the wicking time. This experiment includes the time for a dry membrane to become fully wetted (measuring the rate of saturation) and the time required for a standard sample volume to be completely absorbed by the device. This increase in the flow time allows for an increased interaction time between the sample and the capture antibody at the test line, therefore producing an enhanced test signal. For the measurements of the second experiment, as seen in Figure 3b, the flow times that were measured at the exit of the constrictions were slightly reduced for the narrowest flow paths. This reduction does not however have a major influence because the total flow times across the entire LFD (as in figure 3a) is almost two orders of magnitude greater. The increased total flow times in figure. 3a result from longer flow times after the constriction because the constriction works as a source that provides a reduced liquid volume, and as reported elsewhere,(Elizalde et al. 2015; Fu et al. 2011) the subsequent flow-rate is determined by the volume of liquid provided by the source.

Figure 3. (a) Time for a 30 μl volume of red dye solution to flow through the whole LFD and be absorbed by the absorbent pad, (b) Time for the fluid front to flow through the constriction, averaged over three measurements and (c) Effect of the constriction width on the area of the test line

Figure 3. (a) Time for a 30 μl volume of red dye solution to flow through the whole LFD and be absorbed by the absorbent pad, (b) Time for the fluid front to flow through the constriction, averaged over three measurements and (c) Effect of the constriction width on the area of the test line

Additionally, the area of each test line was measured from the individual ImageJ images and it was observed that the test area is reduced when the flow path is constricted, as seen in figure 3c. Since the area of the test line is reduced, this presents an enhanced concentration of the capture antibodies for binding with the analyte in the sample directed through this confined area. This constriction also leads to the concentrating of the sample along it. A larger number of analyte molecules will therefore be bound to this smaller area, producing an increased colour intensity, and thus an enhanced sensitivity and limit of detection.

3.2. Increased sensitivity measurements for a CRP assay.

3.2.1. Preparation of LFDs

Photopolymer lines forming the individual strips with different constriction widths (1, 2, 3, 4, 5 mm) were prepared by dispensing polymer with the Nordson dispenser and subsequently polymerising the material according to the LDW technique described earlier. The photopolymer dispensing was done at a speed of 40 mm/s thus ensuring that a continuous polymer line and hence impenetrable fluidic walls are formed within the nitrocellulose membranes. The temperature of the dispensing head was set at 70° C to reduce the viscosity and facilitate the proper flow of the photopolymer through the dispenser. The test and control lines were subsequently produced by dispensing the capture antibody (mouse anti-human CRP in 1% BSA/PBS) and the control antibody (anti-mouse IgG in 1% BSA/PBS) with a Biodot dispenser at a concentration of 1 mg/ml. The test lines and control lines were dispensed on the nitrocellulose membrane in the middle of the constriction, at a distance of 12 mm from the inlet of the device. The reagent droplets dispensed to produce the test and control lines had a volume of 20 nl and a separation of 250 μm, and this ensured the formation of a continuous test/control line. The nitrocellulose membrane was then attached to a standard LFD backing card, and this was followed by the attachment of the cellulose-based absorbent pad such that it overlapped the nitrocellulose membrane by 2 mm. The striped devices were left to dry at room temperature overnight and then the card was cut along the polymerised lines with a manual cutter to obtain individual LFDs.

3.2.2 Assay implementation on LFDs

The CRP-assay which was implemented via the LFDs had a sandwich format, and works as follows. The biotinylated anti-CRP detection antibodies attached with streptavidin-gold nanoparticle complexes first bind to the CRP analyte from the sample. The CRP analyte then binds to the anti-CRP capture antibodies (in the test line), thus producing a colour signal at the test line only when the analyte is present in the sample. The labelled antibodies also bind to the anti-mouse antibodies in the control line, thus producing a coloured signal regardless of the presence or absence of the CRP analyte in the sample. 15 μl of the detection antibody (biotinylated mouse anti-human CRP) was added to a microtiter well at a concentration of 1.6 μg/ml along with 15 μl of streptavidin coupled with 40 nm gold nanoparticles at OD 1, and 10 μl of recombinant human CRP standard at various concentrations ranging from 2.5 μg/ml to 39 ng/ml in a 1% BSA/PBS solution. The gold-conjugated detection antibodies were premixed with the CRP standard in a microtiter well plate for several minutes, and this was done so that the complex achieved an equilibrium condition before it was introduced to the LFD. The LFD was immersed in the conjugate antibody – sample solution, and the device was then left to run for several minutes (the time ranged between 3 min to 6 min depending on the test zone constriction) until all of the solution was wicked through the membrane and the result was immediately documented using a standard scanner (Figure 4). All assays were performed in triplicate using different batches of the same analyte concentrations

3.2.3. Evaluation of the assay results

Following the experiments with varying constriction widths and analyte concentrations, the effect of the constrictions on the colour intensity produced at the test/control lines was evaluated. In order to quantify the results of the CRP assays and produce the colour intensity graphs (Figure 5), the photos were processed using the ImageJ software. The coloured test line was selected and the mean grey values were extracted. These values were then normalised by removing the background noise of each individual LFD strip. The normalised colour intensity was then plotted against the channel width at the constriction, as can be seen in figure 5. For an analyte concentration of 2.5 μg/ml the normalised colour intensity increased ~2-fold for a decrease in the test zone width from 5 mm (the usual dimensions of a standard LFD) to 1 mm. It was observed that for smaller analyte concentrations, there was a substantial increase in colour intensity, for example, 5.4x, from 6.4 to 34.4 a.u., for 312.5 ng/ml and 62x, from 0.42 to 26.1 a.u., for 78 ng/ml. The increase in colour intensities for the different constriction widths becomes more apparent for reduced analyte concentrations. At high analyte concentrations, both the standard 5 mm device and the constricted-flow LFDs perform similarly. At lower analyte concentrations the standard device cannot perform equally well due to its low limits of detection. The ratio of the average colour intensity of the 1 mm devices to that of the average colour intensity of the 5 mm devices was calculated for the same range of analyte concentrations and it was found to be exponentially increasing towards smaller concentrations.The average intensity ratio of a 1 mm constricted device over the standard 5 mm device was 2, 2.2, 3.9, 5.4, 7.9, 62 and 95 for analyte concentrations of 2.5, 1.25, 0.625, 0.3125, 0.156, 0.078 and 0.039 μg/ml respectively. This proves that the constricted flow devices are more suitable for detection of analytes in low concentrations when compared to a standard LFD.

Figure 4. Scanned images of the CRP LFDs with different constrictions showing a) the test and control lines and b) test lines of various CRP analyte concentrations

Figure 4. Scanned images of the CRP LFDs with different constrictions showing a) the test and control lines and b) test lines of various CRP analyte concentrations

Although achieving a decrease of the limit of detection for LFDs has been an objective that has seen considerable research effort, so far LFDs have not been found to be as sensitive as other standard

Figure 5. Increase of the colour intensity measurements by decreasing the width of the test zone at different concentrations of CRP

Figure 5. Increase of the colour intensity measurements by decreasing the width of the test zone at different concentrations of CRP

laboratory techniques, and this has been one of the primary reasons that has hindered their widespread adoption for clinical diagnostic purposes. For our experiments using constricted flow paths, the limit of detection was also observed to be greatly enhanced. As can be seen in figure 6, the limit of detection of the CRP assay on a standard 5 mm wide LFD was ~150 ng/ml. However, on decreasing the width of the constrictions, the limit of detection is also decreased considerably, and for the case of the device with a 1 mm wide constriction, the limit of detection was found to be ~5 ng/ml, a 30-fold enhancement when compared to the standard 5mm wide LFD.

Figure 6. The limit of detection for the CRP lateral flow assay depending on the constriction width on the membrane for three devices of each constriction width

Figure 6. The limit of detection for the CRP lateral flow assay depending on the constriction width on the membrane for three devices of each constriction width

- Conclusion

We have presented a novel strategy for improving the sensitivity and limit of detection of LFDs by fabricating spatial constrictions along their flow path within the nitrocellulose membranes of these devices. From our understanding and interpretation, the enhanced signal that is produced within these constriction-enabled LFDs is because of two reasons - firstly, the test zone is decreased in this area which means that both the AuNP-conjugated antibodies and the sample is therefore concentrated, and secondly, the overall flow rate is decreased, therefore allowing for longer interaction time between the capture antibody and the sample. This therefore allows for the fabrication of LFDs with enhanced sensitivity as they can produce increased colour intensity for the same analyte concentration. We have demonstrated a sandwich CRP assay in order to evaluate designs with different widths and depending on the analyte concentration, these designs exhibited a significant improvement in the observed signal. Additionally, the limit of detection of such a device was found to be lowered considerably, allowing for use of such devices with samples with very low analyte concentrations that would otherwise be impossible to detect. Sensitivity and limit of detection are the two key parameters that LFDs have been inadequate at, and by using a uniquely designed flow path created with a simple laser manufacturing technique we believe we can enable the drive for the use of LFDs as efficient, standard laboratory diagnostic tools.

Conflicts of interest

There are no conflicts to declare.

Acknowledgements

The authors acknowledge the funding received via the Engineering and Physical Sciences Research Council (EPSRC) Grant Nos. EP/N004388/1, EP/P025757/1 and EP/M027260/1. The underpinning RDM data for this paper can be found at . https://dor.org/10.5258/SOTON/D0387

References

Ahn, J.S., Choi, S., Jang, S.H., Chang, H.J., Kim, J.H., Nahm, K.B., Oh, S.W., Choi, E.Y., 2003. Development of a point-of-care assay system for high-sensitivity C-reactive protein in whole blood. Clinica Chimica Acta 332(1), 51-59.

Anfossi, L., Di Nardo, F., Giovannoli, C., Passini, C., Baggiani, C., 2013. Increased sensitivity of lateral flow immunoassay for ochratoxin A through silver enhancement. Analytical and Bioanalytical Chemistry 405(30), 9859-9867.

Bahadır, E.B., Sezgintürk, M.K., 2016. Lateral flow assays: Principles, designs and labels. TrAC Trends in Analytical Chemistry 82, 286-306.

Black, S., Kushner, I., Samols, D., 2004. C-reactive Protein. Journal of Biological Chemistry 279(47), 48487-48490.

Carrilho, E., Martinez, A.W., Whitesides, G.M., 2009. Understanding Wax Printing: A Simple Micropatterning Process for Paper-Based Microfluidics. Analytical Chemistry 81(16), 7091-7095.

Cate, D.M., Adkins, J.A., Mettakoonpitak, J., Henry, C.S., 2015. Recent Developments in Paper-Based Microfluidic Devices. Analytical Chemistry 87(1), 19-41.

Elizalde, E., Urteaga, R., Berli, C.L.A., 2015. Rational design of capillary-driven flows for paper-based microfluidics. Lab on a Chip 15(10), 2173-2180.

Fu, E., Kauffman, P., Lutz, B., Yager, P., 2010. Chemical signal amplification in two-dimensional paper networks. Sensors and actuators. B, Chemical 149(1), 325-328.

Fu, E., Ramsey, S.A., Kauffman, P., Lutz, B., Yager, P., 2011. Transport in two-dimensional paper networks. Microfluidics and Nanofluidics 10(1), 29-35.

He, P.J.W., Katis, I.N., Eason, R.W., Sones, C.L., 2015a. Engineering fluidic delays in paper-based devices using laser direct-writing. Lab on a Chip 15(20), 4054-4061.

He, P.J.W., Katis, I.N., Eason, R.W., Sones, C.L., 2015b. Laser-based patterning for fluidic devices in nitrocellulose. Biomicrofluidics 9(2), 026503.

He, P.J.W., Katis, I.N., Eason, R.W., Sones, C.L., 2016. Laser direct-write for fabrication of three-dimensional paper-based devices. Lab on a Chip 16(17), 3296-3303.

Koczula, Katarzyna M., Gallotta, A., 2016. Lateral flow assays. Essays in Biochemistry 60(1), 111-120.

Leung, W., Chan, C.P., Leung, M., Lehmann, K., Renneberg, I., Lehmann, M., Hempel, A., Glatz, J.F.C., Renneberg, R., 2005. Novel “Digital‐Style” Rapid Test Simultaneously Detecting Heart Attack and Predicting Cardiovascular Disease Risk. Analytical Letters 38(3), 423-439.

Li, X., Tian, J., Garnier, G., Shen, W., 2010. Fabrication of paper-based microfluidic sensors by printing. Colloids and Surfaces B: Biointerfaces 76(2), 564-570.

Li, X., Tian, J., Nguyen, T., Shen, W., 2008. Paper-Based Microfluidic Devices by Plasma Treatment. Analytical Chemistry 80(23), 9131-9134.

Martinez, A.W., Phillips, S.T., Butte, M.J., Whitesides, G.M., 2007. Patterned Paper as a Platform for Inexpensive, Low Volume, Portable Bioassays. Angewandte Chemie (International ed. in English) 46(8), 1318-1320.

Martinez, A.W., Phillips, S.T., Wiley, B.J., Gupta, M., Whitesides, G.M., 2008. FLASH: A rapid method for prototyping paper-based microfluidic devices. Lab on a chip 8(12), 2146-2150.

Olkkonen, J., Lehtinen, K., Erho, T., 2010. Flexographically Printed Fluidic Structures in Paper. Analytical Chemistry 82(24), 10246-10250.

Parolo, C., de la Escosura-Muñiz, A., Merkoçi, A., 2013a. Enhanced lateral flow immunoassay using gold nanoparticles loaded with enzymes. Biosensors and Bioelectronics 40(1), 412-416.

Parolo, C., Medina-Sanchez, M., de la Escosura-Muniz, A., Merkoci, A., 2013b. Simple paper architecture modifications lead to enhanced sensitivity in nanoparticle based lateral flow immunoassays. Lab on a Chip 13(3), 386-390.

Posthuma-Trumpie, G.A., Korf, J., van Amerongen, A., 2009. Lateral flow (immuno)assay: its strengths, weaknesses, opportunities and threats. A literature survey. Analytical and Bioanalytical Chemistry 393(2), 569-582.

Putto, A., Meurman, O., Ruuskanen, O., 1985. C-Reactive protein in viral and bacterial infections. Pediatr Res 19(10), 1103-1103.

Rivas, L., Medina-Sanchez, M., de la Escosura-Muniz, A., Merkoci, A., 2014. Improving sensitivity of gold nanoparticle-based lateral flow assays by using wax-printed pillars as delay barriers of microfluidics. Lab on a Chip 14(22), 4406-4414.

Sharma, S., Zapatero-Rodríguez, J., Estrela, P., Kennedy, R., 2015. Point-of-Care Diagnostics in Low Resource Settings: Present Status and Future Role of Microfluidics. Biosensors 5(3).

Sones, C.L., Katis, I.N., He, P.J.W., Mills, B., Namiq, M.F., Shardlow, P., Ibsen, M., Eason, R.W., 2014. Laser-induced photo-polymerisation for creation of paper-based fluidic devices. Lab on a Chip 14(23), 4567-4574.

Tang, D., Sauceda, J.C., Lin, Z., Ott, S., Basova, E., Goryacheva, I., Biselli, S., Lin, J., Niessner, R., Knopp, D., 2009. Magnetic nanogold microspheres-based lateral-flow immunodipstick for rapid detection of aflatoxin B2 in food. Biosensors and Bioelectronics 25(2), 514-518.