**Sensing the unreachable: challenges and opportunities in biofilm detection in medical settings**

**Yikang Xu1,2, Yousr Dhaouadi1,2, Paul Stoodley3,4, Dacheng Ren1,2,5,6**

1Department of Biomedical and Chemical Engineering, Syracuse University, Syracuse, NY 13244

2Syracuse Biomaterials Institute, Syracuse University, Syracuse, NY 13244, United States

3 Departments of Microbial Infection and Immunity and Orthopedics and the Infectious Diseases Institute, Ohio State University, Columbus, OH 43210, United States

4National Centre for Advanced Tribology at Southampton (nCATS), Mechanical Engineering, University of Southampton, United Kingdom

5Department of Civil and Environmental Engineering, Syracuse University, Syracuse, NY 13244, United States

6Department of Biology, Syracuse University, Syracuse, NY 13244, United States

\*Corresponding author:

Dacheng Ren: Phone +1-315-443-4409. Fax +1-315-443-9175. Address : 329 Link Hall, Syracuse University, Syracuse, NY 13244, United States. Email : dren@syr.edu

**Abstract**

Bacteria can attach to essentially all materials and form multicellular biofilms with high-level tolerance to antimicrobials. Detrimental biofilms are responsible for a variety of problems ranging from food and water contamination, bio-corrosion, to drug resistant infections. Besides the challenges in control, biofilms are also difficult to detect due to the lack of biofilm-specific biomarkers and methods for non-destructive imaging. In this article, we present a concise review of recent advancements in this field, with a focus on medical device-associated infections. We also discuss the technologies that have potential for non-destructive detection of bacterial biofilms.

**Bacterial biofilms.** As the oldest life form on earth, bacteria have developed remarkable capabilities to survive in harsh environments. One such strategy is to colonize surfaces and form biofilms, which are complex structures of bacterial cells embedded in an extracellular matrix comprised of polysaccharides, proteins, DNA and lipids produced intrinsically by these bacteria [1]. Biofilm cells are up to 1,000 times more tolerant to antibiotics than their planktonic counterparts, rendering biofilm-related infections largely unresponsive to antibiotic therapies [2]. Thus, detrimental biofilms are of great concern especially when formed by pathogenic species on implanted medical devices and biomaterials such as indwelling catheters, orthopedic implants and cochlear implants, among others [3-6].

**Challenges in biofilm detection.** Although a number of methods have become quite routine for biofilm characterization in controlled laboratory settings [7], biofilm detection in clinical settings has multiple challenges [8-10]. In general, biofilm associated infections (BAI) are chronic and remain local to the infection sites such as an implant. Additionally, clinical evidence suggests that for long periods of time biofilms can produce occult, or sub clinical, infections in which the inflammatory symptoms are less pronounced than acute infections, only revealing themselves when biofilm cells shed off resulting in bacteremia [11]. Usually BAIs involving foreign bodies are not confirmed until the device is explanted and even then often a diagnosis of biofilm is based on anecdotal evidence [9]. When a biofilm is formed on an implanted medical device, it cannot be directly sampled without a surgical procedure. This essentially precludes the detection using standard microbiological methods such as bacterial culturing on agar plates until intra-operative access. Even if the cells can be sampled, slow-growing variants of bacteria and dormant persister cells [12] may not form colonies under routine clinical culture and thus cause a false negative result [13]. Culturing methods also face challenges with the highly heterogeneous distribution of biofilms and the involvement of fastidious strains or mixed-species biofilms that require specific growth factors [9]. A standardized, reliable method for the detection of biofilms is still missing. Besides BAI diagnosis *in vivo*, microbial detection on explanted medical devices is also challenging if the cells are difficult to sample (if trapped in crevices such as those in endoscopes) or culture (if dormant or missing growth factors). Sampling may not be effective without removing bacteria using a stronger force such as sonication [14]. In this article, we provide a concise review of recent advancements (focusing on the past three years) in biofilm detection and possible future directions (Fig. 1).

|  |
| --- |
|   |
| **Figure 1. Schametic of biofilm formation and potential targets for detection.** Biofilm formation causes significant changes in bacterial gene expression and metabolism, triggers host immue response, and alters the chemical/physical properties of the substrate material. These changes are potential targets for biofilm detection. |

**Molecular methods.** Compared to culturing methods, DNA-based analyses are more sensitive in detecting microbes including unculturable cells and samples with mixed species. Conventional polymerase chain reaction (PCR)-based approaches target ribosomal RNA for detection [15]. The inclusion of specific gene targets can reveal additional information such as antibiotic resistance [16]. Compared to PCR tests that target individual genes, meta-genomic sequencing can reveal more information regarding antibiotic resistance, virulence factors and other important physiological traits. Recent development of whole-genome sequencing (WGS) has drastically improved the throughput and accuracy, and lowered the cost [17]. DNA-based technologies do have limitations, one being the lack of information about the viability of cells and not being able to distinguish if the organisms are in a biofilm or planktonic phenotype (if the cells cannot be separated during sampling). Combining DNA detection with messenger RNA analysis for biofilm-specific gene expression may be able to distinguish the phenotype of the organisms. However, contaminating DNA from the clinical environment (including the patient’s skin, surgical instruments, gloves and irritants) are also a concern.

**Biofilm associated biomarkers.** Although biofilms formed on implanted devices cannot be directly sampled, biofilm growth may produce unique molecules, or stimulate biofilm-specific host responses, that can be detected using standard methods. Antibodies may not be detectable during acute infections, but can be used to help with chronic BAI diagnosis. For example, alpha defensin, an antimicrobial peptide produced by the body to fight infection, has been found in the synovial fluid of infected hip and knee joints and shows good sensitivity and specificity for periprosthetic joint infection (PJI) [18]. However, like antibody tests, it is not biofilm specific.

Besides host factors, identifying biofilm-specific markers on bacterial cells will enable effective detection of biofilms. Since the discovery of biofilm-associated protein (Bap) in *Staphylococcus aureus*, Bap homologs have been found in many bacterial species. These proteins are present on bacterial surfaces and many are involved in biofilm formation [19] and chronic infection in mammary glands [20]. Recombinant subunits of *Acinetobacter baumannii* Bap has been shown to stimulate immune response in mice [21]. However, *bap* was not detected in *S. aureus* isolates from patients with urinary tract infections (UTI) [22]. Further studies are needed to identify makers that are unique to bacterial biofilms especially those present in multiple species.

In addition to cellular targets, biofilm matrix components provide potential markers of biofilms and may offer species identification. One of such possible biomarkers is cellulose, a major component of biofilm matrix of uropathogenic *Escherichia coli* (UPEC). Antypas et al. [23] developed an assay based on optotracing that can detect cellulose in urine in less than 45 min, which can help determine if biofilm is involved in UTI. Similarly, exopolysaccharide (EPS) is a possible marker for detecting *Histophilus somni* biofilms [24]. Chronic lung infection by mucoid *P. aeruginosa* biofilms leads to an IgG antibody response against *P. aeruginosa* components including alginate, lipopolysaccharides and proteins, which is currently used for diagnosis [25]. Interestingly, many species of bacteria incorporate their own extracellular DNA (eDNA) into the matrix and thus cell-free DNA analysis [26] in combination with polysaccharide analysis [27] might provide evidence of biofilm and identify the species involved. Another possibility is to quantify biomass by measuring proteins in biofilms. Guan et al. [28] modified o-phthalaldehyde (OPA) protein assay to achieve extraction-free detection of biofilms.

Quorum sensing (QS) is a well-known system found in numerous microbial species, which enables the cells to regulate cell density-dependent activities by sensing and responding to signaling molecules named autoinducers [29]. At least for lung infection by *P. aeruginosa*, QS signal profiling has been shown to be a potential biomarker for biofilm detection [30]. QS signals are also known to alter the level of immune factors and immune cell proliferation [31,32], which may help identify new biomarkers for biofilm detection. The chronic nature of BAIs indicates that host immunity is unable to eradicate biofilm cells. Thus, it is believed that the innate and acquired immune responses coexist [33]. Deciphering the types and dynamics of immune factors involved in such infections may provide new biomarkers for better diagnosis.

**Biofilm imaging.** Laboratory methods for biofilm culturing, imaging and analysis have been well summarized in recent reviews [7]. Confocal laser scanning microscopy (CLSM), fluorescence in situ hybridization (FISH), and later improved peptide nucleic acid (PNA)-FISH [34] and locked nucleic acids (LNA)-FISH [35], can provide spatial information about biofilms and the location of different species. But these advanced imaging techniques require a clear line of sight with the specimen, which are destructive and not applicable for *in situ* detection/diagnosis. Similar issues exist for scanning electron microscopy (SEM), time-of-flight secondary ion mass spectrometry (TOF-SIMS), transmission electron microscope (TEM) and Mass Spectrometry Imaging (MSI).

In comparison, Raman spectroscopy and surface enhanced Raman spectroscopy (SERS) offer non-destructive molecular detection with high sensitivity [36]. By detecting the photons that change energy when scattering off a material (Raman scattering), Raman spectroscopy is highly sensitive in detecting low-abundant biomolecules, e.g. 5 ppb of *P. aeruginosa* pyocyanin in sputa [37] and quorum sensing signalsin *P. aeruginosa* biofilms in mice [38].

For pre and intra-operative guidance, imaging techniques that can “see” through tissues into the body will arguably have the greatest medical utility. A useful tool for non-invasive biofilm detection is Near Infra-Red (NIR) imaging. Systems with such capabilities have been commercialized, such as the Spectrum In Vivo Imaging System (IVIS) from PerkinElmer, which combines 2D optical imaging and 3D NIR tomography, achieving 3D tracking of bioluminescence and fluorescence signals. Combining IVIS with biofilm-specific markers may improve biofilm diagnosis.

Hyperspectral imaging is a label-free method that has been used in wound monitoring and to detect biofilms in the natural environment. This technology provides visible and near IR spectra in a 2D image and may have future potential for detecting color changes associated with biofilms themselves or reaction of the host tissues to biofilms.

While signs of infection can be diagnosed from observing host tissues in X-rays, direct detection of biofilm will likely require a contrast agent. For example, with appropriate contrast agents such as iron sulfate, X-ray tomography can differentiate biofilm from surrounding water and allow 3D quantification of biofilm structures [39]. Further improvement in biofilm characterization can be achieved using X-ray micro-force computed tomography (μCT), which has been shown for non-destructive analysis of biofilm grown in central venous catheters (CVC) [40]. Sellmyer et al. [41] synthesized [18F] fluoropropyl-trimethoprim and showed it can label bacteria but not inflammatory or cancer cells in rodents, allowing direct visualization of infection using positron emission tomography imaging. Future work to improve sensitivity and reduce background signal from bowel and bone uptake may bring exciting opportunities for non-invasive imaging of biofilm infections.

Currently medical imaging modalities do not have the spatial resolution to detect biofilms which can exist as aggregates with a diameter of 10-100 µm diameter. If biofilm specific labels can be found, radiolabeling has the potential to increase sensitivity in signal detection and would also offer theranostic potential, in which localized imaging is combined with a therapeutic benefit [42].

**Unconventional methods**

**Bio-impedance based sensing.** Biofilm formation on an implant and associated inflammation can significantly alter the environment near the device surface and device-host tissue interface, which provides an opportunity for biofilm-specific sensing. One approach to characterize such changes is to use electric impedance, which represents the retardation of a circuit under alternating current voltage applied to the electric current at specific frequencies [43]. Previous improvements in instrumentation and data analysis [44,45] have made EIS a promising technology for biofilm detection.

Recently, a CMOS on-chip impedance analysis system was engineered using lithography. These devices are highly sensitive, compact, and can be reproduced at a large scale by modern fabrication processes [46]. *S*ingle-frequency impedance spectroscopy has been developed as a label-free system for monitoring biofilm growth and treatment [43,47]. EIS could become a promising monitoring method in the future since it does not require sample preparation and can be done continuously without physically sampling the biofilm, a major advantage over many other methods.

Recent advances have also made it possible to integrate impedimetric sensors with existing electronic medical devices. The most critical addition in such integration is the electrode/probe, which has been demonstrated in the form of a microfluidic chip [48]. This device combines real-time monitoring and a threshold activated treatment. It was later adapted by the same group into a wireless monitoring enabled urinary catheter [49]. It is worth noting that synergy between antibiotics and electrochemical treatment has been reported to have significant effect on killing biofilms including persister cells [50-52]. With the capability to combine treatment and monitoring, it is possible to engineer smart medical devices that can deliver on-demand control of biofilms.

**Surface Acoustic Waves.** Surface Acoustic Waves (SAW) are referred to a form of vibrational wave propagating through a solid material affected by its elasticity. Combing with Love waves (LW), which travel across a surface, LW-SAW devices can provide label-free real-time monitoring of high-molecular weight molecules with high sensitivity [53].

LW-SAW sensors have been developed for early-detection of biofilm formation on surfaces with pg level sensitivity and real-time biofilm monitoring [54]. Using gold nanoparticles as a signal enhancer, LW-SAW can also be used to detect antigen at pg/mL level [55]. Such high sensitivity might be useful for detecting biofilm markers in the future.

A new device based on SAW was engineered recently, which can be clamped onto a liquid filled tube and act as an actuator to monitor soft layer deposition on the tube surface. This device is capable of distinguishing between a soft layer such as biofilm or a hard layer such as limescale in a metal tube [56]. This adaptable device has possible applications in piping inspection and medical catheter check-ups, demonstrating an alternative to integrating into devices. Instead of putting SAW sensor onto the surface of interest, it rather turns the said surface into a substrate for analysis. Besides sensing, SAW has also been used in biofilm treatment [57], presenting a future possibility to engineer smart devices.

**Other approaches.** With advancements in analytic biotechnology and device fabrication, researchers continuously push the boundary towards more sensitive and portable detection systems. For example, a disposable sensor consisting of a small segment of optical fiber and antibody surface coating developed for C-reactive protein sensing has a consistent linear response to the target protein between 0.01 – 20 µg/mL [58]. Meanwhile, Zhang et al. [59] reported the capability to monitor biofilm formation and treatment in real-time using surface plasmon resonance waveguide mode. Cantilever technology can be used to form an array sensor, and achieve simultaneous detection of multiple targets of interest [60]. Although some of these are not directly applicable to biofilm monitoring *in vivo*, they have potential for biomarker detection, analysis of explanted medical devices and other applications, such as detection of biofilms in bioreactors.

**Conclusions**

With the increasing challenges associated with microbial biofilms, there is an urgent need to develop the capability for non-destructive real-time detection of microbial biofilms. However, the patchy nature, size scale and lack of biofilm specific targets provide obstacles in meeting this goal. Conventional methods for microbial detection are largely based on culturing methods and the detection of immunoglobulin antibodies in the blood or other body fluids. Although these methods are effective in diagnosing acute infections, they commonly fail in diagnosing BAIs. Similar challenges exist in detecting microbes from explanted medical devices. Many of the future development will rely on the discovery of new biomarkers and engineering more sensitive detection systems. Most BAIs are culturing negative and involve multiple species, which require universal biomarkers rather than species-specific molecules for detection. It is also important to improve the knowledge of *in vivo* biofilm formation, which has significant differences than that in *in vitro* pure-culture systems. Understanding the dynamics and inter-species interactions will be essential for identifying the right biomarkers. Achieving effective biofilm detection also requires low-cost, easy-to-manufacture, portable/wearable/implantable devices that have a long-life span and require minimal maintenance. Addressing these challenges will bring exciting new technologies for safer medical devices and better healthcare.

**Acknowledgements**

The authors thank the U.S. National Science Foundation for an ERC planning grant (1936926), which facilitates discussions of future research on microbial control and partially led to this review article.

**References**

1. Singh S, Singh SK, Chowdhury I, Singh R: **Understanding the Mechanism of Bacterial Biofilms Resistance to Antimicrobial Agents**. *Open Microbiol J* 2017, **11**:53-62.

2. Lebeaux D, Ghigo JM, Beloin C: **Biofilm-related infections: bridging the gap between clinical management and fundamental aspects of recalcitrance toward antibiotics**. *Microbiol Mol Biol Rev* 2014, **78**:510-543.

3. Donlan RM: **Biofilms: microbial life on surfaces**. *Emerg Infect Dis* 2002, **8**:881-890.

4. Pawlowski KS, Wawro D, Roland PS: **Bacterial biofilm formation on a human cochlear implant**. *Otol Neurotol* 2005, **26**:972-975.

5. Zimmerli W, Sendi P: **Orthopaedic biofilm infections**. *APMIS* 2017, **125**:353-364.

6. Veerachamy S, Yarlagadda T, Manivasagam G, Yarlagadda PK: **Bacterial adherence and biofilm formation on medical implants: a review**. *Proc Inst Mech Eng H* 2014, **228**:1083-1099.

\*7. Azeredo J, Azevedo NF, Briandet R, Cerca N, Coenye T, Costa AR, Desvaux M, Di Bonaventura G, Hébraud M, Jaglic Z, et al.: **Critical review on biofilm methods**. *Crit Rev Microbiol* 2017, **43**:313-351.

Author reviewed the traditional and cutting-edge methods for studying biofilms, from the measurement of biomass, viability, and composition, to the imaging and monitoring of cellular-level behaviors, covering a wide range of techniques for biofilm research.

8. Høiby N, Bjarnsholt T, Moser C, Bassi GL, Coenye T, Donelli G, Hall-Stoodley L, Holá V, Imbert C, Kirketerp-Møller K, et al.: **ESCMID guideline for the diagnosis and treatment of biofilm infections 2014**. *Clin Microbiol Infect* 2015, **21 Suppl 1**:S1-25.

9. Hall-Stoodley L, Stoodley P, Kathju S, Høiby N, Moser C, Costerton JW, Moter A, Bjarnsholt T: **Towards diagnostic guidelines for biofilm-associated infections**. *FEMS Immunol Med Microbiol* 2012, **65**:127-145.

10. Li C, Renz N, Trampuz A: **Management of Periprosthetic Joint Infection**. *Hip Pelvis* 2018, **30**:138-146.

11. Donlan RM, Costerton JW: **Biofilms: survival mechanisms of clinically relevant microorganisms**. *Clin Microbiol Rev* 2002, **15**:167-193.

12. Fisher RA, Gollan B, Helaine S: **Persistent bacterial infections and persister cells**. *Nat Rev Microbiol* 2017, **15**:453-464.

13. Shen Y, Stojicic S, Haapasalo M: **Bacterial viability in starved and revitalized biofilms: comparison of viability staining and direct culture**. *J Endod* 2010, **36**:1820-1823.

14. Bjerkan G, Witsø E, Bergh K: **Sonication is superior to scraping for retrieval of bacteria in biofilm on titanium and steel surfaces in vitro**. *Acta Orthop* 2009, **80**:245-250.

15. Liu Y, Zhang J, Ji Y: **PCR-based Approaches for the Detection of Clinical Methicillin-resistant *Staphylococcus aureus***. *Open Microbiol J* 2016, **10**:45-56.

\*\*16. Tsalik EL, Bonomo RA, Fowler VG: **New Molecular Diagnostic Approaches to Bacterial Infections and Antibacterial Resistance**. *Annu Rev Med* 2018, **69**:379-394.

This paper reviews new methods for molecular diagnostics, both host-focused and pathogen-focused approahes. The review covers the technologies in development and those that have been commercialized.

\*17. Quainoo S, Coolen JPM, van Hijum SAFT, Huynen MA, Melchers WJG, van Schaik W, Wertheim HFL: **Whole-Genome Sequencing of Bacterial Pathogens: the Future of Nosocomial Outbreak Analysis**. *Clin Microbiol Rev* 2017, **30**:1015-1063.

This is a review of conventional molecular approaches for characterizing bacterial strains, and various sequencing technologies that can be intergrated with bioinformatic tools to achieve whole-genome sequencing based nosocomial outbreak analysis.

\*18. Deirmengian C, Kardos K, Kilmartin P, Gulati S, Citrano P, Booth RE: **The Alpha-defensin Test for Periprosthetic Joint Infection Responds to a Wide Spectrum of Organisms**. *Clin Orthop Relat Res* 2015, **473**:2229-2235.

19. Lasa I, Penadés JR: **Bap: a family of surface proteins involved in biofilm formation**. *Res Microbiol* 2006, **157**:99-107.

20. Cucarella C, Tormo MA, Ubeda C, Trotonda MP, Monzón M, Peris C, Amorena B, Lasa I, Penadés JR: **Role of biofilm-associated protein bap in the pathogenesis of bovine *Staphylococcus aureus***. *Infect Immun* 2004, **72**:2177-2185.

21. Noori E, Rasooli I, Owlia P, Mousavi Gargari SL, Ebrahimizadeh W: **A conserved region from biofilm associated protein as a biomarker for detection of Acinetobacter baumannii**. *Microb Pathog* 2014, **77**:84-88.

22. Yousefi M, Pourmand MR, Fallah F, Hashemi A, Mashhadi R, Nazari-Alam A: **Characterization of *Staphylococcus aureus* Biofilm Formation in Urinary Tract Infection**. *Iran J Public Health* 2016, **45**:485-493.

\*23. Antypas H, Choong FX, Libberton B, Brauner A, Richter-Dahlfors A: **Rapid diagnostic assay for detection of cellulose in urine as biomarker for biofilm-related urinary tract infections**. *NPJ Biofilms Microbiomes* 2018, **4**:26.

This study describes a new method for rapid detection of uropathogenic *Escherichia coli* using cellulose as a biomaker of biofilms. The detection was achieved by using optotracing.

24. Pan Y, Fisher T, Olk C, Inzana TJ: **Detection of antibodies to the biofilm exopolysaccharide of Histophilus somni following infection in cattle by enzyme-linked immunosorbent assay**. *Clin Vaccine Immunol* 2014, **21**:1463-1467.

\*25. Høiby N, Bjarnsholt T, Moser C, Jensen P, Kolpen M, Qvist T, Aanaes K, Pressler T, Skov M, Ciofu O: **Diagnosis of biofilm infections in cystic fibrosis patients**. *APMIS* 2017, **125**:339-343.

This article summarizes the current methods for the diagnosis of bacterial infections in cystic fibrosis patients, with *P. aeruginosa* infections as a focus.

26. Nagler M, Insam H, Pietramellara G, Ascher-Jenull J: **Extracellular DNA in natural environments: features, relevance and applications**. *Appl Microbiol Biotechnol* 2018, **102**:6343-6356.

27. Castillo Pedraza MC, Novais TF, Faustoferri RC, Quivey RG, Terekhov A, Hamaker BR, Klein MI: **Extracellular DNA and lipoteichoic acids interact with exopolysaccharides in the extracellular matrix of *Streptococcus mutans* biofilms**. *Biofouling* 2017, **33**:722-740.

\*\*28. Guan A, Wang Y, Phillips KS: **An extraction free modified o-phthalaldehyde assay for quantifying residual protein and microbial biofilms on surfaces**. *Biofouling* 2018, **34**:925-934.

This article describes a modified o-phthalaldehyde (OPA) protein assay, which can quantify proteins in surface-attached biofilms without extraction. This has potential applications in biofilm detection in hard-to-reach areas such as those in endoscopes.

29. Li YH, Tian X: **Quorum sensing and bacterial social interactions in biofilms**. *Sensors (Basel)* 2012, **12**:2519-2538.

30. Singh PK, Schaefer AL, Parsek MR, Moninger TO, Welsh MJ, Greenberg EP: **Quorum-sensing signals indicate that cystic fibrosis lungs are infected with bacterial biofilms**. *Nature* 2000, **407**:762-764.

31. Skindersoe ME, Zeuthen LH, Brix S, Fink LN, Lazenby J, Whittall C, Williams P, Diggle SP, Froekiaer H, Cooley M, et al.: ***Pseudomonas aeruginosa* quorum-sensing signal molecules interfere with dendritic cell-induced T-cell proliferation**. *FEMS Immunol Med Microbiol* 2009, **55**:335-345.

32. Jensen P, Givskov M, Bjarnsholt T, Moser C: **The immune system vs. *Pseudomonas aeruginosa* biofilms**. *FEMS Immunol Med Microbiol* 2010, **59**:292-305.

\*33. Moser C, Pedersen HT, Lerche CJ, Kolpen M, Line L, Thomsen K, Høiby N, Jensen P: **Biofilms and host response - helpful or harmful**. *APMIS* 2017, **125**:320-338.

The article discusses various types of immune response and their interaction with biofilm infection, both acute and chronic; and the possibility of diagnosing biofilm infections by monitoring acquired immue response.

34. Azevedo NF, Vieira MJ, Keevil CW: **Establishment of a continuous model system to study *Helicobacter pylori* survival in potable water biofilms**. *Water Sci Technol* 2003, **47**:155-160.

35. You Y, Moreira BG, Behlke MA, Owczarzy R: **Design of LNA probes that improve mismatch discrimination**. *Nucleic Acids Res* 2006, **34**:e60.

\*36. Nguyen CQ, Thrift WJ, Bhattacharjee A, Ranjbar S, Gallagher T, Darvishzadeh-Varcheie M, Sanderson RN, Capolino F, Whiteson K, Baldi P, et al.: **Longitudinal Monitoring of Biofilm Formation via Robust Surface-Enhanced Raman Scattering Quantification of *Pseudomonas aeruginosa*-Produced Metabolites**. *ACS Appl Mater Interfaces* 2018, **10**:12364-12373.

This study reports a micorfluidic SERS device for quantitative detection of quorum sensing in *P. pseudomonas* biofilms, with a capability to detect 1 ng/mL of pyocyanin.

37. Wu X, Chen J, Li X, Zhao Y, Zughaier SM: **Culture-free diagnostics of *Pseudomonas aeruginosa* infection by silver nanorod array based SERS from clinical sputum samples**. *Nanomedicine: Nanotechnology, Biology and Medicine* 2014, **10**:1863-1870.

38. Bodelón G, Montes-García V, López-Puente V, Hill EH, Hamon C, Sanz-Ortiz MN, Rodal-Cedeira S, Costas C, Celiksoy S, Pérez-Juste I, et al.: **Detection and imaging of quorum sensing in *Pseudomonas aeruginosa* biofilm communities by surface-enhanced resonance Raman scattering**. *Nature Materials* 2016, **15**:1203.

\*\*39. Carrel M, Beltran MA, Morales VL, Derlon N, Morgenroth E, Kaufmann R, Holzner M: **Biofilm imaging in porous media by laboratory X-Ray tomography: Combining a non-destructive contrast agent with propagation-based phase-contrast imaging tools**. *PLOS ONE* 2017, **12**:e0180374.

Authors reported a new contrast agent (iron sulfate) for X-Ray tomography that is non-reactive compared to conventional contrast agents. This new contrast agent is integrated into the extracellular matrix during biofilm growth, enabling biofilm imaging in opaque porous media.

40. Niehaus WL, Howlin RP, Johnston DA, Bull DJ, Jones GL, Calton E, Mavrogordato MN, Clarke SC, Thurner PJ, Faust SN, et al.: **Development of X-ray micro-focus computed tomography to image and quantify biofilms in central venous catheter models in vitro**. 2016, **162**:1629-1640.

41. Sellmyer MA, Lee I, Hou C, Weng CC, Li S, Lieberman BP, Zeng C, Mankoff DA, Mach RH: **Bacterial infection imaging with [***Proc Natl Acad Sci U S A* 2017, **114**:8372-8377.

42. Levine R, Krenning EP: **Clinical History of the Theranostic Radionuclide Approach to Neuroendocrine Tumors and Other Types of Cancer: Historical Review Based on an Interview of Eric P. Krenning by Rachel Levine**. *J Nucl Med* 2017, **58**:3S-9S.

43. van Duuren JBJH, Müsken M, Karge B, Tomasch J, Wittmann C, Häussler S, Brönstrup M: **Use of Single-Frequency Impedance Spectroscopy to Characterize the Growth Dynamics of Biofilm Formation in *Pseudomonas aeruginosa***. *Scientific Reports* 2017, **7**:5223.

44. Liu L, Xu Y, Cui F, Xia Y, Chen L, Mou X, Lv J: **Monitoring of bacteria biofilms forming process by in-situ impedimetric biosensor chip**. *Biosensors and Bioelectronics* 2018, **112**:86-92.

45. Furst AL, Francis MB: **Impedance-Based Detection of Bacteria**. *Chem Rev* 2019, **119**:700-726.

46. Yu H, Yan, M. and Huang, X: **CMOS Impedance Sensor**. *CMOS Integrated Lab‐on‐a‐Chip System for Personalized Biomedical Diagnosis* 2018:60-75.

47. Gutiérrez D, Hidalgo-Cantabrana C, Rodríguez A, García P, Ruas-Madiedo P: **Monitoring in Real Time the Formation and Removal of Biofilms from Clinical Related Pathogens Using an Impedance-Based Technology**. *PLOS ONE* 2016, **11**:e0163966.

\*\*48. Subramanian S, Tolstaya EI, Winkler TE, Bentley WE, Ghodssi R: **An Integrated Microsystem for Real-Time Detection and Threshold-Activated Treatment of Bacterial Biofilms**. *ACS Applied Materials & Interfaces* 2017, **9**:31362-31371.

This study combined impedance-based detection and bioelectric-based treatment of biofilm in one integrated chip targeting *E. coli* K-12 biofilms.

49. Huiszoon RC, Subramanian S, Rajasekaran PR, Beardslee LA, Bentley WE, Ghodssi R: **Flexible Platform for In Situ Impedimetric Detection and Bioelectric Effect Treatment of *Escherichia Coli* Biofilms**. *IEEE Transactions on Biomedical Engineering* 2019, **66**:1337-1345.

50. Freebairn D, Linton D, Harkin-Jones E, Jones DS, Gilmore BF, Gorman SP: **Electrical methods of controlling bacterial adhesion and biofilm on device surfaces**. *Expert Rev Med Devices* 2013, **10**:85-103.

51. Niepa TH, Gilbert JL, Ren D: **Controlling *Pseudomonas aeruginosa* persister cells by weak electrochemical currents and synergistic effects with tobramycin**. *Biomaterials* 2012, **33**:7356-7365.

52. Sultana ST, Call DR, Beyenal H: **Eradication of *Pseudomonas aeruginosa* biofilms and persister cells using an electrochemical scaffold and enhanced antibiotic susceptibility**. *npj Biofilms and Microbiomes* 2016, **2**:2.

53. Puiu M, Gurban A-M, Rotariu L, Brajnicov S, Viespe C, Bala C: **Enhanced sensitive love wave surface acoustic wave sensor designed for immunoassay formats**. *Sensors (Basel, Switzerland)* 2015, **15**:10511-10525.

54. Kim YW, Meyer MT, Berkovich A, Subramanian S, Iliadis AA, Bentley WE, Ghodssi R: **A surface acoustic wave biofilm sensor integrated with a treatment method based on the bioelectric effect**. *Sensors and Actuators A: Physical* 2016, **238**:140-149.

55. Li S, Wan Y, Su Y, Fan C, Bhethanabotla VR: **Gold nanoparticle-based low limit of detection Love wave biosensor for carcinoembryonic antigens**. *Biosensors and Bioelectronics* 2017, **95**:48-54.

56. Tietze S, Singer F, Lasota S, Ebert S, Landskron J, Schwuchow K, Drese KS, Lindner G: **Monitoring of Soft Deposition Layers in Liquid-Filled Tubes with Guided Acoustic Waves Excited by Clamp-on Transducers**. *Sensors* 2018, **18**.

57. Wang H, Teng F, Yang X, Guo X, Tu J, Zhang C, Zhang D: **Preventing microbial biofilms on catheter tubes using ultrasonic guided waves**. *Sci Rep* 2017, **7**:616.

58. Wang W, Mai Z, Chen Y, Wang J, Li L, Su Q, Li X, Hong X: **A label-free fiber optic SPR biosensor for specific detection of C-reactive protein**. *Scientific Reports* 2017, **7**:16904.

59. Zhang P, Guo J-S, Yan P, Chen Y-P, Wang W, Dai Y-Z, Fang F, Wang G-X, Shen Y: **Dynamic Dispersal of Surface Layer Biofilm Induced by Nanosized TiO&lt;sub&gt;2&lt;/sub&gt; Based on Surface Plasmon Resonance and Waveguide**. *Applied and Environmental Microbiology* 2018, **84**:e00047-00018.

60. Zheng F, Wang P, Du Q, Chen Y, Liu N: **Simultaneous and Ultrasensitive Detection of Foodborne Bacteria by Gold Nanoparticles-Amplified Microcantilever Array Biosensor**. *Frontiers in Chemistry* 2019, **7**:232.