Impact of a routine molecular point-of-care ‘test and treat’ strategy for influenza in hospitalised adults: a multi-centre, randomised controlled trial (FluPOC)

Tristan W Clark1,2,3,4*, Kate R Beard1, Nathan J Brendish1, 2, Ahalya K Malachira2, Samuel Mills2, Cathleen Chan2, Stephen Poole1,2,3, Sean Ewings5, Nick Cortes6, Esther Nyimbili3 and Laura Presland3

1. School of Clinical and Experimental Sciences, Faculty of Medicine, University of Southampton, Southampton, UK

2. Department of Infection, University Hospital Southampton NHS Foundation Trust, Southampton, UK.

3. NIHR Southampton Biomedical Research Centre, University Hospital Southampton NHS Foundation Trust, Southampton, UK

4. NIHR Post Doctoral Fellowship Programme, UK

5. School of Primary Care and Population Sciences, University of Southampton, Southampton UK

6. Department of Microbiology, Hampshire Hospitals Foundation NHS Trust, Winchester, UK

*Corresponding author

Dr Tristan William Clark, Associate Professor and Consultant in Infectious Diseases.

NIHR Post-doctoral Fellow.

Address: LF101, South Academic block, Southampton General Hospital, Southampton, SO16 6YD.

Tel: 0044(0)23818410	E-mail address: T.W.Clark@soton.ac.uk

Summary

Background The diagnosis of Influenza in hospitalised patients is delayed due to long turnaround times with laboratory testing, leading to inappropriate and late antiviral and isolation facility use. Molecular point-of-care test (mPOCT) are highly accurate, easy to use and generate results in under 1 hour but high quality evidence for their clinical impact is lacking. The aim of this study was to evaluate the clinical impact of routine mPOCT on influenza detection, antiviral use, infection control measures and clinical outcome, in adults hospitalised with acute respiratory illness.

Methods In this multicentre, randomised controlled trial we enrolled adults hospitalised with acute respiratory illness during influenza seasons. Patients were randomised (1:1) to receive mPOCT for influenza or routine clinical care. The primary outcome was the proportion of influenza-infected patients who received antivirals. Secondary outcomes included time to antivirals, isolation facility use, and clinical outcomes. This study is registered with ISRCTN:17197293 and has completed.

Findings Between December 2017 and May 2019, 613 patients were enrolled (307 assigned to mPOCT and 306 to routine care) and all were analysed. 100 (33%) of 307 patients in the mPOCT group and 102 (33%) of 306 in the control group had influenza. 100 (100%) of 100 influenza-infected patients were diagnosed in the mPOCT group and 60 (59%) of 102 were diagnosed though routine clinical care (relative risk 1·7, 95%CI 1·7 to 1·7; p<0·0001). 99 (99%) of 100 influenza-infected patients received antivirals in the mPOCT group versus 63 (62%) 102 in the control group (relative risk 1·6, 95%CI 1·4 to 1·9; p<0·0001). Median time to antivirals was 1·0 hour in the mPOCT group versus 6·0 hours in the control group (difference of 5·0 hours, 95%CI 0 to 6·0; p=0·004). 70 (70%) of 100 influenza-infected patients in the mPOCT group were nursed in single room accommodation versus 39 (38%) of 102 in the control group (relative risk 1·8, 95%CI 1·4 to 2·4; p<0·0001).

Interpretation Routine mPOCT for influenza was associated with improved influenza detection and improvements in appropriate and timely antiviral and isolation facility use. Routine mPOCT should replace laboratory based diagnostics for acute admissions to hospital during influenza season.

Funding NIHR

Introduction

The influenza virus causes seasonal epidemics every winter in temperate climates leading to large numbers of hospitalisations, principally in adults.1-4 Adults hospitalised with influenza are frequently admitted to critical care units and 3-15% will die in hospital.5 Influenza often remains undiagnosed in patients admitted to hospital due to a lack of routine testing and the inaccuracy of testing based on clinical suspicion.6 Many influenza-infected patients are therefore not identified and treated appropriately with antiviral agents or isolated correctly in single room accommodation, to prevent the spread of influenza to others. Even when patients are tested for influenza, the diagnosis is delayed due to the long turnaround times of centralised laboratory polymerase chain reaction (PCR) testing.7 This delay also contributes to the inappropriate and late use of antivirals use and isolation facilities and impairs patient flow though acute areas reducing hospital operational capacity.

Antiviral treatment with neuraminidase inhibitors (NAIs) is recommended for all hospitalised patients with suspected and proven influenza by UK and US national guidelines8, 9 and evidence suggests that earlier use is associated with better outcomes.10, 11 New rapid molecular point-of-care tests (mPOCT) for influenza are highly accurate, easy to use and can generate a result in under 1 hour.12-15 Routinely testing patient for influenza at the point-of-care may improve influenza detection, the use of antivirals and isolation facilities by allowing confident, real-time decision making.7, 16-20

Our previous trial suggested that the use of routine molecular point-of-care testing for respiratory viruses improves the use of antivirals and isolation facilities for influenza, but was hampered by small numbers of influenza positive patients and a lack of universal viral testing in the control group.7 Several retrospective observational studies have also suggested that mPOCT may improve the detection of influenza, the use of influenza antivirals and isolation facilities and may reduce nosocomial transmission15-19, however these studies are highly prone to bias. As mPOCT for influenza is now starting to be used in some NHS hospitals, high quality evidence for its clinical impact is urgently needed.  The aim of the study was to prospectively evaluate the impact of a routine molecular point-of-care ‘test and treat’ strategy for influenza in adults hospitalised with acute respiratory illness, on clinical management and patient outcome.

Methods

Study design and participants

We performed a multicentre, pragmatic, parallel group, open-label, randomised controlled superiority trial with 1:1 allocation to intervention and control groups, in two acute hospitals in Hampshire, UK.  The trial took place over two successive winter seasons in 2017/18 and 2018/19 when influenza was circulating according to national (Public Health England) surveillance systems. All patients were recruited from the Acute Medical Unit (AMU) and Emergency Department (ED) of Southampton General Hospital (a large acute teaching hospital in the South of the UK serving a population of 650,000 for secondary care) run by University Hospital Southampton Foundation NHS Trust (who were the sponsor for the trial) and from Hampshire County Hospital, a large district general hospital in Hampshire run by Hampshire Hospitals Foundation NHS Trust. The study was approved by the South Central - Hampshire A Research Ethics Committee: REC reference 17/SC/0368, approved on the 7th September 2017. The protocol is published and freely available at:

https://bmjopen.bmj.com/content/9/12/e031674.long.20 The protocol was amended a single time on the 23rd November 2017 to allow the inclusion of a second study site (Hampshire Hospitals Foundation NHS Trust) and this was communicated to the trial registry.

Eligible patients were: aged 18 years or over; had the capacity to give informed, written consent or where capacity was lacking consultee assent was obtained; were a patient admitted to the hospital and physically in either the AMU or ED; could be recruited to the study within a 16 hours of admission; had an acute respiratory illness; and had a duration of illness less than or equal to 10 days prior to hospitalisation. An episode of acute respiratory illness was defined as a provisional diagnosis of  acute pulmonary illness including pneumonia, bronchitis (non-pneumonic lower respiratory tract infection - NPLRTI) and influenza-like illness (ILI), or an acute exacerbation of a chronic respiratory illness (including exacerbation of COPD, asthma or bronchiectasis). The exclusion criteria were: patients not fulfilling inclusion criteria; a palliative approach being taken by the treating clinicians; patient declines nasal/pharyngeal swabbing; previously included in this study and re-presenting within the last 30 days after hospital discharge; or consent declined or consultee assent declined.

Randomisation and masking

Patient were consecutively assigned a unique participant identification number by study team members who then used an internet-based randomisation service (sealedenvelope.com, which uses random permuted blocks of varying sizes; 4, 6 and 8) to generate the allocation sequence, and assigned the participants to the groups in a 1:1 ratio. Patients assigned to the intervention group were immediately swabbed and tested as detailed below. No attempt at blinding trial participants, research staff or clinical care providers was made. Data analysts and statisticians were blinded to group allocation.

Procedures

Participants randomised to the intervention (mPOCT) group had a combined nose and throat swab sample (and sputum samples where available) taken by research staff according to standard protocols and then analysed immediately using the FilmArray Respiratory Panel 2. The testing units were located in the AMUs and the test took around 45 minutes to generate a result. The clinical and infection control teams were informed directly of all results (positive and negative). Patients randomised to the control group were managed according to routine clinical care where testing for respiratory viruses was at the discretion of the responsible clinical team and was performed using laboratory PCR by conventional methods, in the on-site laboratory facilities. Patients in the control group also had a nose and throat swab obtained at enrolment and stored for subsequent analysis (using the FilmArray Respiratory Panel 2) to allow subsequent assessment of missed diagnoses, at least 30 days post enrolment.  Clinical management decisions were made independently by the responsible clinical team and where influenza was detected by mPOCT, research teams signposted national guidelines for the treatment of influenza [9].  Demographic and clinical data were collected at recruitment and outcome data was collected retrospectively from paper case-notes, electronic medical records and electronic prescribing systems. All data were collected on a standardised case report form.

Outcomes

The primary outcome measure was the proportion of influenza-infected patients treated with antivirals (NAIs) during their hospital stay, within 5 days of admission. This was selected as Public Health England guidelines recommend that all adults requiring hospital admission for influenza should receive NAIs. This guidance is consistent with those of several other countries.8, 9 Although no placebo controlled trials of NAIs have ever been performed in hospitalized adults, there is a large body of observational evidence demonstrating that NAI use is associated with improvements in clinical outcome in this group. 10, 11, 26-28

Secondary outcomes were: The proportion of influenza-infected (PCR positive) patients identified during hospitalisation, turnaround time for respiratory virus testing, proportion of all antiviral use occurring in influenza-infected patients, time to antiviral commencement, duration of antiviral use in influenza-infected patients, duration of antiviral use in influenza uninfected  patients, proportion of patients treated with antibiotics, proportion of patients treated with single doses or brief courses (less than 24 hours) of antibiotics, duration of antibiotic use,  proportion of patients isolated in a side room, duration of isolation facility use, proportion of influenza-infected patients isolated, time to isolation for influenza-infected patients, time from admission to de-isolation of influenza negative patients, duration of hospitalisation, time on supplementary oxygen, time to clinical improvement*, median hospital recovery score (HRS) at day 4 and 7**, proportion of patients requiring intensive care unit (ICU) or high dependency unit (HDU) admission, proportion re-presenting to hospital within 30 days , proportion readmitted to hospital within 30 days, in-hospital and 30 day mortality.

All outcomes were measured for the duration of hospitalisation or up to 30 days (whichever was shortest) unless specified otherwise and the duration of antimicrobials includes medication (antibiotics and antivirals) that patients were discharged home with.

*The outcome measure time to clinical stability (definition in the published protocol) was subsequently changed to time to clinical improvement (TCCI), a recently developed outcome measure for trials of influenza antivirals. Full details and references for TCCI are provided in the appendix p2.

** Hospital Recovery Scale (HRS) is a 6 point ordinal scale developed as an outcome measure for trials of influenza therapeutics. Full details and references for HRS are provided in the appendix p2.

Statistical analysis

Sample size was based upon the primary outcome measure of the proportion of influenza positive patients treated with NAIs (within 5 days of hospital admission). The study initially aimed to recruit a total of up to 840 patient-participants (420 patients per group) over up to 3 influenza seasons, on the conservative assumption of a 25% positivity rate for influenza during influenza seasons (based on our previous locally conducted studies over several seasons), i.e. ~100 influenza positive patients in each group. These group sizes would give a 90% power at a 0.05 significance level to detect a difference of 20% in neuraminidase inhibitor use for influenza-infected patients (from 65% to 85%) [8]. As 100 patients or greater with confirmed influenza were recruited to each group by the end of the second season the trial was stopped at this point having recruited 613 patients in total.

Analysis was by intention-to-treat and the frame work was superiority. Analysis was performed by a dedicated statistician from the University of Southampton who was blinded to group allocation. Analysis was conducted using Prism version 7·0 (GraphPad Software Inc; La Jolla, California) and Stata version 16·0 (StataCorp, College Station, Texas). Data for the primary outcome was complete and missing data was minimal (<1%) overall. The primary outcome (proportion of influenza infected patients treated with NAIs) was measured only in influenza-infected patients, i.e. the intention-to-treat influenza-infected (ITTI) population. Difference in proportions and unadjusted risk ratios (relative risk) were used to compare the groups. The pre-specified plan was to analyse antiviral use in a logistic regression model; however, the model could not be fitted due to separation, at least partly as a result of only one person in the mPOCT not receiving antivirals. No further multivariable analyses were attempted given the strong signal. Secondary outcome measures were measured both in all patients (ITT population) and also in influenza-infected (ie the ITTI population). For secondary outcomes, the intervention and control groups were compared for equality of proportions for binary data and t-tests or Mann-Whitney U tests for continuous data (e.g. turnaround time), as appropriate; the choice between the latter two was based on the distribution of the observed data and the sample size. Where 95% confidence intervals are presented, Stata version 16.0 defaults are used.

This study was prospectively registered with the ISRCTN 17197293, on the 13th November 2017.

Role of funding source

This study was funded by the NIHR Post Doctoral Fellowship Programme and was supported by the NIHR Southampton Respiratory Biomedical Research Centre and the NIHR Clinical Research Network, Wessex. The molecular analysers were leased and the test kits purchased independently from the UK distributer bioMérieux UK ltd. The manufactures of the molecular test platform (BioFire diagnostics, a bioMérieux company, Salt Lake City, Utah, US) had no role in the study conception, design, data analysis or manuscript preparation. The corresponding author had full access to all of the data and the final responsibility to submit for publication.

Results

Between 12th December 2017 and 3rd May 2019, 613 patients were randomised during 2 influenza seasons (2017/18 and 2018/19), 307 to the mPOCT group and 306 to the control group. As 100 patients or greater with confirmed influenza were recruited to each group by the end of the second season, the trial was stopped at this point, having recruited 613 patients in total. All patients were analysed in the intention-to-treat (ITT) analysis and all influenza-infected patients were analysed in the intention-to-treat infected (ITTI) analysis.

All patients (ITT)

The median (IQR) turnaround time for results was 1·2 (1·1 to 1·4) hours in the mPOCT group and 22.6 [16·0 to 28·7] hours in the control group (difference of 21·4 hours, 95%CI 20·6 to 22·4; p<0·0001; table2). All antiviral use was with the NAI oseltamivir and there was no difference in the proportion of patients receiving antivirals between the groups; 139 (45%) of 307 in the mPOCT group versus 130 (43%) 306 in the control group (relative risk 1·1 , 95% CI 0·9 to 1·3; p=0·49; table 2) however, more antiviral-treated patients in the mPOCT  group were influenza-infected; 99 (71%) of 139 versus 63 (49%) of 130 in the control group (relative risk 1·5, 95%CI 1·2 to 1·8; p<0·0001; table 2). Influenza-uninfected patients receiving antivirals had a median (IQR) course length of 0·1 (0 to 0·5) days (i.e. a single dose) in the mPOCT group versus 1·5 (0·5 to 4·5) days in the control group (difference of 1·4 days, 95% CI 0·7 to 2·1; p<0·0001; table 2). More patients in the mPOCT group; 106 (35%) of 307, were isolated in single room accommodation versus 67 (22%) of 306 in the control group (relative risk 1·6, 95%CI 1·2 to 2·1; p=0·0006; table 2).

Influenza infected patients (ITTI)

100 (33%) of 307 patients in the mPOCT group and 102 (33%) of 306 patients in the control group were influenza infected. 100 (100%) of 100 influenza-infected patients were diagnosed with influenza whilst in hospital in the mPOCT group (all by PCR testing) but only 60 (59%) of 102 were diagnosed in the control group, 55 were PCR positive and 5 had a clinical diagnosis of influenza but had no confirmatory test performed (relative risk 1·7, 95%CI 1·7 to 1·7; p<0·0001; table 3).

For the primary outcome, 99 (99%) of 100 influenza-infected patients received antiviral treatment in the mPOCT group versus 63 (62%) 102 in the control group (relative risk 1·6, 95%CI 1·4 to 1·9; p<0·0001; number needed to test = 2·7; table 3). Median (IQR) time to receipt of antivirals was 1·0 (0 to 2·0) hour in the mPOCT group versus 6·0 (0 to 12·0) hours in the control group (difference of 5·0 hours, 95%CI 0 to 6·0; p=0·004; table 3 and figure 2).  Excluding those patients treated empirically (i.e. before results were available) the median (IQR) time to receipt of antivirals was 2·0 [1·0 to 3·0] hours in the mPOCT group versus 30·5 [12·0 to 59·0] hours in the control group (difference of 28·5 hours, 95%CI 18·0 to 31·0; p<0·0001; table 3 and appendix p6)

Very few influenza-infected patients were correctly isolated in single room accommodation empirically (i.e. prior to the result of influenza testing) in either group; 10 (10%) of 100 and 9 (9%) of 102 (relative risk 1·1, 95%CI 0·5 to 2·7; p=0·77; table 3).  During hospitalisation, 70 (70%) of 100 influenza-infected patients in the mPOCT group were correctly nursed in single room accommodation versus 39 (38%) of 102 in the control group (relative risk 1·8, 95%CI 1·4 to 2·4; p<0·0001; table 3). Excluding those patients rapidly discharged on the same day as admitted or nursed in an area without single room accommodation (such as high dependency areas), 70 (92%) of 76 influenza-infected patients in the mPOCT group were correctly nursed in single room accommodation versus 39 (43%) of 90 in the control group (relative risk 2·1, 95%CI 1·7 to 2·8; p<0·0001). Patients were isolated after a median (IQR) of 4·5 (1·9 to 9·8) hours in the mPOCT group verses 25·0 (3·6 to 45·0) hours in the control group, (difference of 20·0 hours, 95%CI 6·0 to 26·0; p=0·0003; table 3 and figure 3). Influenza-infected patients not initially isolated, were isolated after a median (IQR) of 5·0 (2·0 to 12·0) hours in the mPOCT group verses 32·0 (16·0 to 55·0) hours in the control group (difference of 27·0 hours, 95%CI 17·0 to 33·0; p<0·0001; table 3, appendix p6).

90 (90%) of 100 influenza-infected patients received antibiotics in the mPOCT group versus 98 (96%) of 102 in the control group (relative risk 0·9, 95%CI 0·9 to 1·0; p=0·089; table 3). Patients in the mPOCT group received antibiotics for a median (IQR) of 6·1 (4·1 to 7·2) days versus 5·8 (4·2 to 7·4) days in the control group (difference -0·3, 95%CI -0·8 to 0·8; p=0·96; table 3). 13 (14%) of 90 patients in the mPOCT group received antibiotics for less than 24 hours versus 7 (7%) of 98 in the control group (relative risk 0·9, 95%CI 0·8 to 1·1; p=0·10; table 3).

Influenza-infected patients were hospitalised for a median (IQR) of 2·7 (1·1 to 4·2) days in the mPOCT group versus 2·7 (1·1 to 5·4) days in the control group (difference of 0·0 days, 95%CI -0·4 to 1·0; p=0·37; table 3). The hospital recovery scale (HRS) is an ordinal categorical scale used to assess patient outcome in interventional influenza trials, at various time points (full details in appendix p2) with a higher score denoting a worse outcome. Median HRS score at 7 days post admission was lower in the mPOCT group verses the control group (p=0·045, appendix p7). Median (IQR) time to clinical improvement (TTCI, defined in the appendix p2) was 8·8 hours in the mPOCT group versus 13·9 hours in the control group (difference of 5·1 hours, 95%CI -0·3 to 9·9; p=0·078; table 3). Time to event analysis showed a trend towards reduced TTCI in the mPOCT group compared to the control group (Log-rank test p=0·056; appendix p8). Influenza-infected patient were treated with a median (IQR) of 32·0 (15·0 to 59·0) hours of supplementary oxygen versus 33·0 (14·0 to 81.0) hours in the control group (difference of 1·0 hour, 95%CI -8·0 to 18·0; p=0·48; table 3 and appendix p9).

19 adverse events occurred in the influenza-infected patients in the mPOCT group compared to 34 in the control group (table 4). 1 (1%) of 100 patients was admitted to a critical care unit in the mPOCT group compared with 6 (6%) of 102 in the control group (relative risk 0·2, 95%CI 0·1 to 1·0; p=0·058; table 4). No patients died in the mPOCT group and 2 (2%) of 102 died in the control group (p=0·16). 10 (10%) of 100 patients had prolonged hospitalisation (greater than 7 days) in the mPOCT group versus 20 (20%) of 102 in the control group (relative risk 0·5, 95%CI 0·3 to 1·0; p=0·055; table 4).

Discussion

The routine use of a molecular point-of-care ‘test and treat’ strategy for influenza during influenza season led to improved detection of influenza-infected patients admitted to hospital and the delivery of results in near real-time to clinical and infection control teams.  NAI use and isolation facility use were improved by mPOCT so that influenza-infected patients were rapidly treated with antivirals and appropriately isolated. In addition routine mPOCT was also associated with improvements in some clinical outcomes measures in influenza-infected patients, including a more rapid recovery.

The high proportion of influenza-infected patients remaining undiagnosed in the control group is likely to relate to a lack of testing by clinicians in patients not displaying the classical symptoms or signs of influenza.  Influenza testing in most institutions is not performed routinely but is instead based on clinical suspicion of influenza, which is known to be highly inaccurate, even during periods of intense transmission.6 Undiagnosed influenza-infected patients clearly represent a high risk for nosocomial transmission of influenza to staff and other patients and in addition undiagnosed patients do not have the opportunity to benefit from antiviral therapy. This makes a compelling argument for health care institution to institute routine molecular testing for influenza in all patients admitted to hospital with ARI, during influenza season.

Our study achieved rapid turnaround times for results of just over 1 hour compared with almost 1 day with laboratory PCR in the control group, a typical laboratory turnaround time reported in studies.7, 21 The importance of very rapid turnaround times in improving clinical management has been demonstrated in several studies22, 23 and is supported by similar findings in studies of other diagnostics in emergency departments.24 It is possible that some of the benefits associated with POCT in this study could also be achieved by routine laboratory PCR testing if turnaround times could be dramatically reduced. However as turnaround times of one hour or less are unlikely to be achievable in most laboratory settings, testing for influenza in hospitals should move to the point-of-care. The mPOCT platform used in this study generated result in 45 minutes but other mPOCT platforms can generate results in under 20 minutes and may further reduce time to therapeutic and infection control interventions.

The use of influenza antivirals was significantly improved with routine mPOCT so that nearly all influenza-infected patients received appropriate antiviral treatment within a few hours of admission. Patients not initially suspected of influenza who subsequently had a positive POCT were rapidly started on NAIs compared with a delay of over 24 hours in the control group. Conversely patients who were started empirically on NAIs because of suspicion of influenza, but who subsequently tested negative by mPOCT, had their NAIs stopped after a single dose compared with a prolonged course of unnecessary drug in the control group. These findings are consistent with earlier randomised controlled trials and observational studies of mPOCT for influenza which consistently show improved use of directed NAIs in patients tested by mPOCT.7, 16-19 With the use of even faster mPOCTs performed at the point of presentation, the need for empirical therapy could be removed altogether without compromising the speed of administration of antivirals and ushering in an era real-time pathogen-directed therapy for influenza and other infections.25

Hospital side rooms, used for isolating potentially infectious patients, are a limited resource in most UK hospitals. Our study shows that in routine clinical care most influenza-infected patients are not isolated appropriately in single room accommodation or in influenza cohort wards, representing a high risk for nosocomial transmission. In comparison most influenza-infected patients in the mPOCT group were correctly isolated. In addition, influenza-infected patients who were not initially isolated but subsequently had a positive laboratory PCR result for influenza were only isolated after a delay of nearly 2 days (reflecting the long turnaround time of laboratory PCR) compared with just a few hours in the mPOCT group. The rational and appropriate use of single rooms is hugely important for UK hospitals in order to prevent nosocomial transmission of infection and to maintain patient flow though acute areas during periods of intense influenza activity.

Routine mPOCT use was not associated with major reductions in antibiotic use in influenza-infected patients in this trial and the majority of patients in both groups received antibiotics, consistent with other similar trials.7, 26 Other studies of mPOCT for respiratory viruses have shown similar results with either no effect of antibiotic use or only modest reductions.7, 15-19 The identification of a virus by PCR does not by itself rule out the presence of a concomitant bacterial infection and so clinicians tend to continue to treat hospitalised ARI patients with antibiotics. Alternative strategies to reduce unnecessary antibiotics are still urgently needed in this group and include combining pathogen detection with biomarker measurement (such as CRP or procalcitonin) or with other markers of host response, all integrated within a robust antibiotic stewardship programmes.

In addition to demonstrating the above improvements in clinical management this is the first randomised trial to show improvements in clinical outcome in influenza-infected patients through the use of mPOCT. Although this study cannot directly attribute the cause of these changes, the improvements in time to clinical improvement and HRS are likely to relate to the rapid, near universal use of oseltamivir in influenza-infected patients in the mPOCT group compared to frequently absent or delayed use with routine care. Although no placebo controlled trials of NAIs have ever been performed in hospitalized adults, these findings are consistent with a large body of observational evidence showing that NAI use is associated with improvements in clinical outcome including length of stay and mortality, in this group10, 11, 27-29 and that earlier use of NAIs is associated with the best outcomes.10, 11

The strengths of the study include; the multicentre nature, conduct over two winter seasons and its pragmatic design, with broad inclusion criteria representing typical patients admitted to UK secondary care, simple intervention and comparison to routine clinical care. These factors suggest that the findings of this study are likely to be generalisable to other similar UK and international centres. The trial was designed and powered to evaluate influenza specific outcomes in hospitalised adults and universal sampling with retrospective testing for influenza in the control group allowed a direct assessment of missed diagnoses and comparison of outcomes in all influenza-infected patients. The use of clinical outcomes measures such as time on supplementary oxygen, time to clinical improvement and hospital recovery score, allowed the assessment of the impact of mPOCT on clinical outcomes in addition to changes in clinical management.

Our study has a number of potential weakness. As the study took place only during winter months when influenza was circulating the findings cannot be extrapolated to times outside of this period when mPOCT for influenza is unlikely to have an impact. In this study no attempt was made at blinding clinical teams or participants to the allocated groups. As the study required clinical teams to be informed of the mPOCT results in real-time, they could not be blinded to which group a participant had been randomised to.  In our trial the mPOCT was performed by research staff rather than clinicians and so uncertainties remain about how testing should be best delivered in routine care. Various implementation models exist including physician or nurse-delivered testing and technician-delivered testing within a POCT testing hub, embedded in acute area.  These models should be studied in implementation trials alongside obtaining real world data to confirm the clinical impact.  Another uncertainty regarding the implementation of mPOCT is the choice between syndromic testing for a comprehensive panel of respiratory viruses or testing for influenza alone. Our previous work suggests that the impact on antibiotic use and length of stay will be greatest when the syndromic approach is used although this must be balanced against the current higher cost and longer turnaround times of most current systemic panels, compared to tests for influenza alone.30

In conclusion a routine molecular point-of-care ‘test and treat’ strategy for influenza in adults hospitalised with ARI improved the detection of influenza, the appropriate use of influenza antivirals and the rational use of isolation facilities, and was associated with improvements in clinical outcome, in influenza-infected patients. Routine mPOCT for influenza should become the standard of care in hospitals during influenza season.

References

1. Thompson WW, Shay DK, Weintraub E, et al. Influenza-associated hospitalizations in the United States. JAMA 2004;292:1333–40.

2. Dao CN, Kamimoto L, Nowell M, et al. Emerging Infections Program Network: adult hospitalizations for laboratory-positive influenza during the 2005-2006 through 2007-2008 seasons in the United States. J Infect Dis 2010;202:881-8.

3. Troeger CE, Blacker BF, Khalil IA, et al. Mortality, morbidity, and hospitalisations due to influenza lower respiratory tract infections, 2017: an analysis for the Global Burden of Disease Study 2017. Lancet Respir Med 2019;7:69-89.

4. Widmer K, Zhu Y, Williams JV, et al. Rates of hospitalizations for respiratory syncytial virus, human metapneumovirus, and influenza virus in older adults. J Infect Dis 2012;206:56-62.

5. Mauskopf J, Klesse M, Lee S, et al. The burden of influenza complications in different high-risk groups: a targeted literature review. J Med Econ 2013;16:264-77.

6. Dugas AF, Valsamakis A, Atreya MR, et al. Clinical diagnosis of influenza in the ED. Am J Emerg Med 2015;33:770-5.

7. Brendish NJ, Malachira AK, Armstrong L, et al. Routine molecular point-of-care testing for respiratory viruses in adults presenting to hospital with acute respiratory illness (ResPOC): a pragmatic, open-label, randomised controlled trial. Lancet Respir Med 2017;5:401-411.

8. Guidance on use of antiviral agents for the treatment and prophylaxis of seasonal influenza. Public Health England. Version 10.0. September 2019. https://assets.publishing.service.gov.uk/government/uploads/system/uploads/attachment_data/file/833572/PHE_guidance_antivirals_influenza_201920.pdf

9. Uyeki TM, Bernstein HH, Bradley JS, Englund JA, File TM and Fry AM. Clinical Practice Guidelines by the Infectious Diseases Society of America: 2018 Update on Diagnosis, Treatment, Chemoprophylaxis, and Institutional Outbreak Management of Seasonal Influenza. Clin Infect Dis 2019;68:e1-e47.

10. Muthuri SG, Venkatesan S, Myles PR, et al. Effectiveness of neuraminidase inhibitors in reducing mortality in patients admitted to hospital with influenza A H1N1pdm09 virus infection: a meta-analysis of individual participant data. Lancet Respir Med 2014;2:395-404.

11. Katzen J, Kohn R, Houk JL, Ison MG. Early Oseltamivir After Hospital Admission Is Associated With Shortened Hospitalization: A 5-Year Analysis of Oseltamivir Timing and Clinical Outcomes. Clin Infect Dis 2019;69:52-58.

12. Brendish NJ, Schiff HF, Clark TW. Point-of-care testing for respiratory viruses in adults: The current landscape and future potential. J Infect 2015;71:501-10.

13. Huang HS, Tsai CL, Chang J, Hsu TC, Lin S, Lee CC. Multiplex PCR system for the rapid diagnosis of respiratory virus infection: systematic review and meta-analysis. Clin Microbiol Infect 2018;24:1055-1063.

14. Merckx J, Wali R, Schiller I, et al. Diagnostic Accuracy of Novel and Traditional Rapid Tests for Influenza Infection Compared With Reverse Transcriptase Polymerase Chain Reaction: A Systematic Review and Meta-analysis. Ann Intern Med 2017;167:394-409.

15. Vos LM, Bruning AHL, Reitsma JB, et al. Rapid Molecular Tests for Influenza, Respiratory Syncytial Virus, and Other Respiratory Viruses: A Systematic Review of Diagnostic Accuracy and Clinical Impact Studies. Clin Infect Dis 2019;69:1243-1253.

16. Trabattoni E, Le V, Pilmis B, et al. Implementation of Alere i Influenza A & B point of care test for the diagnosis of influenza in an emergency department. Am J Emerg Med 2018. 36:916-921.

17. Vecino-Ortiz AI, Goldenberg SD, Douthwaite ST, et al. Impact of a multiplex PCR point-of-care test for influenza A/B and respiratory syncytial virus on an acute pediatric hospital ward. Diagn Microbiol Infect Dis 2018;91:331-335.

18. Egilmezer E, Walker GJ, Bakthavathsalam P, et al. Systematic review of the impact of point-of-care testing for influenza on the outcomes of patients with acute respiratory tract infection. Rev Med Virol 2018;28:e1995.

19. Youngs J, Marshall B, Farragher M et al. Implementation of influenza point-of-care testing and patient cohorting during a high-incidence season: a retrospective analysis of impact on infection prevention and control and clinical outcomes. Hosp Infect 2019;101:276-284.

20. Beard K, Brendish N, Malachira A, et al. Pragmatic multicentre randomised controlled trial evaluating the impact of a routine molecular Point-of-Care ‘test-and-treat’ strategy for influenza in adults hospitalised with acute respiratory illness (FluPOC): trial protocol. BMJ Open 2019;9:e031674.

21. Chu HY, Englund JA, Huang D, et al. Impact of rapid influenza PCR testing on hospitalization and antiviral use: a retrospective cohort study. J Med Virol 2015;87:2021‐2026.

22. Brendish NJ, Malachira AK, Beard KR, Ewings S, Clark TW. Impact of turnaround time on outcome with point-of-care testing for respiratory viruses: a post hoc analysis from a randomised controlled trial. Eur Respir J 2018;52:pii:1800555.

23. Nasir Wabe, Ling Li, Maria R Dahm, et al. Timing of respiratory virus molecular testing in emergency departments and its association with patient care outcomes: a retrospective observational study across six Australian hospitals. BMJ Open 2019;9:e030104.

24. Li L, Georgiou A, Vecellio E, et al. The effect of laboratory testing on emergency department length of stay: a multihospital longitudinal study applying a cross-classified random-effect modeling approach. Acad Emerg Med 2015;22:38-46.

25. Brendish NJ, Clark TW. Molecular Point-of-care Testing for Influenza to Improve Early Neuraminidase Inhibitor Treatment and Outcomes in Hospitalized Adults. Clin Infect Dis 2019;68:2154-2155.

26. Branche AR, Walsh EE, Vargas R, et al. Serum Procalcitonin Measurement and Viral Testing to Guide Antibiotic Use for Respiratory Infections in Hospitalized Adults: A Randomized Controlled Trial. J Infect Dis 2015;212:1692-700.

27. McGeer A, Green KA, Plevneshi A, et al. Antiviral therapy and outcomes of influenza requiring hospitalization in Ontario, Canada. Clin Infect Dis 2007;45:1568-75.

28. Venkatesan S, Myles PR, Bolton KJ, et al. Neuraminidase Inhibitors and Hospital Length of Stay: A Meta-analysis of Individual Participant Data to Determine Treatment Effectiveness Among Patients Hospitalized With Nonfatal 2009 Pandemic Influenza A(H1N1) Virus Infection. J Infect Dis 2020;221;356-366.

29. Lee N, Chan PK, Choi KW, et al. Factors associated with early hospital discharge of adult influenza patients. Antivir Ther 2007;12:501-8.

30. Brendish NJ, Mills S, Ewings S, Clark TW. Impact of point-of-care testing for respiratory viruses on antibiotic use in adults with exacerbation of airways disease. J Infect 2019;79:357-362.

Author contributions

TWC reviewed the medical literature, conceived of and designed the study, oversaw the conduct of the study, participated in the data of collection and interpretation or results, drafted and wrote the manuscript. KRB reviewed the medical literature participated in the trial design, recruited patients, generated and collected data, assisted with drafting the manuscript. NJB, AKM, SP, SM, CC, EN, and LP screened and recruited patients and collected data. NC acted as PI for the second trial site and oversaw the running of the study there, SE designed the statistical analysis plan and analysed the data. All authors reviewed and contributed to the manuscript during its development.

Declaration of interests

TWC has received speaker fees, honoraria, travel reimbursement, and equipment and consumables free of charge for the purposes of research outside of this submitted study, from BioFire diagnostics LLC and BioMerieux. TWC has received consultancy fees from Synairgen research Ltd, Randox laboratories Ltd and Cidara therapeutics. He a member of an advisory board for Roche and a member of two independent data monitoring committees for trials sponsored by Roche. He has acted as the UK chief investigator for an IMP study sponsored by Janssen. SM has received speaker fees and travel reimbursement from BioMerieiux.  All other authors declare that they have no conflicts of interests.

Data sharing

All de-identified participant data analysed and presented in this study are available from the corresponding author following publication, on reasonable request providing this meets local ethical and research governance criteria.

Acknowledgements

We thank all of the patients recruited at the Southampton General Hospital and Hampshire County Hospital and all the clinical staff involved including doctors, nurses, and laboratory technicians. We thank the directors, research nurses, data managers, clinical trials assistants and laboratory staff at the NIHR Clinical Research Facility and the NIHR Southampton Biomedical Research Centre. We thank the staff at the R&D department, University Hospital Southampton NHS Foundation Trust and the NIHR Clinical Research Network, Wessex for their support throughout the trial. In particular we thank the data analysts Ms. Florina Borca and Dr. Hang Phan at the CIRU-AXIS team for their help with data analysis. Lastly we thank Dr. Kordo Saeed, consultant microbiologist and Ana Maria Arias, research nurse, at the Royal Hampshire County Hospital for all their assistance during the trial.

This report is independent research supported by the National Institute for Health Research (NIHR Post Doctorial Fellowship, Dr Tristan Clark, PDF 2016-09-061). The views expressed in this publication are those of the author(s) and not necessarily those of the NHS, the National Institute for Health Research or the Department of Health.

Table 1: Baseline characteristics of all patients (ITT population) and influenza-infected patients (ITTI population).

| Characteristic | All patients (ITT) |  | Influenza-infected patients (ITTI) |  |
| --- | --- | --- | --- | --- |
|  | mPOCT (n=307) | Control (n=306) | mPOCT (n=100) | Control (n=102) |
| Age (years) Sex | 62 (41 to 74) | 63 (47 to 76) | 63 (42 to 73) | 65 (51 to 78) |
| Female Male | 175 (57%) 132 (43%) | 157 (51%) 148 (49%) | 53 (53%) 47 (47%) | 47 (46%) 55 (54%) |
| Ethic origin White British Other | 285 (93%) 22 (7%) | 284 (93%) 21 (7%) | 95 (95%) 5 (5%) | 94 (93%) 6 (7%) |
| Current smoker Yes No | 71 (23%) 235 (77%) | 70 (23%) 235 (77%) | 23 (23%) 77 (77%) | 24 (24%) 78 (76%) |
| Influenza vaccine* Yes No | 199 (65%) 107 (35%) | 185 (61%) 117 (39%) | 60 (60%) 40 (40%) | 56 (55%) 46 (45%) |
| Pregnant Yes No | 4 (1%) 302 (99%) | 9 (3%) 297 (97%) | 1 (1%) 99 (99%) | 2 (2%) 100 (98%) |
| Duration of symptoms (days)† | 4 (3 to 6) | 4 (3 to 7) | 4 (3 to 7) | 4 (3 to 6) |
| Antibiotics within 14 days Yes No | 95 (31%) 209 (69%) | 99 (33%) 205 (67%) | 30 (30%) 70 (70%) | 34 (33%) 68 (67%) |
| Antivirals within 14 days Yes No | 0 (0) 305 (100%) | 2 (1%) 302 (99%) | 0 (0) 100 (100%) | 1 (1%) 101 (99%) |
| Comorbidity |  |  |  |  |
| Hypertension Cardiovascular disease Respiratory disease Renal disease Liver disease Diabetes mellitus Cancer Immunosuppression Charlson Comorbidity Index score | 82 (27%) 61 (20%) 216 (70%) 16 (5%) 2 (1%) 36 (12%) 18 (6%) 14 (5%) 4 (4 to 7) | 89 (29%) 64 (21%) 189 (62%) 18 (6%) 6 (2%) 57 (19%) 21 (7%) 12 (4%) 4 (3 to 8) | 26 (26%) 22 (22%) 60 (60%) 4 (4%) 1 (1%) 12 (12%) 7 (7%) 3 (3%) 4 (0 to 4) | 29 (28%) 17 (17%) 49 (48%) 6 (6%) 0 (0%) 15 (15%) 9 (9%) 4 (4%) 4 (0 to 7) |
| Observations at admission |  |  |  |  |
| Temperature (°C ) Temp ≥38°C Pulse rate (bpm) Respiratory rate (bpm) Oxygen saturations Supplementary O2 Blood pressure (mmHg) Systolic Diastolic NEWS2 score | 37·1 (36·5 to 38·1) 82 (27%) 100 (88 to 117) 23 (20 to 26) 95 (92 to 97) 54 (18%) 134 (120 to 150) 74 (63 to 84) 4 (3 to 6) | 37·1 (36·4 to 37·9) 70 (23%) 100 (85 to 115) 22 (19 to 27) 95 (92 to 97) 71 (23%) 135 (120 to 148) 75 (65 to 83) 4 (3 to 6) | 37·3 (36·7 to 38·1) 30 (30%) 103 (90 to 116) 22 (19 to 27) 95 (93 to 97) 21 (21%) 135 (122 to 148) 71 (64 to 81) 4 (3 to 7) | 37·5 (36·7 to 38·3) 35 (35%) 100 (85 to 115) 22 (19 to 28) 95 (91 to 97) 22 (22%) 136 (117 to 146) 75 (65 to 84) 5 (3 to 7) |
| Laboratory and radiology |  |  |  |  |
| CRP (mg/L) WCC (109/L) Neutrophils (109/L) Lymphocytes (109/L) CXR Yes No CT chest Yes No | 48 (13 to 127) 10·5 (7·9 to 14·6) 8·1 (5·8 to 12·1) 1·1 (0·7 to 1·6) 302 (98%) 5 (2%) 19 (6%) 288 (94%) | 38 (14 to 99) 9·9 (7·3 to 13·3) 7·8 (5·2 to 10·9) 1·1 (0·7 to 1·7) 299 (98%) 25 (8%) 25 (8%) 281 (92%) | 51 (17 to 138) 8·7 (6·7 to 11·5) 6·6 (5·1 to 9·6) 0·9 (0·6 to 1·4) 98 (98%) 2 (2%) 2 (2%) 98 (98%) | 49 (18 to 98) 8·1 (5·8 to 11·0) 6·0 (4·0 to 8·1) 0·8 (0·6 to 1·4) 102 (100%) 0 (0) 4 (4%) 98 (96%) |
| Final Diagnosis |  |  |  |  |
| Asthma COPD exacerbation Exacerbation other chronic lung disease Pneumonia Influenza-like illness NPLRTI Other | 59 (19%) 77 (25%) 8 (3%) 88 (29%) 27 (9%) 31 (10%) 17 (6%) | 55 (18%) 79 (26%) 9 (3%) 71 (23%) 38 (12%) 38 (12%) 16 (5%) | 22 (22%) 19 (19%) 2 (2%) 23 (23%) 21 (21%) 12 (12%) 1 (1%) | 14 (14%) 23 (23%) 1 (1%) 20 (20%) 32 (31%) 11 (11%) 1 (1%) |

Data are n (%) or median (IQR). NEWS2=national early warning score 2, CRP=C-reactive protein, WCC=white cell count, CXR=chest X-ray,

COPD=chronic obstructive airways disease, NPLRTI=non-pneumonic lower respiratory tract infection. *For the current influenza season.

†Prior to admission to hospital.

Table 2: Influenza testing, antiviral use and isolation facility use in all patients (ITT population).

| Outcome | mPOCT (n=307) | Control (n=306) | Difference or Relative Risk (95% CI) | p value |
| --- | --- | --- | --- | --- |
| Tested for influenza during admission | 307 (100%) | 212 (69%) | 1·4 (1·4 to 1·5) | <0·0001 |
| Turnaround time for results (hours)* | 1·2 (1·1 to 1·4) | 22·6 (16·0 to 28·7) | 21·4 (20·6 to 22·4) | <0·0001 |
| Influenza infected Influenza A Influenza B Influenza A+B | 100 (33%) 63/100 (63%) 37/100 (37%) 0/100 (0%) | 102 (33%) 76/102 (75%) 24/102 (24%) 2/102 (2%) | 1·0 (0·9 to 1·2) | 0·86 |
| Antiviral use |  |  |  |  |
| Total antiviral use Oseltamivir Other Influenza positive patients Duration of antivirals (days) All antiviral treated pts Influenza positive Influenza negative | 139 (45%) 139 (45%) 0 (0%) 99/139 (71%) 4·7 (0·7 to 5·1) 4·8 (4·6 to 5·5) 0·1 (0·1 to 0·5) | 130 (43%) 130 (43%) 0 (0%) 63/130 (49%) 4·5 (1·0 to 5·0) 4·5 (4·2 to 5·0) 1·5 (0·5 to 4·5) | 1·1 (0·9 to 1·3) 1·5 (1·2 to 1·8) -0·2 (-0·5 to 0·0) -0·3 (-0·6 to -0·1) 1·4 (0·7 to 2.·1) | 0·49 <0·0001 0·11 0·0024 <0·0001 |
| Isolation facility use† |  |  |  |  |
| Isolated at presentation Isolated at any time Influenza positive Duration of isolation (days) All isolated patients Influenza positive Influenza negative | 30 (10%) 106 (35%) 70/106 (66%) 2·5 (1·3 to 4·5) 2·8 (1·9 to 4·5) 1·4 (0·5 to 3·5) | 26 (9%) 67 (22%) 39/67 (58%) 1·8 (1·0 to 4·4) 2·4 (1·0 to 4·7) 1·4 (1·0 to 3·9) | 1.2 (0.7 to 1.9) 1·6 (1·2 to 2·1) 1·1 (0·9 to 1·5) -0·7 (-0·9 to 0·3) -0·4 (-1·2 to 0·3) 0·0 (-0·5 to 0·8) | 0.58 0·0006 0·33 0·29 0·21 0·42 |

Data are n (%) or median (IQR). CI=confidence interval.*Based on those that were tested; for mPOCT group this is the turnaround time for mPOCT results, for the control group it is for laboratory PCR results. †Isolation facility use relates to the use of single room accommodation or dedicated influenza cohort bays for patients.

Table 3: Outcomes in influenza-infected patients (ITTI population).

| Outcome | mPOCT (n=100) | Control (n=102) | Difference or Relative Risk (95% CI) | p value |
| --- | --- | --- | --- | --- |
| Diagnosis of influenza during admission PCR positive Clinical diagnosis* | 100 (100%) 100 (100%) 0 (0%) | 60 (59%) 55 (54%) 5 (5%) | 1·7 (1·7 to 1·7) | <0·0001 |
| Turnaround time of results (hours) | 1·2 (1·1 to 1·4) | 23·1 (13·8 to 31·5) | 21·9 (20·2 to 23·8) | <0·0001 |
| Antiviral use |  |  |  |  |
| Antiviral use Empirical antiviral use† Directed use‡ Time to antiviral use (hours) Time to antiviral use, excluding those treated empirically (hours) | 99 (99%) 30/99 (30%) 69/99 (70%) 1·0 (0 to 2·0) 2·0 (1·0 to 3·0) | 63 (62%) 41/63 (65%) 22/63 (35%) 6·0 (0 to 12·0) 30·5 (12·0 to 59·0) | 1·6 (1·4 to 1·9) 0·6 (0·4 to 0·7) 2.0 (1.4 to 2.9) 5·0 (0 to 6·0) 28·5 (18·0 to 31·0) | <0·0001 <0·0001 - 0·0039 <0·0001 |
| Isolation facility use§ |  |  |  |  |
| Isolated at presentation Isolated during admission Isolated during admission excluding exceptions¶ Time to isolation (hours) Time to isolation, excluding those isolated empirically (hours) | 10 (10%) 70 (70%) 70/76 (92%) 4·5 (1·9 to 9·8) 5·0 (2·0 to 12) | 9 (9%) 39 (38%) 39/90 (43%) 25·0 (3·6 to 45·1) 32·0 (16·0 to 55·0) | 1·1 (0·5 to 2·7) 1·8 (1·4 to 2·4) 2·1 (1·7 to 2·8) 20·5 (6·0 to 26·5) 27·0 (17·0 to 33·0) | 0·77 <0·0001 <0·0001 <0·0003 <0·0001 |
| Antibiotic use |  |  |  |  |
| Antibiotic use (any) Duration of antibiotics (days) Single dose of antibiotics only Antibiotics given for <24 hours IV antibiotics Duration of IV antibiotics (days) Single dose of IV antibiotics only IV antibiotics given for <24 hours | 90 (90%) 6·1 (4·1 to 7·2) 9 (10%) 13 (14%) 68 (68%) 0·5 (0·1 to 1·8) 29/68 (43%) 40/68 (59%) | 98 (96%) 5·8 (4·2 to 7·4) 6 (6%) 7 (7%) 69 (68%) 1·1 (0·1 to 2·0) 19/68 (28%) 34/68 (50%) | 0·9 (0·9 to 1·0) -0·3 (-0·8 to 0·8) 0·9 (0·9 to 1·1) 0·9 (0·8 to 1·1) 1·0 (0·8 to 1·2) 0·6 (0 to 0·5) 0.8 (0·6 to 1·0) 0·8 (0·6 to 1·2) | 0·089 0·96 0·32 0·10 0·96 0·31 0·064 0·26 |
| Clinical outcome measures |  |  |  |  |
| Length of stay (days) Discharged within 24 hours | 2·7 (1·1 to 4·2) 23 (23%) | 2·7 (1·1 to 5·4) 18 (18%) | 0·0 (-0·4 to 0·9) 0·9 (0·8 to 1·1) | 0·37 0·34 |
| Duration of supplementary O2 (hours) | 32·0 (15·0 to 59·0) | 33·0 (14·0 to 81·0) | 1·0 (-8·0 to 18·0) | 0·48 |
| Time to clinical improvementǁ (hours) | 8·8 (3·2 to 23·9) | 13·9 (5·1 to 42·5) | 5·1 (-0·3 to 9·9) | 0·077 |
| HRS scoreǁ, day 7 1. Home 2. Hospital ward not requiring O2 3. Hospital ward requiring O2 4. ICU not ventilated 5. ICU ventilated 6. Death | 90 (90%) 8 (8%) 2 (2%) 0 (0%) 0 (0%) 0 (0%) | 82 (80%) 12 (12%) 5 (5%) 1 (1%) 1 (1%) 1 (1%) |  | 0·045** |

Data are n (%) or median (IQR). PCR=polymerase chain reaction, IV=intravenous, ICU=intensive care unit, O2=Oxygen *A clinical diagnosis of influenza is defined as either influenza listed as a diagnosis on the discharge summary or a full 5 day course of oseltamivir given, in the absence of diagnostic testing performed. †Empirical antiviral use denotes antivirals started in patients with a presumptive clinical diagnosis of influenza before diagnostic results were available. ‡Directed antiviral use denotes antiviral use started once diagnosis has been confirmed by a positive diagnostic test. All antiviral use was with oseltamivir. §Isolation facility use relates to the use of single room accommodation or dedicated influenza cohort bays for patients. ¶Exceptions are those patients discharged rapidly or nursed in areas without isolation facilities. ǁTime to clinical improvement and the Hospital Recovery Score (HRS) are defined in the supplementary appendices. **Mann Whitney U test.

Table 4: Adverse events in influenza-infected patients (ITTI).

| Adverse event | mPOCT (n=100) | Control (n=102) | Relative Risk (95% CI) | p value |
| --- | --- | --- | --- | --- |
| Total number of Adverse Events (AEs) | 19 | 34 |  |  |
| Patients with AE | 16 (16%) | 23 (23%) | 0·7 (0·4 to 1·3) | 0·24 |
| ICU or HDU admission | 1 (10%)* | 6 (6%)† | 0·2 (0·0 to 1·4) | 0·058 |
| Death In hospital Within 30 days | 0 (0%) 0 (0%) | 2 (2%)‡ 2 (2%)‡ | NA NA | 0·16 0·16 |
| Readmission within 30 days | 6 (6%) | 5 (5%) | 1·2 (0·4 to 3·8) | 0·76 |
| Representation to hospital but without admission, within 30 days | 2 (2%) | 1 (1%) | 2·0 (0·2 to 21·7) | 0·56 |
| Prolonged hospitalisation (>7 days) | 10 (10%) | 20 (20%) | 0·5 (0·3 to 1·0) | 0·055 |

Data are n (%) unless stated otherwise. ICU=intensive care unit, HDU=high dependency unit. *Influenza A H3. †Influenza A H3 x2, Influenza A H1 x1, Influenza A untypable x1, Influenza B x2. ‡Influenza B x1, Influenza A untypable x1.

Figure legends

Figure 1: Trial profile

Figure 2: Time-to-event curve showing antiviral use over time in influenza-infected patients

Figure 3: Time-to-event curve showing isolation facility use over time in influenza-infected patients

Supplementary Figures

Figure S1.Details for subgroup analysis for time to antiviral and time to isolation facility use in influenza infected patients

Figure S2. Hospital Recovery Scale (HRS) for influenza-infected patients on

day 4 (a) and day 7 (b)

Figure S3. Time to clinical improvement (TCCI) for influenza-infected patients

Figure S4. Time on supplementary Oxygen for influenza-infected patients

Figure S5. Time to discharge for influenza-infected patients

Research in context

Evidence before this study

We searched PubMed, the Cochrane Controlled Clinical Trials Register and ClinicalTrials.gov database, for relevant published articles and ongoing trials assessing the clinical impact of molecular POCT (mPOCT) for influenza in hospitalised adults with acute respiratory illness. We used the search terms; ”point-of-care testing” or “rapid PCR testing” or ‘’rapid molecular testing’’ and “influenza” and “hospital” and “adult” and “clinical trial” or “randomised controlled trial” or ‘’trial’’ or “study”. We limited the search to studies published between Jan 1, 1980, and Jan 1, 2020, in English.  We excluded studies using antigen-based (including digital immunoassays) point-of-care tests for influenza, studies in children, and studies reporting only diagnostic accuracy. We found no Cochrane systematic reviews for mPOCT in adults. We found 2 randomised controlled trials of mPOCT for influenza and other respiratory viruses evaluating clinical outcomes and 1 quasi-randomised trial of low quality. One of these studies was performed in China and did not specifically report the impact of influenza detection on outcomes. The other two studies suggested improvement in neuraminidase inhibitor (NAI) use and isolation facility use with mPOCT but were constrained by low number of influenza infected patients or other methodological issues. We found several retrospective observational studies in hospitalised adults evaluating the impact of mPOCT on clinical outcomes including NAI use, antibiotics, length of stay and number of investigations. These all broadly suggested improvements in NAI use and investigations but did not reliably evaluate isolation facility use and were conflicting regarding the impact on antibiotic use. One study suggested a reduction in nosocomial spread of influenza. No studies reported on the impact of mPOCT for influenza on clinical outcomes other than length of stay. We found a systematic review of studies evaluating clinical impact and diagnostic accuracy of rapid molecular test for respiratory viruses, only some of which were performed at the point of care and all are already mentioned above.

Added value of this study

This multicentre, pragmatic, randomised control trial of mPOCT for influenza was designed and powered to evaluate influenza-specific outcomes in hospitalised adults. Retrospective testing for influenza in all patients in the control group allows a direct assessment of missed diagnoses and comparison of outcomes in all influenza-infected patients managed with mPOCT or routine clinical care. The use of clinical outcomes measures such as time to clinical improvement, time on supplementary oxygen, and the hospital recovery scale score, allows assessment of the impact of mPOCT on clinical outcomes in addition to the impact on clinical management (use of antivirals, antibiotics, isolation facility use etc). Whilst this study does not definitively exclude the possibility that routine laboratory PCR testing could lead to some of the improvements seen with POCT, this would require dramatic reductions in current turnaround times that are unlikely to be achievable in most laboratories due to batch testing and other constraints.

Implications of all the available evidence

Routine mPOCT for influenza in hospitalised adults improves the detection rate of influenza, and the appropriate use of influenza antivirals and isolation facilities. It was also associated with improvements in clinical outcome in influenza-infected patients. Routine mPOCT for influenza should become standard of care in hospitals during influenza season.