## Propagation loss
Measuring the propagation loss for silicone waveguides in light guiding catheters
818-sl detector
Waveguide 1 | LS-6946 | Batch 1 | 1 mm fiber 0.5 na
Cut back position | Waveguide length (cm) | Background | Optical power (mW) | Optical power (mW) | Dbm | Loss Db/cm | Propagation length from normalisation start (cm) | Normalised loss dB
1 | 13.4 | 0.00254 | 0.01695 | 0.01441 | -18.4134 |  | 13.5 | -15.6291
2 | 5.4 | 0.00559 | 0.5323 | 0.52671 | -2.7843 | -1.9536 | 0 | 0
3 | 0 | 0.002 | 15.4598 | 15.4578 | 11.8915 | -2.7177
[5 empty] | Mean Propagation loss | -1.9536
[5 empty] | Coupling loss | -4.1261
Waveguide 4 | LS-6946 | Batch 1 | 1 mm fiber 0.5 na
Cut back position | Waveguide length (cm) | Background | Optical power (mW) | Optical power (mW) | Dbm | Loss Db/cm | Propagation length from normalisation start (cm) | Normalised loss dB
1 | 25.2 | 0.001 | 0.01695 | 0.0036 | -24.437 |  | 15.2 | -16.5052
2 | 10 | 0.001 | 0.5323 | 0.161 | -7.9317 | -1.0859 | 0 | 0
[5 empty] | Mean Propagation loss | -1.0859
Coupling loss
Waveguide 2 | LS-6946 | Batch 2 | 1mm fibre 0.5 na
Cut back position | Waveguide length (cm) | Background | Optical power (mW) | Optical power (mW) | Dbm | Loss Db/cm | Length | Normalised loss dB
1 | 27.4 | 0.002 | 0.0077 | 0.0057 | -22.4413 |  | 16 | -17.4873
2 | 19.4 | 0.002 | 0.0345 | 0.0325 | -14.8812 | -0.94501 | 8 | -9.9272
3 | 11.4 | 0.003 | 0.3226 | 0.3196 | -4.9539 | -1.2409 | 0 | 0
4 | 3.4 | 0.003 | 3.6379 | 3.6349 | 5.605 | -1.3199
5 | 0 | 0.003 | 24.3333 | 24.3303 | 13.8615
[5 empty] | Mean Propagation loss | -1.093
[5 empty] | Coupling loss | -6.3557
Waveguide 3 | LS-69436 | Batch 2 | 600 um 0.22 NA
Cut back position | Waveguide length (cm) | Background | Optical power (mW) | Optical power (mW) | Dbm | Loss Db/cm | Length | Normalised loss dB
1 | 24.8 | 0.001 | 0.0122 | 0.0112 | -19.5078 |  | 16 | -15.8526
2 | 16.8 | 0.001 | 0.06516 | 0.06416 | -11.9274 | -0.94756 | 8 | -8.2721
3 | 8.8 | 0.002 | 0.433 | 0.431 | -3.6552 | -1.034 | 0 | 0
4 | 0.8 | 0.003 | 3.31 | 3.307 | 5.1943 | -1.1062
5 | 0 | 0.003 | 4.791 | 4.788 | 6.8015
[5 empty] | Mean Propagation loss | -0.99079
[5 empty] | Coupling loss | -1.7378
[5 empty] | LS-69456 | Average |  | Std
[5 empty] | Propagation loss dB/cm | -1.2808 |  | 0.39054

## Side emission
Measuring the side emission irradiance and calculating the irradiance at the far wall of the catheter lumen
|  | Photodiode measurement (mW) |  |  |  | Estimated Intensity at far wall of lumen (mW) Z=2.9mm | [6 empty] | Relative Photodiode measurement (mW)
|  | Waveguide |  |  |  | Waveguide | [6 empty] | Waveguide
|  | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 |  |  |  | 1 | 2 | 3 | 4
[10 empty] | Mean | Stdev | Running mean |  |  |  |  | Mean | Stdev
Displacement i (cm) | 0 | 4.288 | 3.4332 | 5.1807 | 8.7362 | 57.7589 | 46.2452 | 69.7836 | 117.6771 | 72.8662 | 31.3817 | 72.8662 | 1 | 1 | 1 | 1 | 1 | 0
| 1 | 2.2456 | 1.9038 | 2.3709 | 2.9697 | 30.2477 | 25.6445 | 31.9358 | 40.0018 | 31.9574 | 5.9857 | 52.4118 | 0.52369 | 0.55453 | 0.45764 | 0.33993 | 0.46895 | 0.095036
| 2 | 0.60132 | 0.69737 | 0.59211 | 0.59211 | 8.0997 | 9.3936 | 7.9757 | 7.9757 | 8.3611 | 0.69075 | 37.7283 | 0.14023 | 0.20312 | 0.11429 | 0.067776 | 0.13136 | 0.05646
| 3 | 0.18268 | 0.25917 | 0.20454 | 0.18462 | 2.4606 | 3.4911 | 2.7552 | 2.4869 | 2.7984 | 0.48054 | 28.9958 | 0.042602 | 0.07549 | 0.039481 | 0.021133 | 0.044677 | 0.022621
| 4 | 0.070476 | 0.12822 | 0.089995 | 0.093189 | 0.94932 | 1.7271 | 1.2122 | 1.2553 | 1.286 | 0.32369 | 23.4538 | 0.016436 | 0.037347 | 0.017371 | 0.010667 | 0.020455 | 0.011645
| 5 | 0.037973 | 0.071546 | 0.048115 | 0.05159 | 0.5115 | 0.96372 | 0.64811 | 0.69492 | 0.70456 | 0.18949 | 19.6623 | 0.0088557 | 0.020839 | 0.0092874 | 0.0059053 | 0.011222 | 0.0065855
| 6 | 0.022174 | 0.044577 | 0.02619 | 0.026332 | 0.29869 | 0.60045 | 0.35279 | 0.35469 | 0.40165 | 0.13505 | 16.9108 | 0.0051713 | 0.012984 | 0.0050554 | 0.0030141 | 0.0065562 | 0.0043983
| 7 | 0.013456 | 0.026748 | 0.017661 | 0.015294 | 0.18125 | 0.3603 | 0.23789 | 0.20601 | 0.24636 | 0.079417 | 14.8277 | 0.0031381 | 0.0077911 | 0.0034089 | 0.0017506 | 0.0040222 | 0.0026155
| 8 | 0.0086291 | 0.017334 | 0.01055 | 0.010319 | 0.11623 | 0.23349 | 0.14211 | 0.139 | 0.15771 | 0.051823 | 13.1977 | 0.0020124 | 0.005049 | 0.0020364 | 0.0011812 | 0.0025698 | 0.0017

## Power budget
LED to fibre coupling loss |  |  |  | Optical Power budget for system
| Output power (mW) |  |  |  | Coupling loss
| LED | Fibre | DB |  | LED >>> fibre | fibre>>>waveguide | fibre>>>catheter body
Max | 675 | 81.6092 | -9.1756 | Min loss (dB) | -9.1756 | -3.6
Min | 675 | 67.8161 | -9.9797 | Max loss (dB) | -9.9797 | -5.8
| 675 | 67.0103 | -10.0316 | Average loss (dB) | -9.711 | -4.7
| Mean | 72.1452 | -9.729 | Cumulative loss (dB) | -9.711 | -14.411
| SD | 8.206 | 0.47991 | Power propagating (%) | 10.6882 | 3.6216
|  |  |  | Power propogating (mW) | 72.1452 | 2.4269 | 69.7183
|  |  |  | mW/cm2 |  | 2.0224 | 58.0986
| 675 | 63.9392 | -10.2354
| 675 | 80.3512 | -9.2431

## SEM statistics
By eye counts from SEM images of intact cells on silicone surfaces
| B11 |  | B41 |  | B42 |  | Control 1 |  | Control 2 |  | Control 3 |  | Welchs T test | [10 empty] | Width | Height
FOV | No Cells | Notes | No Cells | Notes | No Cells | Notes | No Cells | Notes | No Cells | Notes | No Cells | Notes |  | Control Catheter (X1) | Light guiding Catheter (X2) | X2-X1 | t value | Degrees freedom | P 1 tailed | P 2 tailed | 95% Confidence Interval of difference Lower | 95% Confidence interval of difference Upper | Field of view | 103.8 | 74.6 | 7743.48 | um2
1 | 177 |  | 82 |  | 498 | globules DISPLAYED | 427 | clusters without mucus | 1208 | evenly spaced bright cells | 1340 | evenly spaced bright cells | Mean (CFU) | 12613534.2667 | 2995761.0222 | -9617773.2444 | -8.6136 | 2.6428 | <0.005 | <0.01 | -14422400.3673 | -4813146.1216 |  |  |  | 7.7435e-05 | cm2
2 | 157 |  | 135 |  | 17 | dim/ sparse cells | 991 | evenly spaced bright cells | 1060 | evenly spaced bright cells | 1290 | evenly spaced bright cells | Stdev (CFU) | 1791726.7895 | 727971.1048 |  |  |  |  | Kill factor>> | Infinity | 1.617 |  |  | Multiplication factor for cell density | 12914.0903
3 | 283 |  | 234 | globules | 261 |  | 710 | Globules with mucus | 866 | evenly spaced bright cells | 1196 | evenly spaced bright cells DISPLAYED | SEM | 1034453.9442 | 420294.3133 | 1116576.1383
4 | 311 |  | 131 | Displayed | 439 | globules | 1128 | DISPLAYED evenly spaced bright cells | 1416 | evenly spaced bright cells | 91 | single clump , some fragments | Technical replicates | 5 | 5
5 | 145 |  | 255 |  | 71 | dim/ sparse cells | 828 | evenly spaced bright cells | 774 | evenly spaced bright cells | 1326 | evenly spaced bright cells | Biological replicates | 3 | 3 |  |  | t value alpha=0.05 | 2.92 | 4.303
6 | 555 | large biofilm mass | [12 empty] | 0.2375 |  |  |  | P=0.007 | P=0.013
Mean (cells/FOV) | 271.3333 |  | 167.4 |  | 257.2 |  | 816.8 |  | 1064.8 |  | 1048.6
StDev (cells/FOV) | 154.9886 |  | 73.7855 |  | 214.1359 |  | 269.5676 |  | 258.8227 |  | 538.2525
[5 empty] | Mean | 231.9778 |  |  |  |  | Mean | 976.7333
[5 empty] | StDev | 56.3707 |  |  |  |  | StDev | 138.743
Mean (cells/cm2) | 3503998.6667 |  | 2161803.6 |  | 3321480.8 |  | 10548155.2 |  | 13750827.2 |  | 13541620.4
Stdev (cells/cm2) | 2001522.8028 |  | 952865.9461 |  | 2765351.505 |  | 3481196.2027 |  | 3342436.5964 |  | 6950993.3398
[5 empty] | Mean | 2995761.0222 |  |  |  |  | Mean | 12613534.2667
[5 empty] | StDev | 727971.1048 |  |  |  |  | StDev | 1791726.7895
[5 empty] | Survival fraction | 0.2375

## MB evaluation of catheters
CFU counts of Proteus Mirabilis from catheter pieces in an invitro study comparing LGC to standard catheter
|  |  | Separate segments CFU | [9 empty] | Whole catheter. Segment 1+ Segment 2 , CFU
Treatment | Catheter | Segment | Intensity | Plating Dilution | Plate 1 CFU | Plate 2 CFU | Plate 3 CFU | Mean | Std | Num Cells | StD | CV | Plate 1 CFU | Plate 2 CFU | Plate 3 CFU | Mean | Std | CV | SEM | Means | Std | t
No light control | 1 | 1 | 0 | 100 | 8 | 10 | 11 | 9.6667 | 1.5275 | 193333.3333 | 30550.5046 | 15.802
No light control | 1 | 2 | 0 | 100 | 21 | 20 | 23 | 21.3333 | 1.5275 | 426666.6667 | 30550.5046 | 7.1603 | 580000 | 600000 | 680000 | 620000 | 52915.0262 | 8.5347 | 30550.5046
No light control | 2 | 1 | 0 | 100 | 5 | 4 | 10 | 6.3333 | 3.2146 | 126666.6667 | 64291.0051 | 50.7561
No light control | 2 | 2 | 0 | 100 | 24 | 26 | 18 | 22.6667 | 4.1633 | 453333.3333 | 83266.64 | 18.3676 | 580000 | 600000 | 560000 | 580000 | 20000 | 3.4483 | 11547.0054
No light control | 3 | 1 | 0 | 100 | 22 | 16 | 11 | 16.3333 | 5.5076 | 326666.6667 | 110151.4109 | 33.7198
No light control | 3 | 2 | 0 | 100 | 3 | 3 | 1 | 2.3333 | 1.1547 | 46666.6667 | 23094.0108 | 49.4872 | 500000 | 380000 | 240000 | 373333.3333 | 130128.142 | 34.8558 | 75129.5178
No light control | 4 | 1 | 0 | 100 | 25 | 20 | 25 | 23.3333 | 2.8868 | 466666.6667 | 57735.0269 | 12.3718
No light control | 4 | 2 | 0 | 100 | 24 | 16 | 24 | 21.3333 | 4.6188 | 426666.6667 | 92376.0431 | 21.6506 | 980000 | 720000 | 980000 | 893333.3333 | 150111.07 | 16.8035 | 86666.6667
No light control | 5 | 1 | 0 | 100 | 6 | 12 | 12 | 10 | 3.4641 | 200000 | 69282.0323 | 34.641 | [7 empty] | X2 | S2
No light control | 5 | 2 | 0 | 100 | 8 | 7 | 4 | 6.3333 | 2.0817 | 126666.6667 | 41633.32 | 32.8684 | 280000 | 380000 | 320000 | 326666.6667 | 50332.2296 | 15.4078 | 29059.3263 | 558666.6667 | 226048.1758
Blue light | 1 | 1 | 17.0573 | 1 | 11 | 18 | 14 | 14.3333 | 3.5119 | 2866.6667 | 702.3769 | 24.5015
Blue light | 1 | 2 | 3.0603 | 1 | 9 | 7 | 8 | 8 | 1 | 1600 | 200 | 12.5 | 4000 | 5000 | 4400 | 4466.6667 | 503.3223 | 11.2684 | 290.5933
Blue light | 2 | 1 | 19.782 | 1 | 0 | 4 | 3 | 2.3333 | 2.0817 | 466.6667 | 416.3332 | 89.2143
Blue light | 2 | 2 | 3.5492 | 1 | 18 | 22 | 10 | 16.6667 | 6.1101 | 3333.3333 | 1222.0202 | 36.6606 | 3600 | 5200 | 2600 | 3800 | 1311.4877 | 34.5128 | 757.1878
Blue light | 3 | 1 | 16.796 | 1 | 10 | 44 | 50 | 34.6667 | 21.5716 | 6933.3333 | 4314.3172 | 62.2257 | [7 empty] | X1 | S1 | Sdelta | X-u | t | 1 | 2 | 3 | df | p
Blue light | 3 | 2 | 1.1124 | 1 | 100 | 108 | 100 | 102.6667 | 4.6188 | 20533.3333 | 923.7604 | 4.4988 | 22000 | 30400 | 30000 | 27466.6667 | 4738.4948 | 17.2518 | 2735.7713 | 11911.1111 | 13475.6296 | 101390.761 | -546755.5556 | -5.3926 | 1.0568e+20 | 1.832e+15 | 2.611e+19 | 4.0472
Welchs T test
| Control Catheter (X1) | Light guiding Catheter (X2) | X2-X1 | t value | Degrees freedom | P 1 tailed | P 2 tailed | 95% Confidence Interval of difference Lower | 95% Confidence interval of difference Upper
Mean (CFU) | 558666.6667 | 11911.1111 | -546755.5556 | -5.3926 | 4.0472 | <0.005 | <0.01 | -828216.3081 | -265294.803
Stdev (CFU) | 226048.1758 | 13475.6296 |  |  |  |  | Kill factor>> | Infinity | 1.9043
SEM | 101091.8175 | 7780.1584 | 101390.761
Technical replicates | 3 | 3
Biological replicates | 5 | 3 |  |  | t value alpha=0.05 | 2.132 | 2.776
[6 empty] | P=0.003 | P=0.006

## Power metre characterisation
Adjacent (cm) | 20.32
| Absolute power (uW) | [7 empty] | Relative power (%)
Angle | Opposite (cm) | No filter 1 | No filter 2 | No filter 3 | OD3 1 | OD3 2 | OD 3 3 | Angle | No filter 1 | No filter 2 | No filter 3 | No filter mean | No filter SD | OD3 1 | OD3 2 | OD 3 3 | OD3 filter Mean | OD3 filter SD | Costheta | Filter Path (mm) | Transmission
0 | 0 | 437.5 | 1202 | 491.1 | 291.1 | 293.7 | 294.6 | 0 | 100 | 100 | 100 | 100 | 0 | 100 | 100 | 100 | 100 | 0 | 100 | 1.41 | 100
5 | 1.7733 | 436 | 1200 | 489.9 | 292 | 293.5 | 294.7 | 5 | 99.6571 | 99.8336 | 99.7557 | 99.7488 | 0.088433 | 100.3092 | 99.9319 | 100.0339 | 100.0917 | 0.19515 | 99.6195 | 1.4154 | 99.6195
10 | 3.5465 | 438.4 | 1203 | 493.4 | 288.6 | 290.5 | 292.1 | 10 | 100.2057 | 100.0832 | 100.4683 | 100.2524 | 0.19677 | 99.1412 | 98.9105 | 99.1514 | 99.0677 | 0.13626 | 98.4808 | 1.4318 | 98.4808
15 | 5.3198 | 437.4 | 1205 | 494.3 | 283.3 | 283.3 | 286.4 | 15 | 99.9771 | 100.2496 | 100.6516 | 100.2928 | 0.3393 | 97.3205 | 96.459 | 97.2166 | 96.9987 | 0.47028 | 96.5926 | 1.4597 | 96.5926
20 | 7.093 | 439.1 | 1205 | 492.9 | 274.2 | 271.8 | 273.8 | 20 | 100.3657 | 100.2496 | 100.3665 | 100.3273 | 0.067283 | 94.1944 | 92.5434 | 92.9396 | 93.2258 | 0.86193 | 93.9693 | 1.5005 | 93.9693
25 | 8.8663 | 437 | 1208 | 490.1 | 266.6 | 261.8 | 264.5 | 25 | 99.8857 | 100.4992 | 99.7964 | 100.0604 | 0.38258 | 91.5836 | 89.1386 | 89.7828 | 90.1683 | 1.2673 | 90.6308 | 1.5558 | 90.6308
30 | 10.6395 | 438.2 | 1209 | 492.5 | 246.1 | 248.3 | 251 | 30 | 100.16 | 100.5824 | 100.2851 | 100.3425 | 0.21695 | 84.5414 | 84.542 | 85.2003 | 84.7612 | 0.38021 | 86.6025 | 1.6281 | 86.6025
35 | 12.4128 | 433.7 | 1213 | 495.3 | 230.6 | 239.1 | 238.8 | 35 | 99.1314 | 100.9151 | 100.8552 | 100.3006 | 1.013 | 79.2168 | 81.4096 | 81.0591 | 80.5618 | 1.178 | 81.9152 | 1.7213 | 81.9152
40 | 14.186 | 434.2 | 1219 | 497.3 | 221 | 223.8 | 225.3 | 40 | 99.2457 | 101.4143 | 101.2625 | 100.6408 | 1.2106 | 75.9189 | 76.2002 | 76.4766 | 76.1986 | 0.27883 | 76.6044 | 1.8406 | 76.6044

## Correction factor
Repeat measurements Direct power measurement with OD3 ND filter vs Integrating sphere measurements.
| 918 OD3 Power (mW) | 818 sphere Power (No multiplication) (uW) | Integration Sphere Power (After multiplication) (mW) | Correction Factor | Sphere Multiplication factor | 53.48
1 | 29.22 | 668.3 | 35.7423 | 0.81752
2 | 31.62 | 668.8 | 35.769 | 0.88401
3 | 30.32 | 649.2 | 34.7207 | 0.87325
4 | 31.3 | 646.2 | 34.5603 | 0.90566
5 | 31.01 | 654 | 34.9775 | 0.88657
Mean correction factor | 30.694 |  | 35.154 | 0.8734
SD | 0.95356 |  | 0.56913 | 0.033352