## Calibration Front Sheet
Calculations used in the code:
float get_voltage(int raw_adc) {
return raw_adc * (AREF / ADC_RESOLUTION);
}
float get_temperature(float voltage) {
return (voltage - 1.25) / 0.005;
}
float get_pressure(float voltage) {
return ((2 * voltage) - 1);
}
Pressure calibration data was recorded for varying sub-atmosheric pressures.
This data is available in the separate calibration raw data Excel document.
Further atmospheric calibration data is taken from actual heat recovery rig tests for further analysis.
Temperature calibration data was also taken, available in the raw data document.
Between 10 and 20 data points were used to take averages.
This was chosen as sufficient to give good average values to proeprly understand the system accuracy.
Greater data points were limited by the method of data reporting making analysis time-consuming.

## Pressure Calibration
Calibration Data: -0.8 barg test 1 |  |  |  | Calibration Data: -0.8 barg test 2 |  |  |  | Calibration Data: -0.8 barg test 3
| T1 | T2 | T3 |  | T1 | T2 | T3 |  | T1 | T2 | T3 | Pressure (barg) | [7 empty] | Pressure Difference (barg)
1 | -0.78 | -0.79 | -0.79 | 1 | -0.79 | -0.8 | -0.79 | 1 | -0.8 | -0.79 | -0.79 | T1 | T2 | T3 | Analogue/Known | Transducer manufacturer given accuracy: | 0.0025 | bar |  | T1 | T2 | T3 | Analogue/Known
2 | -0.81 | -0.8 | -0.79 | 2 | -0.8 | -0.8 | -0.79 | 2 | -0.79 | -0.79 | -0.79 | -0.219 | -0.223 | -0.205 | -0.2 | Analogue gauge precision: | 0.01 | bar |  | -0.019 | -0.023 | -0.005 | -0.2
3 | -0.82 | -0.8 | -0.79 | 3 | -0.8 | -0.81 | -0.79 | 3 | -0.81 | -0.8 | -0.79 | -0.51017 | -0.512 | -0.496 | -0.5 |  |  |  |  | -0.010167 | -0.012 | 0.004 | -0.5
4 | -0.8 | -0.8 | -0.79 | 4 | -0.79 | -0.8 | -0.79 | 4 | -0.81 | -0.8 | -0.79 | -0.8 | -0.80083 | -0.78967 | -0.8 |  |  |  |  | 0 | -0.00083333 | 0.010333 | -0.8
5 | -0.79 | -0.8 | -0.79 | 5 | -0.8 | -0.8 | -0.79 | 5 | -0.79 | -0.8 | -0.78 | [7 empty] | avg | -0.0097222 | -0.011944 | 0.0031111
6 | -0.8 | -0.8 | -0.79 | 6 | -0.79 | -0.8 | -0.79 | 6 | -0.8 | -0.81 | -0.79
7 | -0.8 | -0.8 | -0.79 | 7 | -0.8 | -0.81 | -0.79 | 7 | -0.8 | -0.81 | -0.79
8 | -0.79 | -0.8 | -0.79 | 8 | -0.81 | -0.81 | -0.79 | 8 | -0.8 | -0.81 | -0.79
9 | -0.81 | -0.8 | -0.78 | 9 | -0.8 | -0.79 | -0.79 | 9 | -0.8 | -0.8 | -0.79
10 | -0.81 | -0.8 | -0.79 | 10 | -0.82 | -0.8 | -0.79 | 10 | -0.8 | -0.8 | -0.79
11 | -0.8 | -0.81 | -0.8 | 11 | -0.79 | -0.8 | -0.79 | 11 | -0.82 | -0.81 | -0.79
12 | -0.79 | -0.8 | -0.79 | 12 | -0.81 | -0.8 | -0.79 | 12 | -0.8 | -0.79 | -0.79
13 | -0.8 | -0.8 | -0.79 | 13 | -0.81 | -0.79 | -0.79 | 13 | -0.79 | -0.81 | -0.79
14 | -0.79 | -0.8 | -0.79 | 14 | -0.79 | -0.8 | -0.79 | 14 | -0.8 | -0.8 | -0.79
15 | -0.79 | -0.81 | -0.79 | 15 | -0.79 | -0.8 | -0.79 | 15 | -0.79 | -0.81 | -0.79
16 | -0.8 | -0.8 | -0.78 | 16 | -0.8 | -0.79 | -0.79 | 16 | -0.8 | -0.8 | -0.79
17 | -0.8 | -0.81 | -0.79 | 17 | -0.79 | -0.8 | -0.79 | 17 | -0.82 | -0.8 | -0.79
18 | -0.8 | -0.81 | -0.79 | 18 | -0.8 | -0.8 | -0.79 | 18 | -0.81 | -0.8 | -0.79
19 | -0.79 | -0.79 | -0.79 | 19 | -0.81 | -0.8 | -0.79 | 19 | -0.8 | -0.8 | -0.79
20 | -0.8 | -0.81 | -0.79 | 20 | -0.8 | -0.8 | -0.79 | 20 | -0.81 | -0.79 | -0.79
AVERAGE | -0.7985 | -0.8015 | -0.7895 | AVERAGE | -0.7995 | -0.8 | -0.79 | AVERAGE | -0.802 | -0.801 | -0.7895
Calibration Data: -0.5 barg test 1 |  |  |  | Calibration Data: -0.5 barg test 2 |  |  |  | Calibration Data: -0.5 barg test 3
| T1 | T2 | T3 |  | T1 | T2 | T3 |  | T1 | T2 | T3
1 | -0.51 | -0.52 | -0.51 | 1 | -0.5 | -0.52 | -0.49 | 1 | -0.51 | -0.51 | -0.49
2 | -0.5 | -0.51 | -0.5 | 2 | -0.51 | -0.52 | -0.5 | 2 | -0.49 | -0.5 | -0.5
3 | -0.51 | -0.51 | -0.5 | 3 | -0.51 | -0.51 | -0.49 | 3 | -0.5 | -0.51 | -0.5
4 | -0.51 | -0.51 | -0.5 | 4 | -0.52 | -0.51 | -0.5 | 4 | -0.49 | -0.52 | -0.49
5 | -0.52 | -0.52 | -0.49 | 5 | -0.5 | -0.5 | -0.5 | 5 | -0.52 | -0.51 | -0.5
6 | -0.5 | -0.51 | -0.49 | 6 | -0.5 | -0.51 | -0.5 | 6 | -0.53 | -0.51 | -0.49
7 | -0.51 | -0.52 | -0.5 | 7 | -0.51 | -0.51 | -0.5 | 7 | -0.49 | -0.52 | -0.49
8 | -0.52 | -0.5 | -0.49 | 8 | -0.53 | -0.51 | -0.48 | 8 | -0.52 | -0.51 | -0.5
9 | -0.51 | -0.52 | -0.5 | 9 | -0.5 | -0.51 | -0.49 | 9 | -0.52 | -0.51 | -0.49
10 | -0.5 | -0.52 | -0.5 | 10 | -0.49 | -0.52 | -0.5 | 10 | -0.51 | -0.51 | -0.5
11 | -0.51 | -0.51 | -0.5 | 11 | -0.51 | -0.51 | -0.49 | 11 | -0.52 | -0.51 | -0.5
12 | -0.51 | -0.51 | -0.5 | 12 | -0.51 | -0.5 | -0.49 | 12 | -0.51 | -0.52 | -0.5
13 | -0.52 | -0.51 | -0.51 | 13 | -0.5 | -0.51 | -0.5 | 13 | -0.52 | -0.51 | -0.49
14 | -0.51 | -0.51 | -0.49 | 14 | -0.52 | -0.51 | -0.5 | 14 | -0.53 | -0.51 | -0.49
15 | -0.51 | -0.51 | -0.49 | 15 | -0.51 | -0.52 | -0.5 | 15 | -0.51 | -0.51 | -0.5
16 | -0.5 | -0.51 | -0.49 | 16 | -0.51 | -0.52 | -0.51 | 16 | -0.51 | -0.51 | -0.5
17 | -0.52 | -0.51 | -0.49 | 17 | -0.52 | -0.52 | -0.5 | 17 | -0.51 | -0.51 | -0.49
18 | -0.5 | -0.5 | -0.49 | 18 | -0.54 | -0.51 | -0.51 | 18 | -0.53 | -0.51 | -0.49
19 | -0.51 | -0.52 | -0.49 | 19 | -0.49 | -0.52 | -0.49 | 19 | -0.51 | -0.51 | -0.48
20 | -0.51 | -0.51 | -0.5 | 20 | -0.51 | -0.52 | -0.5 | 20 | -0.5 | -0.51 | -0.5
AVERAGE | -0.5095 | -0.512 | -0.4965 | AVERAGE | -0.5095 | -0.513 | -0.497 | AVERAGE | -0.5115 | -0.511 | -0.4945
Calibration Data: -0.2 barg test 1 |  |  |  | Calibration Data: -0.2 barg test 2 |  |  |  | Calibration Data: -0.2 barg test 3
| T1 | T2 | T3 |  | T1 | T2 | T3 |  | T1 | T2 | T3
1 | -0.21 | -0.22 | -0.2 | 1 | -0.23 | -0.22 | -0.21 | 1 | -0.2 | -0.22 | -0.2
2 | -0.22 | -0.22 | -0.2 | 2 | -0.22 | -0.22 | -0.2 | 2 | -0.2 | -0.23 | -0.2
3 | -0.23 | -0.23 | -0.2 | 3 | -0.23 | -0.22 | -0.2 | 3 | -0.23 | -0.23 | -0.21
4 | -0.22 | -0.22 | -0.2 | 4 | -0.22 | -0.22 | -0.2 | 4 | -0.24 | -0.23 | -0.21
5 | -0.22 | -0.22 | -0.21 | 5 | -0.21 | -0.23 | -0.19 | 5 | -0.21 | -0.22 | -0.21
6 | -0.21 | -0.23 | -0.22 | 6 | -0.21 | -0.22 | -0.2 | 6 | -0.2 | -0.21 | -0.21
7 | -0.21 | -0.22 | -0.21 | 7 | -0.21 | -0.21 | -0.2 | 7 | -0.21 | -0.24 | -0.22
8 | -0.22 | -0.23 | -0.21 | 8 | -0.23 | -0.21 | -0.21 | 8 | -0.22 | -0.22 | -0.2
9 | -0.23 | -0.23 | -0.22 | 9 | -0.22 | -0.22 | -0.19 | 9 | -0.23 | -0.23 | -0.2
10 | -0.21 | -0.23 | -0.19 | 10 | -0.24 | -0.23 | -0.2 | 10 | -0.23 | -0.22 | -0.2
11 | -0.24 | -0.22 | -0.2 | 11 | -0.24 | -0.22 | -0.21 | 11 | -0.21 | -0.23 | -0.2
12 | -0.22 | -0.22 | -0.2 | 12 | -0.22 | -0.23 | -0.21 | 12 | -0.22 | -0.23 | -0.2
13 | -0.21 | -0.22 | -0.21 | 13 | -0.2 | -0.23 | -0.2 | 13 | -0.2 | -0.22 | -0.22
14 | -0.23 | -0.22 | -0.21 | 14 | -0.22 | -0.22 | -0.21 | 14 | -0.23 | -0.23 | -0.21
15 | -0.21 | -0.23 | -0.23 | 15 | -0.24 | -0.22 | -0.21 | 15 | -0.22 | -0.21 | -0.21
16 | -0.22 | -0.21 | -0.22 | 16 | -0.23 | -0.22 | -0.19 | 16 | -0.21 | -0.23 | -0.2
17 | -0.21 | -0.22 | -0.2 | 17 | -0.22 | -0.22 | -0.21 | 17 | -0.22 | -0.22 | -0.21
18 | -0.22 | -0.23 | -0.21 | 18 | -0.21 | -0.23 | -0.2 | 18 | -0.23 | -0.22 | -0.2
19 | -0.22 | -0.22 | -0.2 | 19 | -0.21 | -0.22 | -0.2 | 19 | -0.19 | -0.22 | -0.22
20 | -0.22 | -0.22 | -0.19 | 20 | -0.23 | -0.23 | -0.2 | 20 | -0.22 | -0.22 | -0.2
AVERAGE | -0.219 | -0.223 | -0.2065 | AVERAGE | -0.222 | -0.222 | -0.202 | AVERAGE | -0.216 | -0.224 | -0.2065

## Temperature Calibration
[13 empty] | Thermocouple Calibration: |  |  | Thermocouple accuracy: | 1.5 | degC
Thermocouple 1: | [15 empty] | Conditioner accuracy: | 2 | degC
[16 empty] | Combined accuracy: | 2.5 | degC
| Test 1 |  |  |  | Test 2 |  |  |  | Test 3 |  |  |  | TC1 | Heater | Difference
| Repeat 1 | Repeat 2 | Repeat 3 |  | Repeat 1 | Repeat 2 | Repeat 3 |  | Repeat 1 | Repeat 2 | Repeat 3 |  | 49.1533 | 50.0667 | -0.91333 | Sous vide heater precision: | 0.05 | degC
1 | 49.8 | 49.8 | 48.83 | 1 | 59.57 | 58.59 | 58.59 | 1 | 69.34 | 69.34 | 69.34 |  | 58.6887 | 60.0333 | -1.3447 | Thermometer precision: | 1 | degC
2 | 48.83 | 49.8 | 48.83 | 2 | 58.59 | 58.59 | 58.59 | 2 | 69.34 | 68.36 | 68.36 |  | 68.7843 | 70 | -1.2157
3 | 48.83 | 48.83 | 48.83 | 3 | 59.57 | 58.59 | 59.57 | 3 | 69.34 | 68.36 | 68.36 |  |  | avg | -1.1579
4 | 48.83 | 49.8 | 48.83 | 4 | 59.57 | 58.59 | 58.59 | 4 | 69.34 | 68.36 | 68.36
5 | 48.83 | 48.83 | 48.83 | 5 | 58.59 | 58.59 | 58.59 | 5 | 69.34 | 69.34 | 68.36 |  | TC2 | Heater | Difference
6 | 48.83 | 49.8 | 49.8 | 6 | 58.59 | 58.59 | 58.59 | 6 | 69.34 | 68.36 | 68.36 |  | 23.3407 | 25.3 | -1.9593
7 | 48.83 | 49.8 | 49.8 | 7 | 58.59 | 58.59 | 59.57 | 7 | 69.34 | 68.36 | 68.36 |  | 33.3327 | 35.1333 | -1.8007
8 | 48.83 | 49.8 | 48.83 | 8 | 58.59 | 58.59 | 58.59 | 8 | 70.31 | 68.36 | 68.36 |  | 43.3943 | 45.1333 | -1.739
9 | 48.83 | 49.8 | 48.83 | 9 | 58.59 | 58.59 | 57.62 | 9 | 69.34 | 68.36 | 68.36 |  |  | avg | -1.833
10 | 48.83 | 48.83 | 48.83 | 10 | 58.59 | 58.59 | 57.62 | 10 | 68.36 | 68.36 | 68.36
AVERAGE | 48.927 | 49.509 | 49.024 | AVERAGE | 58.884 | 58.59 | 58.592 | AVERAGE | 69.339 | 68.556 | 68.458
Thermometer | 48 | 47 | 46 | Thermometer | 57 | 57 | 56 | Thermometer | 67 | 66 | 67
Heater | 50.2 | 50.1 | 49.9 | Heater | 60.1 | 60.1 | 59.9 | Heater | 70.1 | 70 | 69.9
Thermocouple 2:
| Test 1 |  |  |  | Test 2 |  |  |  | Test 3
| Repeat 1 | Repeat 2 | Repeat 3 |  | Repeat 1 | Repeat 2 | Repeat 3 |  | Repeat 1 | Repeat 2 | Repeat 3
1 | 23.44 | 22.46 | 23.44 | 1 | 35.16 | 31.25 | 33.2 | 1 | 43.95 | 42.97 | 43.95
2 | 22.46 | 23.44 | 23.44 | 2 | 34.18 | 32.23 | 32.23 | 2 | 41.99 | 43.95 | 43.95
3 | 24.41 | 24.41 | 22.46 | 3 | 34.18 | 33.2 | 33.2 | 3 | 42.97 | 42.97 | 42.97
4 | 23.44 | 23.44 | 24.41 | 4 | 33.2 | 32.23 | 32.23 | 4 | 42.97 | 43.95 | 45.9
5 | 22.46 | 23.44 | 23.44 | 5 | 33.2 | 33.2 | 33.2 | 5 | 43.95 | 42.97 | 43.95
6 | 23.44 | 23.44 | 22.46 | 6 | 32.23 | 33.2 | 33.2 | 6 | 42.97 | 41.99 | 43.95
7 | 23.44 | 23.44 | 23.44 | 7 | 33.2 | 33.2 | 33.2 | 7 | 42.97 | 43.95 | 42.97
8 | 22.46 | 24.41 | 23.44 | 8 | 37.11 | 34.18 | 32.23 | 8 | 42.97 | 42.97 | 43.95
9 | 23.44 | 23.44 | 22.46 | 9 | 33.2 | 33.2 | 35.16 | 9 | 42.97 | 42.97 | 42.97
10 | 23.44 | 23.44 | 23.44 | 10 | 34.18 | 33.2 | 33.2 | 10 | 42.97 | 43.95 | 43.95
AVERAGE | 23.243 | 23.536 | 23.243 | AVERAGE | 33.984 | 32.909 | 33.105 | AVERAGE | 43.068 | 43.264 | 43.851
Thermometer | 24 | 24 | 24 | Thermometer | 34 | 33 | 34 | Thermometer | 43 | 44 | 43
Heater | 25.3 | 25.3 | 25.3 | Heater | 35.1 | 35.2 | 35.1 | Heater | 45.2 | 45.1 | 45.1
20
19
18
17
16
15
14
13
12
11
10
9
8
7
6
5
4
3
2
1