## Load Cell _ Calibrated W's
Calibration Test 1 |  | Test 2 |  | Test 3 |  | Test 4 |  | Test 5 |  |  |  | Average Recorded Weight (kg)
TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) |  | Weight (kg) | Test 1 | Test 2 | Test 3 | Test 4 | Test 5 | AVERAGE | SU (dev from mean) | k = | 95% Confidence in variability
838 | 0.002 | 825 | 0.002 | 796 | 0.002 | 780 | 0.002 | 806 | 0.002 | Phase 1 | 0 | 0.0023968 | 0.0013953 | 0.0021468 | 0.0019725 | 0.0020799 | 0.0019983 | 0.00016609 | 2 | 0.00033219
1021 | 0.002 | 1007 | 0.002 | 978 | 0.002 | 963 | 0.002 | 990 | 0.002 | Phase 2 | 0.2 | 0.20181 | 0.20115 | 0.20156 | 0.20209 | 0.20247 | 0.20182 | 0.00022563 | 2 | 0.00045126
1138 | 0.002 | 1125 | 0.002 | 1096 | 0.002 | 1080 | 0.002 | 1106 | 0.002 | Phase 3 | 0 | 0.0020394 | 0.0011499 | 0.0020217 | 0.0020495 | 0.0023564 | 0.0019234 | 0.00020308 | 2 | 0.00040616
1256 | 0.002 | 1243 | 0.002 | 1214 | 0.002 | 1198 | 0.002 | 1224 | 0.002 | Phase 4 | 0.4 | 0.40184 | 0.40099 | 0.40241 | 0.40207 | 0.40248 | 0.40196 | 0.0002686 | 2 | 0.0005372
1373 | 0.002 | 1359 | 0.002 | 1331 | 0.002 | 1314 | 0.002 | 1341 | 0.002 | Phase 5 | 0 | 0.0014565 | 0.00099206 | 0.0014832 | 0.0018422 | 0.002004 | 0.0015556 | 0.00017547 | 2 | 0.00035094
1490 | 0.002 | 1477 | 0.002 | 1448 | 0.002 | 1432 | 0.002 | 1459 | 0.002 | Phase 6 | 1 | 1.0017 | 1.0007 | 1.0027 | 1.002 | 1.0017 | 1.0018 | 0.00031259 | 2 | 0.00062518
1607 | 0.002 | 1594 | 0.002 | 1565 | 0.002 | 1549 | 0.002 | 1575 | 0.002 | Phase 7 | 0 | 0.0022285 | 0.0010417 | 0.002041 | 0.0021813 | 0.0020276 | 0.001904 | 0.00021907 | 2 | 0.00043814
1725 | 0.002 | 1712 | 0.002 | 1683 | 0.002 | 1667 | 0.002 | 1693 | 0.002
1842 | 0.002 | 1828 | 0.002 | 1800 | 0.002 | 1783 | 0.002 | 1810 | 0.002
1959 | 0.002 | 1946 | 0.002 | 1917 | 0.002 | 1901 | 0.002 | 1928 | 0.002 |  |  | Offset (kg)
2076 | 0.002 | 2063 | 0.002 | 2034 | 0.002 | 2019 | 0.002 | 2044 | 0.002 |  | Weight (kg) | Test 1 | Test 2 | Test 3 | Test 4 | Test 5
2194 | 0.002 | 2181 | 0.002 | 2152 | 0.002 | 2136 | 0.002 | 2162 | 0.002 | Phase 1 | 0 | 0.0023968 | 0.0013953 | 0.0021468 | 0.0019725 | 0.0020799
2311 | 0.002 | 2297 | 0.002 | 2269 | 0.002 | 2253 | 0.002 | 2280 | 0.002 | Phase 2 | 0.2 | 0.0018091 | 0.0011477 | 0.0015621 | 0.0020947 | 0.0024713
2428 | 0.002 | 2415 | 0.002 | 2386 | 0.002 | 2370 | 0.002 | 2397 | 0.002 | Phase 3 | 0 | 0.0020394 | 0.0011499 | 0.0020217 | 0.0020495 | 0.0023564
2545 | 0.002 | 2532 | 0.002 | 2503 | 0.002 | 2488 | 0.002 | 2514 | 0.002 | Phase 4 | 0.4 | 0.0018413 | 0.0009901 | 0.0024147 | 0.002073 | 0.0024783
2663 | 0.002 | 2650 | 0.002 | 2621 | 0.002 | 2605 | 0.002 | 2631 | 0.002 | Phase 5 | 0 | 0.0014565 | 0.00099206 | 0.0014832 | 0.0018422 | 0.002004
2781 | 0.002 | 2766 | 0.002 | 2742 | 0.002 | 2722 | 0.002 | 2749 | 0.002 | Phase 6 | 1 | 0.0016909 | 0.00072098 | 0.0026598 | 0.002004 | 0.0017339
2897 | 0.002 | 2884 | 0.002 | 2859 | 0.002 | 2839 | 0.002 | 2866 | 0.002 | Phase 7 | 0 | 0.0022285 | 0.0010417 | 0.002041 | 0.0021813 | 0.0020276
3015 | 0.002 | 3001 | 0.002 | 2976 | 0.002 | 2957 | 0.002 | 2983 | 0.002
3135 | 0.002 | 3119 | 0.002 | 3093 | 0.002 | 3074 | 0.002 | 3103 | 0.002
3253 | 0.002 | 3235 | 0.002 | 3211 | 0.002 | 3191 | 0.002 | 3221 | 0.002 |  | Weight (kg) | Average Load Cell (kg)
3371 | 0.002 | 3353 | 0.002 | 3329 | 0.002 | 3308 | 0.002 | 3339 | 0.002 |  | 0 | 0.0018453
3487 | 0.002 | 3470 | 0.002 | 3445 | 0.002 | 3429 | 0.002 | 3456 | 0.002 |  | 0.2 | 0.20182
3605 | 0.002 | 3591 | 0.002 | 3563 | 0.002 | 3547 | 0.002 | 3573 | 0.002 |  | 0.4 | 0.40196
3722 | 0.002 | 3708 | 0.002 | 3681 | 0.002 | 3663 | 0.002 | 3691 | 0.002 |  | 1 | 1.0018
3840 | 0.002 | 3826 | 0.002 | 3798 | 0.002 | 3781 | 0.002 | 3808 | 0.002 |  | Offset | 0.0018459
3957 | 0.002 | 3943 | 0.002 | 3915 | 0.002 | 3899 | 0.002 | 3926 | 0.002
4074 | 0.002 | 4061 | 0.002 | 4032 | 0.002 | 4016 | 0.002 | 4043 | 0.002
4192 | 0.002 | 4177 | 0.002 | 4150 | 0.002 | 4133 | 0.002 | 4160 | 0.002 |  | Combined accuracy |  | 0.0018997 | kg
4310 | 0.002 | 4295 | 0.002 | 4268 | 0.002 | 4251 | 0.002 | 4278 | 0.002
4426 | 0.002 | 4413 | 0.002 | 4384 | 0.002 | 4368 | 0.002 | 4395 | 0.002
4544 | 0.002 | 4530 | 0.002 | 4502 | 0.002 | 4486 | 0.002 | 4512 | 0.002
4661 | 0.002 | 4647 | 0.002 | 4620 | 0.002 | 4602 | 0.002 | 4630 | 0.002
4779 | 0.002 | 4765 | 0.002 | 4737 | 0.002 | 4720 | 0.002 | 4747 | 0.002
4896 | 0.002 | 4882 | 0.002 | 4854 | 0.002 | 4838 | 0.002 | 4865 | 0.002
5013 | 0.002 | 5000 | 0.002 | 4971 | 0.002 | 4955 | 0.002 | 4982 | 0.002
5131 | 0.002 | 5116 | 0.002 | 5089 | 0.002 | 5072 | 0.002 | 5099 | 0.002
5249 | 0.002 | 5234 | 0.002 | 5207 | 0.002 | 5190 | 0.002 | 5217 | 0.002
5365 | 0.002 | 5352 | 0.002 | 5323 | 0.002 | 5307 | 0.002 | 5334 | 0.002
5483 | 0.002 | 5469 | 0.002 | 5441 | 0.002 | 5425 | 0.002 | 5451 | 0.002
5600 | 0.002 | 5586 | 0.002 | 5559 | 0.002 | 5541 | 0.002 | 5569 | 0.002
5718 | 0.002 | 5704 | 0.002 | 5676 | 0.002 | 5659 | 0.002 | 5686 | 0.002
5835 | 0.002 | 5821 | 0.002 | 5793 | 0.003 | 5777 | 0.002 | 5804 | 0.002
5952 | 0.002 | 5939 | 0.002 | 5910 | 0.003 | 5894 | 0.002 | 5921 | 0.002
6070 | 0.002 | 6055 | 0.002 | 6028 | 0.003 | 6011 | 0.002 | 6038 | 0.002
6188 | 0.002 | 6173 | 0.002 | 6146 | 0.003 | 6129 | 0.002 | 6156 | 0.002
6304 | 0.002 | 6291 | 0.002 | 6262 | 0.002 | 6246 | 0.002 | 6273 | 0.002
6422 | 0.002 | 6408 | 0.002 | 6380 | 0.003 | 6364 | 0.002 | 6390 | 0.002
6539 | 0.002 | 6525 | 0.002 | 6498 | 0.003 | 6480 | 0.002 | 6508 | 0.002
6657 | 0.002 | 6643 | 0.002 | 6615 | 0.002 | 6598 | 0.002 | 6625 | 0.002
6774 | 0.002 | 6760 | 0.002 | 6732 | 0.002 | 6716 | 0.002 | 6743 | 0.002
6891 | 0.002 | 6878 | 0.002 | 6849 | 0.002 | 6833 | 0.002 | 6860 | 0.002
7009 | 0.002 | 6994 | 0.002 | 6967 | 0.003 | 6950 |  | 6977 | 0.002
7127 | 0.002 | 7112 | 0.002 | 7085 | 0.003 | 7068 |  | 7095 | 0.002
7243 | 0.002 | 7230 | 0.002 | 7201 | 0.003 | 7185 | 0.002 | 7213 | 0.002
7361 | 0.002 | 7347 | 0.002 | 7319 | 0.003 | 7303 | 0.002 | 7329 | 0.002
7479 | 0.002 | 7464 | 0.002 | 7437 | 0.003 | 7420 | 0.002 | 7447 | 0.002
7596 | 0.002 | 7582 | 0.002 | 7554 | 0.003 | 7537 | 0.002 | 7564 | 0.002
7713 | 0.002 | 7699 | 0.002 | 7671 | 0.003 | 7655 | 0.002 | 7682 | 0.002
7830 | 0.002 | 7817 | 0.002 | 7792 | 0.003 | 7773 | 0.002 | 7799 | 0.002
7948 | 0.002 | 7933 | 0.002 | 7909 | 0.003 | 7889 | 0.002 | 7916 | 0.002
8066 | 0.002 | 8051 | 0.002 | 8027 | 0.003 | 8007 | 0.002 | 8034 | 0.002
8186 | 0.002 | 8169 | 0.002 | 8144 | 0.003 | 8124 | 0.002 | 8155 | 0.002
8303 | 0.002 | 8286 | 0.002 | 8261 | 0.003 | 8242 | 0.002 | 8271 | 0.002
8421 | 0.002 | 8403 | 0.002 | 8379 | 0.003 | 8359 | 0.002 | 8389 | 0.002
8538 | 0.002 | 8521 | 0.002 | 8497 | 0.003 | 8480 | 0.002 | 8507 | 0.002
8655 | 0.002 | 8641 | 0.002 | 8613 | 0.003 | 8597 | 0.002 | 8624 | 0.002
8773 | 0.002 | 8759 | 0.002 | 8731 | 0.003 | 8715 | 0.002 | 8741 | 0.002
8890 | 0.002 | 8877 | 0.002 | 8848 | 0.003 | 8832 | 0.002 | 8859 | 0.002
9008 | 0.002 | 8993 | 0.002 | 8966 | 0.003 | 8949 | 0.002 | 8976 | 0.002
9125 | 0.002 | 9111 | 0.002 | 9083 | 0.003 | 9067 | 0.002 | 9094 | 0.002
9242 | 0.002 | 9229 | 0.002 | 9200 | 0.003 | 9184 | 0.002 | 9210 | 0.002
9360 | 0.002 | 9346 | 0.002 | 9318 | 0.003 | 9302 | 0.002 | 9328 | 0.003
9477 | 0.002 | 9463 | 0.002 | 9436 | 0.003 | 9419 | 0.002 | 9446 | 0.002
9594 | 0.002 | 9580 | 0.002 | 9552 | 0.003 | 9536 | 0.002 | 9563 | 0.003
9712 | 0.002 | 9698 | 0.002 | 9670 | 0.003 | 9654 | 0.002 | 9680 | 0.002
9829 | 0.002 | 9816 | 0.002 | 9787 | 0.003 | 9772 | 0.002 | 9798 | 0.003
9947 | 0.002 | 9932 | 0.002 | 9905 | 0.003 | 9888 | 0.002 | 9915 | 0.002
10064 | 0.002 | 10050 | 0.002 | 10022 | 0.003 | 10006 | 0.002 | 10033 | 0.002
10181 | 0.002 | 10168 | 0.002 | 10139 | 0.003 | 10124 | 0.002 | 10150 | 0.002
10299 | 0.002 | 10285 | 0.002 | 10257 | 0.003 | 10241 | 0.002 | 10267 | 0.002
10417 | 0.002 | 10402 | 0.002 | 10375 | 0.002 | 10358 | 0.002 | 10385 | 0.002
10533 | 0.002 | 10520 | 0.002 | 10491 | 0.003 | 10476 | 0.002 | 10503 | 0.002
10651 | 0.002 | 10637 | 0.002 | 10609 | 0.003 | 10593 | 0.002 | 10619 | 0.002
10769 | 0.002 | 10755 | 0.002 | 10727 | 0.003 | 10711 | 0.002 | 10737 | 0.002
10886 | 0.002 | 10872 | 0.002 | 10844 | 0.003 | 10828 | 0.002 | 10855 | 0.002
11003 | 0.002 | 10989 | 0.002 | 10961 | 0.003 | 10945 | 0.002 | 10972 | 0.002
11121 | 0.002 | 11107 | 0.002 | 11079 | 0.003 | 11063 | 0.002 | 11089 | 0.003
11238 | 0.002 | 11225 | 0.002 | 11196 | 0.003 | 11181 | 0.002 | 11207 | 0.003
11356 | 0.002 | 11341 | 0.002 | 11314 | 0.003 | 11297 | 0.002 | 11324 | 0.002
11473 | 0.002 | 11459 | 0.002 | 11431 | 0.003 | 11415 | 0.002 | 11442 | 0.003
11590 | 0.002 | 11577 | 0.002 | 11549 | 0.003 | 11533 | 0.002 | 11580 | 0.002
11708 | 0.002 | 11695 | 0.002 | 11684 | 0.003 | 11650 | 0.002 | 11698 | 0.002
11826 | 0.002 | 11811 | 0.002 | 11801 | 0.003 | 11767 | 0.002 | 11815 | 0.002
11942 | 0.002 | 11929 | 0.002 | 11919 | 0.003 | 11885 | 0.002 | 11932 | 0.002
12060 | 0.002 | 12047 | 0.002 | 12037 | 0.002 | 12002 | 0.002 | 12050 | 0.002
12178 | 0.002 | 12164 | 0.002 | 12153 | 0.003 | 12139 | 0.002 | 12168 | 0.002
12295 | 0.002 | 12281 | 0.002 | 12271 | 0.003 | 12256 | 0.002 | 12284 | 0.003
12428 | 0.002 | 12399 | 0.002 | 12389 | 0.003 | 12374 | 0.002 | 12402 | 0.002
12546 | 0.002 | 12516 | 0.002 | 12506 | 0.003 | 12491 | 0.002 | 12520 | 0.002
12662 | 0.002 | 12634 | 0.002 | 12626 | 0.002 | 12608 | 0.002 | 12637 | 0.003
12780 | 0.002 | 12768 | 0.002 | 12744 | 0.003 | 12726 | 0.002 | 12754 | 0.002
12898 | 0.002 | 12884 | 0.002 | 12862 | 0.003 | 12844 | 0.002 | 12872 | 0.002
13019 | 0.002 | 13002 | 0.002 | 12979 | 0.003 | 12960 | 0.002 | 12993 | 0.002
13135 | 0.002 | 13120 | 0.002 | 13096 | 0.002 | 13078 | 0.002 | 13110 | 0.002
13253 | 0.002 | 13237 | 0.002 | 13214 | 0.003 | 13199 | 0.002 | 13228 | 0.003
13371 | 0.002 | 13358 | 0.002 | 13331 | 0.003 | 13317 | 0.002 | 13345 | 0.002
13488 | 0.002 | 13475 | 0.002 | 13449 | 0.003 | 13433 | 0.003 | 13462 | 0.002
13605 | 0.002 | 13593 | 0.002 | 13566 | 0.003 | 13551 | 0.003 | 13580 | 0.002
13723 | 0.002 | 13710 | 0.002 | 13683 | 0.003 | 13669 | 0.003 | 13698 | 0.002
13841 | 0.002 | 13828 | 0.002 | 13801 | 0.003 | 13786 | 0.003 | 13814 | 0.002
13958 | 0.002 | 13945 | 0.002 | 13919 | 0.003 | 13903 | 0.003 | 13932 | 0.002
14075 | 0.002 | 14062 | 0.002 | 14035 | 0.003 | 14021 | 0.002 | 14050 | 0.002
14193 | 0.002 | 14180 | 0.002 | 14153 | 0.003 | 14138 | 0.003 | 14167 | 0.002
14310 | 0.002 | 14298 | 0.002 | 14271 | 0.003 | 14256 | 0.002 | 14284 | 0.002
14428 | 0.002 | 14414 | 0.002 | 14388 | 0.003 | 14373 | 0.002 | 14402 | 0.002
14545 | 0.002 | 14532 | 0.002 | 14505 | 0.003 | 14490 | 0.003 | 14519 | 0.002
14662 | 0.002 | 14650 | 0.002 | 14623 | 0.003 | 14608 | 0.003 | 14637 | 0.002
14780 | 0.002 | 14767 | 0.002 | 14740 | 0.002 | 14726 | 0.003 | 14754 | 0.002
14898 | 0.002 | 14884 | 0.002 | 14858 | 0.002 | 14842 | 0.003 | 14871 | 0.002
15014 | 0.002 | 15002 | 0.002 | 14976 | 0.003 | 14960 | 0.003 | 14989 | 0.002
15132 | 0.002 | 15119 | 0.002 | 15092 | 0.002 | 15078 | 0.003 | 15107 | 0.002
15250 | 0.002 | 15237 | 0.002 | 15210 | 0.002 | 15195 | 0.003 | 15224 | 0.002
15367 | 0.002 | 15354 | 0.002 | 15328 | 0.003 | 15312 | 0.003 | 15341 | 0.002
15484 | 0.002 | 15472 | 0.002 | 15446 | 0.003 | 15430 | 0.003 | 15459 | 0.002
15602 | 0.002 | 15589 | 0.002 | 15562 | 0.003 | 15547 | 0.003 | 15577 | 0.002
15719 | 0.002 | 15707 | 0.002 | 15680 | 0.002 | 15665 | 0.003 | 15693 | 0.002
15837 | 0.002 | 15824 | 0.002 | 15798 | 0.002 | 15782 | 0.003 | 15811 | 0.002
15954 | 0.002 | 15941 | 0.002 | 15915 | 0.002 | 15899 | 0.002 | 15929 | 0.002
16071 | 0.002 | 16059 | 0.002 | 16032 | 0.002 | 16017 | 0.002 | 16046 | 0.002
16189 | 0.002 | 16177 | 0.002 | 16150 | 0.002 | 16135 | 0.003 | 16163 | 0.002
16307 | 0.002 | 16293 | 0.002 | 16267 | 0.003 | 16251 | 0.003 | 16281 | 0.002
16423 | 0.002 | 16411 | 0.001 | 16385 | 0.002 | 16369 | 0.003 | 16398 | 0.002
16541 | 0.002 | 16529 | 0.001 | 16502 | 0.002 | 16487 | 0.003 | 16516 | 0.002
16659 | 0.002 | 16646 | 0.002 | 16619 | 0.002 | 16604 | 0.003 | 16633 | 0.002
16776 | 0.002 | 16763 | 0.002 | 16737 | 0.002 | 16721 | 0.002 | 16750 | 0.002
16893 | 0.002 | 16881 | 0.002 | 16855 | 0.002 | 16839 | 0.002 | 16868 | 0.002
17011 | 0.002 | 16998 | 0.002 | 16971 | 0.002 | 16956 | 0.003 | 16986 | 0.002
17128 | 0.002 | 17116 | 0.002 | 17089 | 0.002 | 17074 | 0.003 | 17102 | 0.002
17246 | 0.002 | 17233 | 0.002 | 17210 | 0.002 | 17191 | 0.003 | 17220 | 0.002
17363 | 0.002 | 17350 | 0.002 | 17328 | 0.003 | 17308 | 0.003 | 17338 | 0.002
17480 | 0.002 | 17468 | 0.002 | 17444 | 0.002 | 17426 | 0.002 | 17455 | 0.002
17601 | 0.002 | 17586 | 0.002 | 17562 | 0.002 | 17544 | 0.003 | 17575 | 0.002
17719 | 0.002 | 17702 | 0.002 | 17680 | 0.002 | 17660 | 0.002 | 17693 | 0.002
17837 | 0.002 | 17820 | 0.002 | 17797 | 0.002 | 17778 | 0.002 | 17811 | 0.002
17953 | 0.002 | 17938 | 0.002 | 17914 | 0.002 | 17899 | 0.002 | 17928 | 0.002
18071 | 0.002 | 18059 | 0.002 | 18032 | 0.002 | 18017 | 0.003 | 18045 | 0.002
18189 | 0.002 | 18176 | 0.002 | 18149 | 0.003 | 18134 | 0.003 | 18163 | 0.002
18306 | 0.002 | 18293 | 0.002 | 18267 | 0.002 | 18251 | 0.002 | 18280 | 0.002
18423 | 0.002 | 18411 | 0.002 | 18384 | 0.002 | 18369 | 0.002 | 18398 | 0.002
18541 | 0.002 | 18528 | 0.002 | 18501 | 0.002 | 18486 | 0.002 | 18515 | 0.002
18658 | 0.002 | 18646 | 0.002 | 18619 | 0.002 | 18604 | 0.002 | 18632 | 0.002
18776 | 0.002 | 18763 | 0.002 | 18737 | 0.002 | 18721 | 0.002 | 18750 | 0.002
18893 | 0.002 | 18881 | 0.002 | 18853 | 0.002 | 18838 | 0.002 | 18868 | 0.002
19011 | 0.002 | 18998 | 0.002 | 18971 | 0.002 | 18956 | 0.002 | 18984 | 0.002
19128 | 0.002 | 19116 | 0.002 | 19089 | 0.002 | 19074 | 0.002 | 19102 | 0.002
19246 | 0.002 | 19233 | 0.001 | 19206 | 0.002 | 19190 | 0.002 | 19220 | 0.002
19363 | 0.002 | 19350 | 0.002 | 19323 | 0.002 | 19308 | 0.002 | 19338 | 0.002
19480 | 0.002 | 19468 | 0.002 | 19441 | 0.002 | 19426 | 0.003 | 19454 | 0.002
19598 | 0.002 | 19586 | 0.002 | 19558 | 0.002 | 19543 | 0.002 | 19572 | 0.002
19716 | 0.002 | 19702 | 0.002 | 19676 | 0.002 | 19660 | 0.002 | 19690 | 0.002
19832 | 0.002 | 19820 | 0.002 | 19793 | 0.002 | 19778 | 0.002 | 19807 | 0.002
19950 | 0.002 | 19938 | 0.002 | 19911 | 0.002 | 19895 | 0.002 | 19924 | 0.002
20068 | 0.002 | 20055 | 0.002 | 20028 | 0.002 | 20013 | 0.002 | 20042 | 0.002
20185 | 0.002 | 20172 | 0.002 | 20146 | 0.002 | 20130 | 0.002 | 20159 | 0.002
20302 | 0.002 | 20290 | 0.002 | 20263 | 0.002 | 20247 | 0.002 | 20277 | 0.002
20420 | 0.002 | 20407 | 0.002 | 20380 | 0.002 | 20365 | 0.002 | 20395 | 0.002
20537 | 0.002 | 20525 | 0.002 | 20498 | 0.002 | 20483 | 0.002 | 20511 | 0.002
20655 | 0.002 | 20642 | 0.002 | 20616 | 0.002 | 20599 | 0.002 | 20629 | 0.002
20772 | 0.002 | 20759 | 0.002 | 20732 | 0.002 | 20717 | 0.002 | 20747 | 0.002
20889 | 0.002 | 20877 | 0.001 | 20850 | 0.002 | 20835 | 0.002 | 20864 | 0.002
21007 | 0.002 | 20995 | 0.001 | 20968 | 0.002 | 20952 | 0.002 | 20981 | 0.002
21125 | 0.002 | 21111 | 0.001 | 21085 | 0.002 | 21069 | 0.002 | 21099 | 0.002
21241 | 0.002 | 21229 | 0.001 | 21202 | 0.002 | 21187 | 0.002 | 21216 | 0.002
21359 | 0.002 | 21347 | 0.001 | 21320 | 0.002 | 21304 | 0.002 | 21334 | 0.002
21477 | 0.002 | 21464 | 0.001 | 21437 | 0.002 | 21422 | 0.002 | 21451 | 0.002
21594 | 0.002 | 21581 | 0.001 | 21555 | 0.002 | 21539 | 0.002 | 21568 | 0.002
21711 | 0.002 | 21699 | 0.001 | 21672 | 0.002 | 21656 | 0.002 | 21686 | 0.002
21829 | 0.002 | 21816 | 0.001 | 21789 | 0.002 | 21774 | 0.002 | 21804 | 0.002
21946 | 0.002 | 21934 | 0.001 | 21910 | 0.002 | 21892 | 0.002 | 21921 | 0.002
22064 | 0.002 | 22051 | 0.001 | 22028 | 0.002 | 22008 | 0.002 | 22038 | 0.002
22184 | 0.002 | 22169 | 0.001 | 22146 | 0.002 | 22126 | 0.002 | 22156 | 0.002
22302 | 0.002 | 22286 | 0.001 | 22262 | 0.002 | 22244 | 0.002 | 22270 | 0.002
22419 | 0.002 | 22404 | 0.001 | 22380 | 0.002 | 22361 | 0.002 | 22388 | 0.002
22537 | 0.002 | 22521 | 0.001 | 22498 | 0.002 | 22481 | 0.002 | 22506 | 0.002
22654 | 0.002 | 22641 | 0.001 | 22615 | 0.002 | 22599 | 0.002 | 22623 | 0.002
22771 | 0.002 | 22759 | 0.001 | 22732 | 0.002 | 22717 | 0.002 | 22740 | 0.002
22889 | 0.002 | 22877 | 0.001 | 22850 | 0.002 | 22834 | 0.002 | 22858 | 0.002
23007 | 0.002 | 22994 | 0.001 | 22967 | 0.002 | 22951 | 0.002 | 22975 | 0.002
23123 | 0.002 | 23111 | 0.001 | 23085 | 0.002 | 23069 | 0.002 | 23093 | 0.002
23241 | 0.002 | 23229 | 0.001 | 23202 | 0.002 | 23186 | 0.002 | 23211 | 0.002
23359 | 0.003 | 23347 | 0.001 | 23319 | 0.002 | 23304 | 0.002 | 23327 | 0.002
23476 | 0.002 | 23463 | 0.001 | 23437 | 0.002 | 23421 | 0.002 | 23445 | 0.002
23593 | 0.003 | 23581 | 0.002 | 23555 | 0.002 | 23538 | 0.002 | 23563 | 0.002
23711 | 0.003 | 23699 | 0.001 | 23671 | 0.002 | 23656 | 0.002 | 23680 | 0.002
23829 | 0.003 | 23816 | 0.001 | 23789 | 0.002 | 23774 | 0.002 | 23797 | 0.002
23946 | 0.003 | 23933 | 0.001 | 23907 | 0.002 | 23890 | 0.002 | 23915 | 0.002
24064 | 0.003 | 24051 | 0.002 | 24024 | 0.002 | 24008 | 0.002 | 24032 | 0.002
24181 | 0.003 | 24168 | 0.002 | 24141 | 0.002 | 24126 | 0.002 | 24171 | 0.002
24298 | 0.002 | 24286 | 0.002 | 24259 | 0.002 | 24243 | 0.002 | 24288 | 0.002
24416 | 0.003 | 24403 | 0.002 | 24376 | 0.002 | 24360 | 0.002 | 24406 | 0.002
24534 | 0.003 | 24520 | 0.001 | 24494 | 0.002 | 24478 | 0.002 | 24523 | 0.002
24650 | 0.003 | 24638 | 0.001 | 24611 | 0.002 | 24595 | 0.002 | 24640 | 0.002
24768 | 0.003 | 24756 | 0.001 | 24747 | 0.002 | 24713 | 0.002 | 24758 | 0.002
24886 | 0.003 | 24872 | 0.002 | 24864 | 0.002 | 24830 | 0.002 | 24876 | 0.002
25003 | 0.003 | 24990 | 0.001 | 24981 | 0.002 | 24967 | 0.002 | 24992 | 0.002
25120 | 0.003 | 25108 | 0.001 | 25099 | 0.002 | 25084 | 0.002 | 25110 | 0.002
25238 | 0.003 | 25225 | 0.002 | 25217 | 0.002 | 25202 | 0.002 | 25228 | 0.002
25355 | 0.003 | 25342 | 0.002 | 25333 | 0.002 | 25319 | 0.002 | 25345 | 0.002
25473 | 0.003 | 25460 | 0.001 | 25451 | 0.002 | 25437 | 0.002 | 25462 | 0.002
25590 | 0.003 | 25577 | 0.001 | 25569 | 0.002 | 25554 | 0.002 | 25580 | 0.002
25707 | 0.003 | 25695 | 0.001 | 25686 | 0.002 | 25671 | 0.002 | 25697 | 0.002
25840 | 0.003 | 25812 | 0.001 | 25803 | 0.002 | 25789 | 0.002 | 25815 | 0.002
25958 | 0.003 | 25947 | 0.001 | 25921 | 0.002 | 25907 | 0.002 | 25932 | 0.002
26075 | 0.003 | 26063 | 0.002 | 26038 | 0.002 | 26023 | 0.002 | 26050 | 0.002
26192 | 0.003 | 26181 | 0.001 | 26156 | 0.002 | 26141 | 0.002 | 26167 | 0.002
26310 | 0.003 | 26299 | 0.001 | 26273 | 0.002 | 26259 | 0.002 | 26285 | 0.002
26427 | 0.003 | 26416 | 0.001 | 26390 | 0.002 | 26376 | 0.002 | 26402 | 0.002
26545 | 0.003 | 26533 | 0.001 | 26511 | 0.002 | 26493 | 0.002 | 26519 | 0.002
26662 | 0.003 | 26651 | 0.001 | 26629 | 0.002 | 26611 | 0.002 | 26637 | 0.002
26780 | 0.003 | 26768 | 0.001 | 26746 | 0.002 | 26728 | 0.002 | 26755 | 0.002
26901 | 0.003 | 26886 | 0.001 | 26874 | 0.002 | 26846 | 0.002 | 26875 | 0.002
27018 | 0.003 | 27003 | 0.001 | 26992 | 0.002 | 26963 | 0.002 | 26992 | 0.002
27136 | 0.003 | 27120 | 0.001 | 27109 | 0.002 | 27084 | 0.002 | 27110 | 0.002
27264 | 0.003 | 27238 | 0.001 | 27227 | 0.002 | 27201 | 0.002 | 27228 | 0.002
27381 | 0.003 | 27354 | 0.001 | 27344 | 0.002 | 27319 | 0.002 | 27356 | 0.002
27498 | 0.003 | 27470 | 0.001 | 27461 | 0.002 | 27437 | 0.002 | 27473 | 0.002
27616 | 0.003 | 27588 | 0.001 | 27579 | 0.002 | 27565 | 0.002 | 27590 | 0.002
27734 | 0.003 | 27706 | 0.001 | 27697 | 0.002 | 27681 | 0.002 | 27708 | 0.002
27850 | 0.003 | 27834 | 0.001 | 27813 | 0.002 | 27799 | 0.002 | 27826 | 0.002
27968 | 0.003 | 27951 | 0.001 | 27931 | 0.002 | 27917 | 0.002 | 27942 | 0.002
28086 | 0.003 | 28068 | 0.001 | 28049 | 0.002 | 28034 | 0.002 | 28060 | 0.002
28204 | 0.003 | 28186 | 0.001 | 28166 | 0.002 | 28151 | 0.002 | 28178 | 0.002
28320 | 0.003 | 28303 | 0.001 | 28283 | 0.002 | 28269 | 0.002 | 28295 | 0.002
28438 | 0.003 | 28421 | 0.001 | 28401 | 0.002 | 28386 | 0.002 | 28412 | 0.002
28556 | 0.003 | 28538 | 0.002 | 28519 | 0.002 | 28504 | 0.002 | 28530 | 0.002
28673 | 0.003 | 28655 | 0.001 | 28636 | 0.002 | 28621 | 0.002 | 28647 | 0.002
28790 | 0.003 | 28773 | 0.001 | 28753 | 0.002 | 28738 | 0.002 | 28765 | 0.002
28908 | 0.003 | 28891 | 0.001 | 28871 | 0.002 | 28856 | 0.002 | 28882 | 0.002
29025 | 0.003 | 29007 | 0.001 | 28988 | 0.002 | 28974 | 0.002 | 28999 | 0.002
29143 | 0.003 | 29125 | 0.001 | 29106 | 0.002 | 29090 | 0.002 | 29117 | 0.002
29260 | 0.003 | 29243 | 0.001 | 29223 | 0.002 | 29208 | 0.002 | 29235 | 0.002
29377 | 0.003 | 29361 | 0.001 | 29340 | 0.002 | 29326 | 0.002 | 29351 | 0.002
29495 | 0.003 | 29477 | 0.001 | 29458 | 0.002 | 29443 | 0.002 | 29469 | 0.002
29613 | 0.003 | 29595 | 0.001 | 29576 | 0.002 | 29560 | 0.002 | 29587 | 0.002
29729 | 0.003 | 29713 | 0.001 | 29692 | 0.002 | 29678 | 0.002 | 29705 | 0.002
29847 | 0.003 | 29830 | 0.001 | 29810 | 0.002 | 29795 | 0.002 | 29821 | 0.002
29965 | 0.003 | 29947 | 0.001 | 29928 | 0.002 | 29913 | 0.002 | 29939 | 0.002
30082 | 0.003 | 30065 | 0.001 | 30045 | 0.002 | 30030 | 0.002 | 30057 | 0.002
30199 | 0.003 | 30182 | 0.001 | 30162 | 0.002 | 30147 | 0.002 | 30174 | 0.002
30317 | 0.003 | 30300 | 0.001 | 30280 | 0.002 | 30265 | 0.002 | 30291 | 0.002
30434 | 0.003 | 30417 | 0.001 | 30397 | 0.002 | 30383 | 0.002 | 30409 | 0.002
30552 | 0.003 | 30534 | 0.001 | 30515 | 0.002 | 30499 | 0.002 | 30526 | 0.002
30669 | 0.003 | 30652 | 0.001 | 30632 | 0.002 | 30617 | 0.002 | 30644 | 0.002
30786 | 0.003 | 30770 | 0.001 | 30749 | 0.002 | 30735 | 0.002 | 30761 | 0.002
30904 | 0.003 | 30886 | 0.001 | 30867 | 0.002 | 30852 | 0.002 | 30878 | 0.002
31022 | 0.003 | 31004 | 0.001 | 30985 | 0.002 | 30969 | 0.002 | 30996 | 0.002
31138 | 0.003 | 31122 | 0.001 | 31101 | 0.002 | 31087 | 0.002 | 31114 | 0.002
31256 | 0.003 | 31239 | 0.001 | 31217 | 0.002 | 31204 | 0.002 | 31230 | 0.002
31374 | 0.003 | 31356 | 0.001 | 31334 | 0.002 | 31322 | 0.002 | 31348 | 0.002
31495 | 0.003 | 31474 | 0.001 | 31452 | 0.002 | 31439 | 0.002 | 31469 | 0.002
31611 | 0.003 | 31591 | 0.001 | 31569 | 0.002 | 31556 | 0.002 | 31587 | 0.002
31729 | 0.003 | 31709 | 0.001 | 31687 | 0.002 | 31674 | 0.002 | 31704 | 0.002
31847 | 0.003 | 31826 | 0.001 | 31804 | 0.002 | 31795 | 0.002 | 31821 | 0.002
31964 | 0.003 | 31947 | 0.001 | 31922 | 0.002 | 31912 | 0.002 | 31939 | 0.002
32081 | 0.003 | 32064 | 0.001 | 32039 | 0.002 | 32029 | 0.002 | 32056 | 0.002
32199 | 0.002 | 32182 | 0.001 | 32156 | 0.002 | 32147 | 0.002 | 32174 | 0.002
32317 | 0.003 | 32300 | 0.001 | 32274 | 0.002 | 32265 | 0.002 | 32291 | 0.002
32434 | 0.002 | 32416 | 0.001 | 32392 | 0.002 | 32381 | 0.002 | 32408 | 0.002
32551 | 0.003 | 32534 | 0.001 | 32508 | 0.002 | 32499 | 0.002 | 32526 | 0.002
32669 | 0.003 | 32652 | 0.001 | 32626 | 0.002 | 32617 | 0.002 | 32644 | 0.002
32786 | 0.003 | 32769 | 0.001 | 32744 | 0.002 | 32734 | 0.002 | 32760 | 0.002
32904 | 0.002 | 32886 | 0.001 | 32861 | 0.002 | 32851 | 0.002 | 32878 | 0.002
33021 | 0.002 | 33004 | 0.001 | 32978 | 0.002 | 32969 | 0.002 | 32996 | 0.002
33138 | 0.002 | 33122 | 0.001 | 33096 | 0.002 | 33086 | 0.002 | 33114 | 0.002
33256 | 0.002 | 33239 | 0.001 | 33213 | 0.002 | 33204 | 0.002 | 33230 | 0.002
33374 | 0.002 | 33356 | 0.001 | 33331 | 0.002 | 33321 | 0.002 | 33348 | 0.002
33490 | 0.002 | 33474 | 0.001 | 33448 | 0.002 | 33438 | 0.002 | 33466 | 0.002
33608 | 0.002 | 33591 | 0.001 | 33565 | 0.002 | 33556 | 0.002 | 33583 | 0.002
33726 | 0.002 | 33709 | 0.001 | 33683 | 0.002 | 33674 | 0.002 | 33700 | 0.002
33843 | 0.002 | 33826 | 0.001 | 33801 | 0.002 | 33790 | 0.002 | 33818 | 0.002
33960 | 0.002 | 33943 | 0.001 | 33917 | 0.002 | 33908 | 0.002 | 33935 | 0.002
34078 | 0.003 | 34061 | 0.001 | 34035 | 0.002 | 34026 | 0.002 | 34053 | 0.002
34195 | 0.003 | 34179 | 0.001 | 34153 | 0.002 | 34143 | 0.002 | 34170 | 0.002
34313 | 0.002 | 34295 | 0.001 | 34270 | 0.002 | 34260 | 0.002 | 34287 | 0.002
34430 | 0.002 | 34413 | 0.001 | 34387 | 0.002 | 34378 | 0.002 | 34405 | 0.002
34547 | 0.003 | 34531 | 0.001 | 34505 | 0.002 | 34495 | 0.002 | 34523 | 0.002
34665 | 0.003 | 34648 | 0.001 | 34622 | 0.002 | 34613 | 0.002 | 34639 | 0.002
34783 | 0.002 | 34765 | 0.001 | 34740 | 0.002 | 34731 | 0.002 | 34757 | 0.002
34899 | 0.002 | 34883 | 0.001 | 34857 | 0.002 | 34847 | 0.002 | 34875 | 0.002
35017 | 0.002 | 35000 | 0.001 | 34974 | 0.002 | 34965 | 0.002 | 34992 | 0.002
35135 | 0.002 | 35118 | 0.001 | 35092 | 0.002 | 35083 | 0.002 | 35109 | 0.002
35252 | 0.003 | 35235 | 0.001 | 35210 | 0.002 | 35200 | 0.002 | 35227 | 0.002
35369 | 0.002 | 35352 | 0.001 | 35326 | 0.002 | 35317 | 0.002 | 35344 | 0.002
35487 | 0.003 | 35470 | 0.001 | 35444 | 0.002 | 35435 | 0.002 | 35462 | 0.002
35604 | 0.003 | 35588 | 0.001 | 35562 | 0.002 | 35552 | 0.002 | 35579 | 0.002
35722 | 0.002 | 35705 | 0.001 | 35680 | 0.002 | 35670 | 0.002 | 35696 | 0.002
35840 | 0.003 | 35822 | 0.001 | 35800 | 0.002 | 35787 | 0.002 | 35814 | 0.002
35956 | 0.003 | 35940 | 0.001 | 35917 | 0.002 | 35904 | 0.002 | 35932 | 0.002
36074 | 0.003 | 36058 | 0.001 | 36035 | 0.002 | 36022 | 0.002 | 36048 | 0.002
36195 | 0.003 | 36174 | 0.001 | 36152 | 0.002 | 36140 | 0.002 | 36169 | 0.002
36313 | 0.003 | 36292 | 0.001 | 36270 | 0.002 | 36256 | 0.002 | 36287 | 0.002
36429 | 0.003 | 36410 | 0.001 | 36387 | 0.002 | 36377 | 0.002 | 36405 | 0.002
36547 | 0.003 | 36531 | 0.001 | 36505 | 0.002 | 36495 | 0.002 | 36521 | 0.002
36665 | 0.003 | 36647 | 0.001 | 36622 | 0.002 | 36613 | 0.002 | 36639 | 0.002
36782 | 0.003 | 36765 | 0.001 | 36740 | 0.002 | 36729 | 0.002 | 36777 | 0.002
36899 | 0.003 | 36883 | 0.001 | 36857 | 0.002 | 36847 | 0.002 | 36895 | 0.002
37017 | 0.003 | 37000 | 0.001 | 36974 | 0.002 | 36965 | 0.002 | 37013 | 0.002
37135 | 0.003 | 37117 | 0.001 | 37092 | 0.002 | 37082 | 0.002 | 37130 | 0.002
37252 | 0.003 | 37235 | 0.001 | 37210 | 0.002 | 37199 | 0.002 | 37248 | 0.002
37369 | 0.003 | 37352 | 0.001 | 37326 | 0.002 | 37317 | 0.002 | 37365 | 0.002
37487 | 0.003 | 37470 | 0.001 | 37444 | 0.002 | 37434 | 0.002 | 37482 | 0.002
37604 | 0.002 | 37587 | 0.001 | 37562 | 0.002 | 37552 | 0.002 | 37600 | 0.002
37722 | 0.003 | 37704 | 0.001 | 37679 | 0.002 | 37689 | 0.002 | 37718 | 0.002
37839 | 0.003 | 37822 | 0.001 | 37814 | 0.002 | 37806 | 0.002 | 37834 | 0.002
37956 | 0.003 | 37940 | 0.001 | 37932 | 0.002 | 37923 | 0.002 | 37952 | 0.002
38074 | 0.003 | 38056 | 0.001 | 38049 | 0.002 | 38041 | 0.002 | 38070 | 0.002
38192 | 0.003 | 38174 | 0.001 | 38166 | 0.002 | 38158 | 0.002 | 38187 | 0.002
38308 | 0.002 | 38292 | 0.001 | 38284 | 0.002 | 38276 | 0.002 | 38304 | 0.002
38426 | 0.003 | 38410 | 0.001 | 38402 | 0.002 | 38393 | 0.002 | 38422 | 0.002
38544 | 0.003 | 38526 | 0.001 | 38518 | 0.002 | 38510 | 0.002 | 38540 | 0.002
38661 | 0.003 | 38644 | 0.001 | 38636 | 0.002 | 38628 | 0.002 | 38657 | 0.002
38778 | 0.003 | 38762 | 0.001 | 38754 | 0.002 | 38746 | 0.002 | 38774 | 0.002
38896 | 0.003 | 38879 | 0.001 | 38871 | 0.002 | 38862 | 0.002 | 38892 | 0.002
39013 | 0.003 | 38996 | 0.001 | 38988 | 0.002 | 38980 | 0.002 | 39009 | 0.002
39146 | 0.003 | 39131 | 0.001 | 39106 | 0.002 | 39098 | 0.002 | 39127 | 0.002
39264 | 0.003 | 39247 | 0.001 | 39223 | 0.002 | 39215 | 0.002 | 39244 | 0.002
39380 | 0.002 | 39365 | 0.001 | 39341 | 0.002 | 39332 | 0.002 | 39361 | 0.002
39498 | 0.003 | 39483 | 0.001 | 39458 | 0.002 | 39450 | 0.002 | 39479 | 0.002
39616 | 0.003 | 39600 | 0.001 | 39575 | 0.002 | 39567 | 0.002 | 39597 | 0.002
39733 | 0.003 | 39717 | 0.001 | 39693 | 0.002 | 39685 | 0.002 | 39713 | 0.002
39851 | 0.002 | 39835 | 0.001 | 39811 | 0.002 | 39802 | 0.002 | 39831 | 0.002
39968 | 0.002 | 39953 | 0.001 | 39928 | 0.002 | 39919 | 0.002 | 39949 | 0.002
40085 | 0.003 | 40070 | 0.001 | 40045 | 0.002 | 40037 | 0.002 | 40066 | 0.002
40203 | 0.003 | 40187 | 0.001 | 40163 | 0.002 | 40155 | 0.002 | 40183 | 0.002
40321 | 0.003 | 40305 | 0.001 | 40281 | 0.002 | 40271 | 0.002 | 40301 | 0.002
40438 | 0.003 | 40422 | 0.001 | 40400 | 0.002 | 40389 | 0.002 | 40418 | 0.002
40555 | 0.003 | 40540 | 0.001 | 40518 | 0.002 | 40507 | 0.002 | 40536 | 0.002
40673 | 0.003 | 40657 | 0.001 | 40636 | 0.002 | 40624 | 0.002 | 40653 | 0.002
40794 | 0.003 | 40774 | 0.001 | 40754 | 0.002 | 40741 | 0.002 | 40774 | 0.002
40911 | 0.003 | 40892 | 0.001 | 40870 | 0.002 | 40859 | 0.002 | 40891 | 0.002
41028 | 0.003 | 41010 | 0.001 | 40988 | 0.002 | 40976 | 0.002 | 41009 | 0.002
41146 | 0.003 | 41126 | 0.001 | 41106 | 0.002 | 41097 | 0.002 | 41126 | 0.002
41264 | 0.003 | 41247 | 0.001 | 41223 | 0.002 | 41214 | 0.002 | 41244 | 0.002
41380 | 0.003 | 41365 | 0.001 | 41340 | 0.002 | 41332 | 0.002 | 41361 | 0.002
41498 | 0.003 | 41483 | 0.001 | 41458 | 0.002 | 41449 | 0.002 | 41479 | 0.002
41616 | 0.003 | 41600 | 0.001 | 41576 | 0.002 | 41567 | 0.002 | 41596 | 0.002
41733 | 0.003 | 41717 | 0.001 | 41693 | 0.002 | 41684 | 0.002 | 41713 | 0.002
41850 | 0.003 | 41835 | 0.001 | 41810 | 0.002 | 41801 | 0.002 | 41831 | 0.002
41968 | 0.003 | 41952 | 0.001 | 41928 | 0.002 | 41919 | 0.002 | 41949 | 0.002
42086 | 0.003 | 42070 | 0.001 | 42045 | 0.002 | 42037 | 0.002 | 42065 | 0.002
42203 | 0.003 | 42187 | 0.001 | 42163 | 0.002 | 42155 | 0.002 | 42183 | 0.002
42320 | 0.003 | 42305 | 0.001 | 42280 | 0.002 | 42271 | 0.002 | 42301 | 0.002
42438 | 0.003 | 42422 | 0.001 | 42397 | 0.002 | 42389 | 0.002 | 42419 | 0.002
42555 | 0.003 | 42540 | 0.001 | 42515 | 0.002 | 42507 | 0.002 | 42535 | 0.002
42673 | 0.003 | 42657 | 0.001 | 42633 | 0.002 | 42624 | 0.002 | 42653 | 0.002
42790 | 0.003 | 42774 | 0.001 | 42750 | 0.002 | 42741 | 0.002 | 42771 | 0.002
42907 | 0.003 | 42892 | 0.001 | 42867 | 0.002 | 42859 | 0.002 | 42888 | 0.002
43025 | 0.003 | 43010 | 0.001 | 42985 | 0.002 | 42976 | 0.002 | 43005 | 0.002
43143 | 0.003 | 43127 | 0.001 | 43103 | 0.002 | 43094 | 0.002 | 43123 | 0.002
43260 | 0.003 | 43244 | 0.001 | 43219 | 0.002 | 43211 | 0.002 | 43241 | 0.002
43377 | 0.003 | 43362 | 0.001 | 43337 | 0.002 | 43328 | 0.002 | 43358 | 0.002
43495 | 0.002 | 43480 | 0.001 | 43455 | 0.002 | 43446 | 0.002 | 43475 | 0.002
43613 | 0.003 | 43596 | 0.001 | 43573 | 0.002 | 43564 | 0.002 | 43593 | 0.002
43729 | 0.003 | 43714 | 0.002 | 43689 | 0.002 | 43681 | 0.002 | 43710 | 0.002
43847 | 0.003 | 43832 | 0.001 | 43807 | 0.002 | 43798 | 0.002 | 43828 | 0.002
43965 | 0.003 | 43949 | 0.001 | 43925 | 0.002 | 43916 | 0.002 | 43945 | 0.002
44082 | 0.003 | 44066 | 0.001 | 44042 | 0.002 | 44034 | 0.002 | 44062 | 0.002
44199 | 0.003 | 44184 | 0.001 | 44160 | 0.002 | 44150 | 0.002 | 44180 | 0.002
44317 | 0.003 | 44302 | 0.001 | 44277 | 0.002 | 44268 | 0.002 | 44298 | 0.002
44435 | 0.003 | 44419 | 0.002 | 44394 | 0.002 | 44386 | 0.002 | 44416 | 0.002
44552 | 0.003 | 44536 | 0.001 | 44512 | 0.002 | 44503 | 0.002 | 44532 | 0.002
44669 | 0.003 | 44654 | 0.002 | 44630 | 0.002 | 44620 | 0.002 | 44650 | 0.002
44787 | 0.003 | 44771 | 0.001 | 44747 | 0.002 | 44738 | 0.002 | 44768 | 0.002
44904 | 0.003 | 44889 | 0.001 | 44864 | 0.002 | 44855 | 0.002 | 44884 | 0.002
45022 | 0.003 | 45006 | 0.002 | 44982 | 0.002 | 44973 | 0.002 | 45002 | 0.002
45139 | 0.003 | 45123 | 0.001 | 45103 | 0.002 | 45090 | 0.002 | 45120 | 0.002
45256 | 0.003 | 45241 | 0.001 | 45220 | 0.002 | 45208 | 0.002 | 45238 | 0.002
45377 | 0.003 | 45359 | 0.001 | 45337 | 0.002 | 45325 | 0.002 | 45358 | 0.002
45495 | 0.003 | 45476 | 0.001 | 45455 | 0.002 | 45443 | 0.002 | 45475 | 0.002
45613 | 0.003 | 45593 | 0.001 | 45573 | 0.002 | 45560 | 0.002 | 45593 | 0.002
45730 | 0.003 | 45711 | 0.002 | 45689 | 0.002 | 45681 | 0.002 | 45711 | 0.002
45847 | 0.003 | 45832 | 0.002 | 45807 | 0.002 | 45798 | 0.002 | 45828 | 0.002
45965 | 0.003 | 45949 | 0.002 | 45925 | 0.002 | 45916 | 0.002 | 45945 | 0.002
46083 | 0.003 | 46066 | 0.002 | 46043 | 0.002 | 46033 | 0.002 | 46063 | 0.002
46199 | 0.003 | 46184 | 0.002 | 46159 | 0.002 | 46150 | 0.002 | 46180 | 0.002
46317 | 0.003 | 46302 | 0.001 | 46277 | 0.002 | 46268 | 0.002 | 46298 | 0.002
46435 | 0.003 | 46418 | 0.002 | 46395 | 0.002 | 46386 | 0.002 | 46415 | 0.002
46552 | 0.003 | 46536 | 0.002 | 46512 | 0.002 | 46502 | 0.002 | 46533 | 0.002
46669 | 0.003 | 46654 | 0.002 | 46629 | 0.002 | 46620 | 0.002 | 46650 | 0.002
46787 | 0.003 | 46772 | 0.002 | 46747 | 0.002 | 46738 | 0.002 | 46768 | 0.002
46905 | 0.003 | 46888 | 0.002 | 46865 | 0.002 | 46856 | 0.002 | 46885 | 0.002
47022 | 0.003 | 47006 | 0.001 | 46982 | 0.002 | 46972 | 0.002 | 47002 | 0.002
47139 | 0.003 | 47124 | 0.001 | 47099 | 0.002 | 47090 | 0.002 | 47120 | 0.002
47257 | 0.003 | 47241 | 0.002 | 47217 | 0.002 | 47208 | 0.002 | 47238 | 0.002
47374 | 0.002 | 47358 | 0.002 | 47334 | 0.002 | 47325 | 0.002 | 47354 | 0.002
47492 | 0.003 | 47476 | 0.002 | 47452 | 0.002 | 47442 | 0.002 | 47472 | 0.002
47609 | 0.002 | 47594 | 0.001 | 47569 | 0.002 | 47560 | 0.002 | 47590 | 0.002
47726 | 0.002 | 47711 | 0.002 | 47686 | 0.002 | 47677 | 0.002 | 47708 | 0.002
47844 | 0.002 | 47828 | 0.002 | 47804 | 0.002 | 47795 | 0.002 | 47824 | 0.002
47962 | 0.002 | 47946 | 0.002 | 47922 | 0.002 | 47912 | 0.002 | 47942 | 0.002
48079 | 0.002 | 48063 | 0.002 | 48039 | 0.002 | 48029 | 0.001 | 48060 | 0.002
48196 | 0.002 | 48181 | 0.002 | 48156 | 0.002 | 48147 | 0.001 | 48177 | 0.002
48314 | 0.002 | 48299 | 0.002 | 48274 | 0.002 | 48265 | 0.002 | 48294 | 0.002
48432 | 0.002 | 48415 | 0.001 | 48392 | 0.002 | 48382 | 0.002 | 48412 | 0.002
48548 | 0.002 | 48533 | 0.001 | 48508 | 0.002 | 48499 | 0.001 | 48530 | 0.002
48666 | 0.003 | 48651 | 0.002 | 48626 | 0.002 | 48617 | 0.002 | 48647 | 0.002
48784 | 0.002 | 48769 | 0.002 | 48744 | 0.002 | 48735 | 0.002 | 48764 | 0.002
48902 | 0.002 | 48885 | 0.002 | 48861 | 0.002 | 48851 | 0.002 | 48882 | 0.002
49018 | 0.002 | 49003 | 0.002 | 48978 | 0.002 | 48969 | 0.001 | 48999 | 0.002
49136 | 0.002 | 49121 | 0.001 | 49096 | 0.002 | 49087 | 0.001 | 49117 | 0.002
49254 | 0.002 | 49238 | 0.001 | 49214 | 0.002 | 49204 | 0.002 | 49234 | 0.002
49371 | 0.002 | 49355 | 0.001 | 49331 | 0.002 | 49321 | 0.002 | 49373 | 0.002
49488 | 0.002 | 49473 | 0.001 | 49448 | 0.002 | 49439 | 0.001 | 49490 | 0.002
49606 | 0.002 | 49591 | 0.001 | 49566 | 0.002 | 49557 | 0.001 | 49607 | 0.002
49723 | 0.002 | 49708 | 0.001 | 49687 | 0.002 | 49674 | 0.001 | 49725 | 0.002
49841 | 0.002 | 49825 | 0.001 | 49804 | 0.002 | 49792 | 0.001 | 49843 | 0.002
49958 | 0.002 | 49943 | 0.001 | 49922 | 0.002 | 49909 | 0.002 | 49959 | 0.003
50079 | 0.002 | 50060 | 0.001 | 50039 | 0.002 | 50026 | 0.002 | 50080 | 0.003
50197 | 0.002 | 50178 | 0.001 | 50156 | 0.002 | 50144 | 0.002 | 50198 | 0.002
50314 | 0.002 | 50295 | 0.001 | 50274 | 0.002 | 50265 | 0.001 | 50316 | 0.002
50432 | 0.002 | 50416 | 0.001 | 50392 | 0.002 | 50382 | 0.001 | 50434 | 0.003
50549 | 0.002 | 50533 | 0.001 | 50509 | 0.002 | 50519 | 0.002 | 50550 | 0.002
50666 | 0.002 | 50651 | 0.001 | 50626 | 0.002 | 50636 | 0.002 | 50668 | 0.002
50784 | 0.002 | 50768 | 0.001 | 50761 | 0.002 | 50754 | 0.002 | 50786 | 0.002
50902 | 0.002 | 50886 | 0.001 | 50879 | 0.002 | 50871 | 0.002 | 50903 | 0.002
51018 | 0.002 | 51003 | 0.001 | 50997 | 0.002 | 50989 | 0.001 | 51020 | 0.003
51136 | 0.002 | 51121 | 0.001 | 51113 | 0.002 | 51106 | 0.001 | 51138 | 0.003
51254 | 0.002 | 51238 | 0.001 | 51231 | 0.002 | 51223 | 0.001 | 51255 | 0.003
51372 | 0.002 | 51355 | 0.001 | 51349 | 0.002 | 51341 | 0.002 | 51373 | 0.002
51488 | 0.003 | 51473 | 0.001 | 51467 | 0.002 | 51459 | 0.001 | 51490 | 0.002
51606 | 0.002 | 51591 | 0.001 | 51584 | 0.002 | 51575 | 0.002 | 51608 | 0.002
51724 | 0.003 | 51707 | 0.001 | 51701 | 0.002 | 51693 | 0.001 | 51725 | 0.003
51841 | 0.003 | 51825 | 0.001 | 51819 | 0.002 | 51811 | 0.001 | 51843 | 0.002
51958 | 0.002 | 51943 | 0.001 | 51936 | 0.002 | 51929 | 0.001 | 51960 | 0.002
52076 | 0.002 | 52061 | 0.001 | 52054 | 0.002 | 52045 | 0.001 | 52077 | 0.002
52193 | 0.003 | 52177 | 0.001 | 52171 | 0.002 | 52163 | 0.001 | 52195 | 0.002
52311 | 0.002 | 52312 | 0.001 | 52288 | 0.002 | 52281 | 0.001 | 52313 | 0.003
52428 | 0.002 | 52429 | 0.001 | 52406 | 0.002 | 52398 | 0.001 | 52430 | 0.003
52561 | 0.002 | 52546 | 0.001 | 52524 | 0.002 | 52515 | 0.001 | 52547 | 0.003
52678 | 0.002 | 52664 | 0.001 | 52641 | 0.002 | 52633 | 0.001 | 52665 | 0.003
52796 | 0.003 | 52782 | 0.001 | 52758 | 0.002 | 52750 | 0.001 | 52783 | 0.002
52914 | 0.002 | 52899 | 0.001 | 52876 | 0.002 | 52868 | 0.001 | 52899 | 0.003
53030 | 0.002 | 53016 | 0.001 | 52994 | 0.002 | 52985 | 0.001 | 53017 | 0.002
53148 | 0.002 | 53134 | 0.001 | 53110 | 0.002 | 53103 | 0.001 | 53135 | 0.003
53266 | 0.003 | 53252 | 0.001 | 53228 | 0.002 | 53220 | 0.001 | 53252 | 0.003
53384 | 0.003 | 53368 | 0.001 | 53346 | 0.002 | 53338 | 0.001 | 53369 | 0.003
53500 | 0.002 | 53486 | 0.001 | 53464 | 0.002 | 53455 | 0.001 | 53487 | 0.003
53618 | 0.002 | 53604 | 0.001 | 53580 | 0.002 | 53572 | 0.001 | 53605 | 0.003
53736 | 0.003 | 53721 | 0.001 | 53698 | 0.002 | 53690 | 0.001 | 53722 | 0.003
53853 | 0.003 | 53838 | 0.001 | 53816 | 0.002 | 53808 | 0.002 | 53839 | 0.003
53970 | 0.003 | 53956 | 0.001 | 53933 | 0.002 | 53924 | 0.002 | 53957 | 0.003
54088 | 0.003 | 54074 | 0.001 | 54050 | 0.002 | 54042 | 0.001 | 54074 | 0.003
54205 | 0.002 | 54191 | 0.001 | 54168 | 0.002 | 54160 | 0.002 | 54192 | 0.003
54323 | 0.002 | 54308 | 0.001 | 54289 | 0.002 | 54277 | 0.001 | 54309 | 0.002
54440 | 0.002 | 54426 | 0.001 | 54406 | 0.002 | 54394 | 0.002 | 54426 | 0.003
54558 | 0.002 | 54543 | 0.001 | 54523 | 0.002 | 54512 | 0.001 | 54544 | 0.003
54679 | 0.002 | 54661 | 0.001 | 54641 | 0.002 | 54630 | 0.002 | 54665 | 0.002
54796 | 0.002 | 54778 | 0.001 | 54759 | 0.002 | 54747 | 0.002 | 54782 | 0.002
54914 | 0.003 | 54895 | 0.001 | 54876 | 0.002 | 54864 | 0.002 | 54900 | 0.002
55031 | 0.002 | 55013 | 0.001 | 54993 | 0.002 | 54985 | 0.002 | 55017 | 0.002
55148 | 0.002 | 55134 | 0.001 | 55111 | 0.002 | 55103 | 0.002 | 55135 | 0.003
55266 | 0.002 | 55251 | 0.001 | 55228 | 0.002 | 55220 | 0.001 | 55252 | 0.002
55384 | 0.002 | 55369 | 0.001 | 55346 | 0.002 | 55337 | 0.002 | 55369 | 0.002
55501 | 0.002 | 55486 | 0.001 | 55463 | 0.002 | 55455 | 0.002 | 55487 | 0.002
55618 | 0.002 | 55604 | 0.001 | 55581 | 0.002 | 55572 | 0.001 | 55605 | 0.002
55736 | 0.002 | 55721 | 0.001 | 55698 | 0.002 | 55690 | 0.002 | 55723 | 0.002
55854 | 0.002 | 55838 | 0.001 | 55827 | 0.002 | 55808 | 0.002 | 55839 | 0.002
55970 | 0.002 | 55956 | 0.001 | 55944 | 0.002 | 55925 | 0.002 | 55957 | 0.002
56088 | 0.002 | 56074 | 0.001 | 56061 | 0.002 | 56042 | 0.002 | 56075 | 0.002
56216 | 0.002 | 56192 | 0.001 | 56179 | 0.002 | 56160 | 0.002 | 56203 | 0.002
56334 | 0.002 | 56308 | 0.001 | 56296 | 0.002 | 56278 | 0.002 | 56321 | 0.002
56452 | 0.002 | 56426 | 0.001 | 56414 | 0.002 | 56395 | 0.001 | 56437 | 0.002
56568 | 0.002 | 56544 | 0.001 | 56531 | 0.002 | 56523 | 0.001 | 56555 | 0.002
56686 | 0.002 | 56672 | 0.001 | 56648 | 0.002 | 56640 | 0.002 | 56673 | 0.003
56804 | 0.002 | 56790 | 0.001 | 56766 | 0.002 | 56758 | 0.002 | 56790 | 0.002
56922 | 0.002 | 56906 | 0.001 | 56884 | 0.002 | 56876 | 0.002 | 56907 | 0.002
57038 | 0.002 | 57024 | 0.001 | 57001 | 0.002 | 56992 | 0.002 | 57025 | 0.002
57156 | 0.002 | 57142 | 0.001 | 57118 | 0.002 | 57110 | 0.002 | 57143 | 0.002
57274 | 0.002 | 57259 | 0.001 | 57236 | 0.002 | 57228 | 0.001 | 57260 | 0.002
57391 | 0.002 | 57376 | 0.001 | 57354 | 0.002 | 57345 | 0.001 | 57377 | 0.002
57508 | 0.002 | 57494 | 0.001 | 57470 | 0.002 | 57462 | 0.002 | 57495 | 0.002
57626 | 0.002 | 57612 | 0.001 | 57588 | 0.002 | 57580 | 0.002 | 57612 | 0.002
57743 | 0.002 | 57729 | 0.001 | 57706 | 0.002 | 57697 | 0.002 | 57730 | 0.002
57861 | 0.002 | 57846 | 0.001 | 57823 | 0.002 | 57815 | 0.001 | 57847 | 0.003
57978 | 0.002 | 57964 | 0.001 | 57940 | 0.002 | 57932 | 0.001 | 57965 | 0.003
58096 | 0.002 | 58081 | 0.001 | 58058 | 0.002 | 58050 | 0.002 | 58082 | 0.002
58213 | 0.003 | 58199 | 0.001 | 58176 | 0.002 | 58167 | 0.002 | 58200 | 0.002
58331 | 0.003 | 58316 | 0.001 | 58293 | 0.002 | 58285 | 0.002 | 58317 | 0.002
58448 | 0.003 | 58433 | 0.001 | 58411 | 0.002 | 58402 | 0.002 | 58434 | 0.002
58565 | 0.003 | 58551 | 0.001 | 58528 | 0.002 | 58519 | 0.002 | 58552 | 0.002
58683 | 0.003 | 58669 | 0.001 | 58645 | 0.002 | 58637 | 0.002 | 58670 | 0.002
58801 | 0.002 | 58786 | 0.001 | 58763 | 0.002 | 58755 | 0.002 | 58786 | 0.002
58917 | 0.002 | 58903 | 0.001 | 58881 | 0.002 | 58871 | 0.002 | 58904 | 0.002
59035 | 0.002 | 59021 | 0.001 | 59001 | 0.002 | 58989 | 0.002 | 59022 | 0.002
59153 | 0.003 | 59139 | 0.001 | 59118 | 0.002 | 59107 | 0.002 | 59140 | 0.003
59271 | 0.002 | 59255 | 0.001 | 59236 | 0.002 | 59225 | 0.002 | 59256 | 0.002
59385 | 0.003 | 59373 | 0.001 | 59354 | 0.002 | 59341 | 0.002 | 59377 | 0.002
59503 | 0.002 | 59491 | 0.001 | 59471 | 0.002 | 59459 | 0.002 | 59495 | 0.002
59620 | 0.002 | 59608 | 0.001 | 59588 | 0.002 | 59580 | 0.002 | 59613 | 0.002
59738 | 0.002 | 59728 | 0.001 | 59706 | 0.002 | 59698 | 0.002 | 59729 | 0.002
59855 | 0.003 | 59846 | 0.001 | 59824 | 0.002 | 59814 | 0.001 | 59847 | 0.002
59972 | 0.003 | 59964 | 0.001 | 59940 | 0.002 | 59932 | 0.002 | 59965 | 0.002
60090 | 0.003 | 60082 | 0.001 | 60058 | 0.002 | 60050 | 0.002 | 60082 | 0.002
60208 | 0.003 | 60198 | 0.001 | 60176 | 0.002 | 60167 | 0.002 | 60199 | 0.002
60325 | 0.002 | 60316 | 0.001 | 60294 | 0.002 | 60284 | 0.002 | 60317 | 0.002
60442 | 0.002 | 60434 | 0.001 | 60410 | 0.002 | 60402 | 0.002 | 60435 | 0.002
60560 | 0.002 | 60551 | 0.001 | 60528 | 0.002 | 60519 | 0.002 | 60552 | 0.002
60678 | 0.003 | 60668 | 0.001 | 60646 | 0.002 | 60637 | 0.002 | 60669 | 0.002
60794 | 0.002 | 60786 | 0.001 | 60763 | 0.002 | 60754 | 0.002 | 60787 | 0.002
60912 | 0.002 | 60903 | 0.001 | 60880 | 0.002 | 60872 | 0.002 | 60904 | 0.003
61030 | 0.002 | 61021 | 0.001 | 60998 | 0.098 | 60989 | 0.002 | 61022 | 0.003
61147 | 0.002 | 61138 | 0.001 | 61115 | 0.187 | 61107 | 0.006 | 61139 | 0.003
61264 | 0.002 | 61256 | 0.001 | 61233 | 0.203 | 61224 | 0.036 | 61257 | 0.026
61382 | 0.008 | 61373 | 0.001 | 61350 | 0.202 | 61341 | 0.131 | 61374 | 0.114
61500 | 0.06 | 61491 | 0.001 | 61468 | 0.202 | 61459 | 0.204 | 61492 | 0.203
61617 | 0.196 | 61608 | 0.012 | 61585 | 0.202 | 61577 | 0.202 | 61609 | 0.202
61734 | 0.202 | 61725 | 0.068 | 61703 | 0.202 | 61693 | 0.202 | 61726 | 0.202
61852 | 0.202 | 61843 | 0.15 | 61820 | 0.202 | 61811 | 0.202 | 61865 | 0.202
61969 | 0.202 | 61961 | 0.201 | 61937 | 0.202 | 61929 | 0.201 | 61982 | 0.202
62087 | 0.202 | 62078 | 0.201 | 62055 | 0.202 | 62047 | 0.202 | 62100 | 0.202
62204 | 0.202 | 62195 | 0.201 | 62173 | 0.202 | 62163 | 0.202 | 62218 | 0.202
62322 | 0.202 | 62313 | 0.201 | 62290 | 0.202 | 62281 | 0.201 | 62334 | 0.202
62439 | 0.202 | 62431 | 0.201 | 62407 | 0.202 | 62399 | 0.202 | 62452 | 0.202
62557 | 0.202 | 62547 | 0.201 | 62525 | 0.202 | 62516 | 0.202 | 62570 | 0.202
62674 | 0.202 | 62665 | 0.201 | 62643 | 0.202 | 62633 | 0.202 | 62688 | 0.202
62791 | 0.202 | 62783 | 0.201 | 62759 | 0.202 | 62751 | 0.202 | 62804 | 0.202
62909 | 0.202 | 62900 | 0.201 | 62877 | 0.202 | 62868 | 0.202 | 62922 | 0.203
63027 | 0.202 | 63017 | 0.201 | 62995 | 0.202 | 62986 | 0.202 | 63040 | 0.203
63144 | 0.202 | 63135 | 0.201 | 63113 | 0.202 | 63104 | 0.202 | 63157 | 0.202
63261 | 0.202 | 63253 | 0.201 | 63229 | 0.202 | 63221 | 0.202 | 63275 | 0.202
63379 | 0.202 | 63370 | 0.201 | 63347 | 0.202 | 63357 | 0.202 | 63392 | 0.203
63497 | 0.202 | 63488 | 0.201 | 63465 | 0.202 | 63475 | 0.202 | 63510 | 0.202
63613 | 0.202 | 63605 | 0.201 | 63586 | 0.202 | 63593 | 0.202 | 63627 | 0.202
63731 | 0.202 | 63722 | 0.201 | 63703 | 0.202 | 63710 | 0.202 | 63745 | 0.202
63849 | 0.202 | 63840 | 0.201 | 63839 | 0.202 | 63827 | 0.202 | 63862 | 0.202
63970 | 0.202 | 63958 | 0.201 | 63955 | 0.202 | 63945 | 0.202 | 63983 | 0.202
64088 | 0.202 | 64075 | 0.201 | 64073 | 0.202 | 64062 | 0.202 | 64100 | 0.202
64204 | 0.202 | 64192 | 0.201 | 64191 | 0.202 | 64180 | 0.202 | 64218 | 0.202
64322 | 0.202 | 64310 | 0.201 | 64308 | 0.202 | 64295 | 0.202 | 64335 | 0.202
64440 | 0.202 | 64431 | 0.201 | 64425 | 0.202 | 64412 | 0.202 | 64453 | 0.202
64557 | 0.202 | 64548 | 0.201 | 64543 | 0.202 | 64530 | 0.202 | 64570 | 0.202
64674 | 0.202 | 64665 | 0.201 | 64661 | 0.202 | 64648 | 0.202 | 64688 | 0.202
64792 | 0.202 | 64783 | 0.201 | 64778 | 0.202 | 64764 | 0.202 | 64805 | 0.202
64910 | 0.202 | 64901 | 0.201 | 64896 | 0.202 | 64882 | 0.202 | 64922 | 0.203
65027 | 0.202 | 65018 | 0.201 | 65013 | 0.202 | 65000 | 0.202 | 65040 | 0.202
65144 | 0.202 | 65135 | 0.201 | 65130 | 0.202 | 65117 | 0.202 | 65158 | 0.202
65262 | 0.202 | 65253 | 0.201 | 65248 | 0.202 | 65234 | 0.202 | 65275 | 0.202
65379 | 0.202 | 65371 | 0.201 | 65366 | 0.202 | 65352 | 0.202 | 65392 | 0.202
65497 | 0.202 | 65504 | 0.201 | 65482 | 0.202 | 65470 | 0.202 | 65510 | 0.202
65614 | 0.202 | 65622 | 0.201 | 65600 | 0.202 | 65587 | 0.202 | 65628 | 0.202
65732 | 0.202 | 65739 | 0.201 | 65718 | 0.202 | 65704 | 0.202 | 65744 | 0.202
65849 | 0.202 | 65857 | 0.201 | 65836 | 0.202 | 65822 | 0.202 | 65862 | 0.202
65982 | 0.202 | 65974 | 0.201 | 65952 | 0.202 | 65939 | 0.202 | 65980 | 0.202
66100 | 0.202 | 66092 | 0.201 | 66070 | 0.202 | 66057 | 0.202 | 66098 | 0.202
66216 | 0.202 | 66209 | 0.201 | 66188 | 0.202 | 66174 | 0.202 | 66214 | 0.202
66334 | 0.202 | 66326 | 0.201 | 66305 | 0.202 | 66291 | 0.202 | 66332 | 0.202
66452 | 0.202 | 66444 | 0.201 | 66422 | 0.202 | 66409 | 0.202 | 66450 | 0.202
66570 | 0.202 | 66562 | 0.201 | 66540 | 0.202 | 66527 | 0.202 | 66567 | 0.203
66686 | 0.202 | 66679 | 0.201 | 66658 | 0.202 | 66644 | 0.202 | 66684 | 0.202
66804 | 0.202 | 66796 | 0.201 | 66775 | 0.202 | 66761 | 0.202 | 66802 | 0.203
66922 | 0.202 | 66914 | 0.201 | 66892 | 0.202 | 66879 | 0.202 | 66920 | 0.202
67039 | 0.202 | 67032 | 0.201 | 67010 | 0.202 | 66997 | 0.202 | 67037 | 0.202
67156 | 0.202 | 67148 | 0.201 | 67127 | 0.202 | 67113 | 0.202 | 67154 | 0.202
67274 | 0.202 | 67266 | 0.201 | 67245 | 0.202 | 67231 | 0.202 | 67272 | 0.202
67392 | 0.202 | 67384 | 0.201 | 67362 | 0.202 | 67349 | 0.202 | 67389 | 0.202
67509 | 0.202 | 67502 | 0.201 | 67480 | 0.202 | 67467 | 0.202 | 67507 | 0.202
67627 | 0.202 | 67618 | 0.201 | 67597 | 0.202 | 67584 | 0.202 | 67624 | 0.202
67744 | 0.202 | 67736 | 0.201 | 67715 | 0.202 | 67701 | 0.202 | 67742 | 0.202
67861 | 0.202 | 67854 | 0.201 | 67832 | 0.202 | 67819 | 0.202 | 67859 | 0.202
67979 | 0.202 | 67971 | 0.201 | 67949 | 0.202 | 67936 | 0.202 | 67977 | 0.202
68097 | 0.202 | 68088 | 0.201 | 68067 | 0.202 | 68054 | 0.202 | 68094 | 0.202
68214 | 0.202 | 68206 | 0.201 | 68185 | 0.202 | 68171 | 0.202 | 68211 | 0.203
68331 | 0.202 | 68324 | 0.201 | 68305 | 0.202 | 68288 | 0.202 | 68329 | 0.203
68449 | 0.202 | 68441 | 0.201 | 68423 | 0.202 | 68406 | 0.202 | 68447 | 0.202
68570 | 0.202 | 68558 | 0.201 | 68540 | 0.202 | 68524 | 0.202 | 68568 | 0.202
68687 | 0.202 | 68676 | 0.201 | 68658 | 0.202 | 68641 | 0.202 | 68685 | 0.202
68804 | 0.202 | 68793 | 0.201 | 68775 | 0.202 | 68758 | 0.202 | 68802 | 0.202
68922 | 0.202 | 68911 | 0.201 | 68892 | 0.202 | 68879 | 0.202 | 68920 | 0.202
69040 | 0.202 | 69031 | 0.201 | 69010 | 0.202 | 68997 | 0.202 | 69038 | 0.202
69157 | 0.202 | 69149 | 0.201 | 69128 | 0.202 | 69114 | 0.202 | 69154 | 0.202
69274 | 0.202 | 69267 | 0.201 | 69245 | 0.202 | 69231 | 0.201 | 69272 | 0.202
69392 | 0.202 | 69384 | 0.201 | 69362 | 0.202 | 69349 | 0.202 | 69390 | 0.202
69510 | 0.202 | 69501 | 0.201 | 69480 | 0.202 | 69467 | 0.202 | 69508 | 0.202
69626 | 0.202 | 69619 | 0.201 | 69598 | 0.202 | 69583 | 0.202 | 69624 | 0.202
69744 | 0.202 | 69736 | 0.201 | 69714 | 0.202 | 69701 | 0.202 | 69742 | 0.202
69862 | 0.202 | 69854 | 0.201 | 69832 | 0.202 | 69819 | 0.201 | 69860 | 0.202
69980 | 0.202 | 69971 | 0.201 | 69950 | 0.202 | 69937 | 0.202 | 69977 | 0.202
70096 | 0.202 | 70089 | 0.201 | 70068 | 0.202 | 70053 | 0.202 | 70094 | 0.202
70214 | 0.202 | 70206 | 0.201 | 70184 | 0.202 | 70171 | 0.202 | 70212 | 0.202
70332 | 0.202 | 70324 | 0.201 | 70302 | 0.202 | 70289 | 0.202 | 70330 | 0.202
70449 | 0.202 | 70441 | 0.201 | 70420 | 0.202 | 70406 | 0.202 | 70447 | 0.202
70566 | 0.202 | 70558 | 0.201 | 70537 | 0.202 | 70523 | 0.202 | 70564 | 0.202
70684 | 0.202 | 70676 | 0.201 | 70654 | 0.202 | 70641 | 0.202 | 70682 | 0.202
70802 | 0.202 | 70794 | 0.201 | 70772 | 0.202 | 70758 | 0.202 | 70799 | 0.202
70919 | 0.202 | 70910 | 0.201 | 70890 | 0.202 | 70876 | 0.202 | 70917 | 0.202
71036 | 0.202 | 71028 | 0.201 | 71007 | 0.202 | 70993 | 0.202 | 71034 | 0.202
71154 | 0.202 | 71146 | 0.201 | 71124 | 0.202 | 71111 | 0.201 | 71152 | 0.202
71271 | 0.202 | 71264 | 0.201 | 71242 | 0.202 | 71228 | 0.201 | 71269 | 0.202
71389 | 0.202 | 71380 | 0.201 | 71359 | 0.202 | 71346 | 0.202 | 71387 | 0.202
71506 | 0.202 | 71498 | 0.201 | 71477 | 0.202 | 71463 | 0.202 | 71504 | 0.202
71624 | 0.202 | 71616 | 0.201 | 71595 | 0.202 | 71580 | 0.202 | 71621 | 0.202
71741 | 0.202 | 71733 | 0.201 | 71712 | 0.202 | 71698 | 0.202 | 71739 | 0.202
71859 | 0.202 | 71851 | 0.201 | 71829 | 0.202 | 71816 | 0.202 | 71857 | 0.202
71976 | 0.202 | 71968 | 0.201 | 71947 | 0.202 | 71933 | 0.202 | 71974 | 0.202
72093 | 0.202 | 72086 | 0.201 | 72065 | 0.202 | 72050 | 0.202 | 72091 | 0.202
72211 | 0.202 | 72203 | 0.201 | 72181 | 0.202 | 72168 | 0.202 | 72209 | 0.202
72329 | 0.202 | 72321 | 0.201 | 72299 | 0.202 | 72286 | 0.202 | 72327 | 0.202
72446 | 0.202 | 72438 | 0.201 | 72417 | 0.202 | 72402 | 0.202 | 72443 | 0.202
72563 | 0.202 | 72555 | 0.201 | 72535 | 0.202 | 72520 | 0.202 | 72561 | 0.202
72681 | 0.202 | 72673 | 0.201 | 72651 | 0.202 | 72638 | 0.202 | 72679 | 0.202
72799 | 0.202 | 72791 | 0.201 | 72769 | 0.202 | 72755 | 0.202 | 72796 | 0.202
72915 | 0.202 | 72908 | 0.201 | 72890 | 0.202 | 72872 | 0.202 | 72913 | 0.202
73033 | 0.202 | 73025 | 0.201 | 73008 | 0.202 | 72990 | 0.202 | 73031 | 0.202
73151 | 0.202 | 73143 | 0.201 | 73124 | 0.202 | 73108 | 0.202 | 73149 | 0.202
73272 | 0.202 | 73261 | 0.201 | 73242 | 0.202 | 73225 | 0.202 | 73270 | 0.202
73390 | 0.202 | 73377 | 0.201 | 73360 | 0.202 | 73342 | 0.202 | 73387 | 0.202
73506 | 0.202 | 73495 | 0.201 | 73478 | 0.202 | 73463 | 0.202 | 73504 | 0.202
73624 | 0.202 | 73616 | 0.201 | 73594 | 0.202 | 73581 | 0.202 | 73622 | 0.202
73742 | 0.202 | 73734 | 0.201 | 73712 | 0.202 | 73698 | 0.202 | 73739 | 0.202
73859 | 0.202 | 73851 | 0.201 | 73830 | 0.202 | 73816 | 0.202 | 73857 | 0.202
73976 | 0.202 | 73968 | 0.201 | 73947 | 0.202 | 73933 | 0.202 | 73974 | 0.202
74094 | 0.202 | 74086 | 0.201 | 74064 | 0.202 | 74050 | 0.202 | 74092 | 0.202
74212 | 0.202 | 74204 | 0.201 | 74182 | 0.202 | 74168 | 0.202 | 74209 | 0.202
74329 | 0.202 | 74320 | 0.201 | 74300 | 0.202 | 74286 | 0.202 | 74327 | 0.202
74446 | 0.202 | 74438 | 0.201 | 74417 | 0.202 | 74403 | 0.202 | 74465 | 0.202
74564 | 0.202 | 74556 | 0.201 | 74534 | 0.202 | 74520 | 0.202 | 74583 | 0.202
74681 | 0.202 | 74674 | 0.201 | 74652 | 0.202 | 74638 | 0.202 | 74700 | 0.202
74799 | 0.202 | 74790 | 0.201 | 74769 | 0.202 | 74756 | 0.202 | 74817 | 0.202
74916 | 0.202 | 74908 | 0.201 | 74887 | 0.202 | 74872 | 0.202 | 74935 | 0.202
75034 | 0.202 | 75026 | 0.201 | 75004 | 0.202 | 74990 | 0.202 | 75053 | 0.202
75151 | 0.202 | 75143 | 0.201 | 75122 | 0.202 | 75108 | 0.202 | 75170 | 0.202
75269 | 0.202 | 75260 | 0.201 | 75239 | 0.202 | 75225 | 0.202 | 75287 | 0.202
75386 | 0.202 | 75378 | 0.201 | 75357 | 0.202 | 75342 | 0.202 | 75405 | 0.202
75503 | 0.202 | 75495 | 0.201 | 75474 | 0.202 | 75460 | 0.201 | 75523 | 0.202
75621 | 0.202 | 75613 | 0.201 | 75591 | 0.202 | 75578 | 0.202 | 75639 | 0.202
75739 | 0.202 | 75730 | 0.201 | 75709 | 0.202 | 75695 | 0.202 | 75757 | 0.202
75856 | 0.202 | 75848 | 0.201 | 75827 | 0.202 | 75812 | 0.202 | 75875 | 0.202
75973 | 0.202 | 75965 | 0.201 | 75944 | 0.202 | 75930 | 0.202 | 75993 | 0.202
76091 | 0.202 | 76083 | 0.201 | 76061 | 0.202 | 76047 | 0.202 | 76109 | 0.202
76209 | 0.202 | 76200 | 0.201 | 76179 | 0.202 | 76184 | 0.202 | 76227 | 0.202
76325 | 0.202 | 76317 | 0.201 | 76297 | 0.202 | 76302 | 0.202 | 76345 | 0.202
76443 | 0.202 | 76435 | 0.201 | 76413 | 0.202 | 76419 | 0.202 | 76462 | 0.202
76561 | 0.202 | 76553 | 0.201 | 76531 | 0.202 | 76536 | 0.202 | 76579 | 0.202
76679 | 0.202 | 76670 | 0.201 | 76649 | 0.202 | 76654 | 0.202 | 76697 | 0.202
76795 | 0.202 | 76787 | 0.201 | 76767 | 0.202 | 76772 | 0.202 | 76815 | 0.202
76913 | 0.202 | 76905 | 0.201 | 76902 | 0.202 | 76889 | 0.201 | 76932 | 0.202
77031 | 0.202 | 77023 | 0.201 | 77019 | 0.202 | 77006 | 0.201 | 77049 | 0.202
77148 | 0.202 | 77139 | 0.201 | 77136 | 0.202 | 77124 | 0.202 | 77167 | 0.202
77265 | 0.202 | 77257 | 0.201 | 77254 | 0.202 | 77241 | 0.202 | 77284 | 0.202
77383 | 0.202 | 77375 | 0.201 | 77371 | 0.202 | 77359 | 0.202 | 77402 | 0.202
77501 | 0.202 | 77493 | 0.201 | 77492 | 0.202 | 77476 | 0.202 | 77519 | 0.202
77618 | 0.202 | 77609 | 0.201 | 77609 | 0.202 | 77594 | 0.202 | 77637 | 0.202
77735 | 0.202 | 77727 | 0.201 | 77727 | 0.202 | 77711 | 0.202 | 77754 | 0.202
77856 | 0.202 | 77845 | 0.201 | 77845 | 0.202 | 77829 | 0.202 | 77875 | 0.202
77974 | 0.202 | 77962 | 0.201 | 77962 | 0.202 | 77946 | 0.202 | 77992 | 0.202
78091 | 0.202 | 78080 | 0.201 | 78080 | 0.202 | 78063 | 0.202 | 78110 | 0.202
78209 | 0.202 | 78197 | 0.201 | 78197 | 0.202 | 78184 | 0.202 | 78227 | 0.202
78326 | 0.202 | 78318 | 0.201 | 78315 | 0.201 | 78302 | 0.202 | 78345 | 0.202
78444 | 0.202 | 78436 | 0.201 | 78432 | 0.201 | 78419 | 0.202 | 78462 | 0.202
78561 | 0.202 | 78553 | 0.201 | 78550 | 0.201 | 78537 | 0.202 | 78580 | 0.202
78679 | 0.202 | 78687 | 0.201 | 78667 | 0.202 | 78654 | 0.202 | 78697 | 0.202
78796 | 0.202 | 78804 | 0.201 | 78785 | 0.202 | 78772 | 0.202 | 78815 | 0.202
78914 | 0.202 | 78922 | 0.201 | 78902 | 0.202 | 78889 | 0.202 | 78932 | 0.202
79031 | 0.202 | 79039 | 0.201 | 79020 | 0.202 | 79006 | 0.202 | 79050 | 0.202
79149 | 0.202 | 79157 | 0.201 | 79137 | 0.202 | 79124 | 0.202 | 79167 | 0.202
79282 | 0.202 | 79275 | 0.202 | 79255 | 0.201 | 79242 | 0.202 | 79285 | 0.202
79399 | 0.202 | 79392 | 0.202 | 79372 | 0.201 | 79360 | 0.202 | 79403 | 0.202
79516 | 0.202 | 79509 | 0.201 | 79490 | 0.201 | 79476 | 0.202 | 79520 | 0.202
79634 | 0.202 | 79627 | 0.201 | 79607 | 0.201 | 79594 | 0.202 | 79637 | 0.202
79752 | 0.202 | 79745 | 0.201 | 79724 | 0.201 | 79712 | 0.202 | 79755 | 0.202
79868 | 0.202 | 79861 | 0.201 | 79842 | 0.201 | 79830 | 0.202 | 79873 | 0.202
79986 | 0.202 | 79979 | 0.201 | 79960 | 0.201 | 79946 | 0.202 | 79989 | 0.202
80104 | 0.202 | 80097 | 0.202 | 80077 | 0.201 | 80064 | 0.202 | 80107 | 0.202
80222 | 0.202 | 80215 | 0.201 | 80194 | 0.201 | 80182 | 0.202 | 80225 | 0.202
80338 | 0.202 | 80331 | 0.201 | 80312 | 0.201 | 80299 | 0.202 | 80343 | 0.202
80456 | 0.202 | 80449 | 0.201 | 80430 | 0.201 | 80416 | 0.202 | 80459 | 0.202
80574 | 0.202 | 80567 | 0.201 | 80547 | 0.202 | 80534 | 0.202 | 80577 | 0.202
80692 | 0.202 | 80685 | 0.201 | 80664 | 0.201 | 80652 | 0.202 | 80695 | 0.202
80808 | 0.202 | 80801 | 0.201 | 80782 | 0.201 | 80769 | 0.202 | 80813 | 0.202
80926 | 0.202 | 80919 | 0.201 | 80900 | 0.201 | 80886 | 0.202 | 80929 | 0.202
81044 | 0.202 | 81037 | 0.201 | 81017 | 0.201 | 81004 | 0.202 | 81047 | 0.202
81162 | 0.202 | 81155 | 0.201 | 81134 | 0.201 | 81121 | 0.202 | 81165 | 0.202
81278 | 0.202 | 81271 | 0.201 | 81252 | 0.201 | 81239 | 0.202 | 81283 | 0.202
81396 | 0.202 | 81389 | 0.201 | 81370 | 0.201 | 81356 | 0.202 | 81399 | 0.202
81514 | 0.202 | 81507 | 0.202 | 81487 | 0.201 | 81474 | 0.202 | 81517 | 0.203
81631 | 0.202 | 81625 | 0.201 | 81604 | 0.201 | 81591 | 0.202 | 81635 | 0.203
81748 | 0.202 | 81741 | 0.201 | 81722 | 0.201 | 81709 | 0.202 | 81753 | 0.203
81866 | 0.202 | 81859 | 0.201 | 81840 | 0.201 | 81826 | 0.202 | 81869 | 0.203
81984 | 0.202 | 81977 | 0.201 | 81956 | 0.201 | 81944 | 0.202 | 81987 | 0.203
82101 | 0.202 | 82094 | 0.201 | 82074 | 0.201 | 82061 | 0.202 | 82105 | 0.202
82219 | 0.202 | 82211 | 0.201 | 82195 | 0.201 | 82179 | 0.202 | 82222 | 0.202
82336 | 0.202 | 82329 | 0.201 | 82313 | 0.201 | 82296 | 0.202 | 82339 | 0.203
82457 | 0.202 | 82447 | 0.201 | 82430 | 0.201 | 82414 | 0.202 | 82457 | 0.203
82575 | 0.202 | 82564 | 0.201 | 82547 | 0.201 | 82531 | 0.202 | 82572 | 0.202
82692 | 0.202 | 82681 | 0.201 | 82665 | 0.201 | 82649 | 0.202 | 82690 | 0.203
82809 | 0.202 | 82799 | 0.201 | 82783 | 0.201 | 82769 | 0.202 | 82807 | 0.203
82927 | 0.202 | 82920 | 0.201 | 82900 | 0.201 | 82887 | 0.202 | 82925 | 0.203
83045 | 0.202 | 83038 | 0.201 | 83017 | 0.201 | 83004 | 0.202 | 83042 | 0.203
83162 | 0.202 | 83154 | 0.201 | 83135 | 0.202 | 83122 | 0.202 | 83160 | 0.203
83279 | 0.202 | 83272 | 0.201 | 83253 | 0.201 | 83239 | 0.202 | 83277 | 0.203
83397 | 0.202 | 83390 | 0.201 | 83371 | 0.201 | 83357 | 0.202 | 83395 | 0.203
83515 | 0.202 | 83508 | 0.201 | 83487 | 0.201 | 83474 | 0.202 | 83512 | 0.203
83632 | 0.202 | 83624 | 0.201 | 83605 | 0.201 | 83592 | 0.202 | 83630 | 0.202
83749 | 0.202 | 83742 | 0.201 | 83723 | 0.201 | 83709 | 0.202 | 83747 | 0.203
83867 | 0.202 | 83860 | 0.201 | 83841 | 0.201 | 83827 | 0.202 | 83865 | 0.203
83984 | 0.202 | 83977 | 0.201 | 83957 | 0.201 | 83944 | 0.202 | 83982 | 0.203
84102 | 0.202 | 84094 | 0.201 | 84075 | 0.201 | 84062 | 0.202 | 84100 | 0.203
84219 | 0.202 | 84212 | 0.201 | 84193 | 0.201 | 84179 | 0.202 | 84217 | 0.202
84337 | 0.202 | 84330 | 0.201 | 84311 | 0.201 | 84296 | 0.202 | 84335 | 0.202
84454 | 0.202 | 84447 | 0.201 | 84427 | 0.201 | 84414 | 0.202 | 84452 | 0.202
84572 | 0.202 | 84564 | 0.201 | 84545 | 0.201 | 84532 | 0.202 | 84570 | 0.202
84689 | 0.202 | 84682 | 0.201 | 84663 | 0.201 | 84649 | 0.202 | 84687 | 0.202
84807 | 0.202 | 84800 | 0.201 | 84791 | 0.201 | 84766 | 0.202 | 84805 | 0.203
84925 | 0.202 | 84917 | 0.201 | 84909 | 0.201 | 84884 | 0.202 | 84922 | 0.202
85042 | 0.202 | 85035 | 0.201 | 85026 | 0.201 | 85002 | 0.202 | 85040 | 0.202
85170 | 0.202 | 85152 | 0.201 | 85143 | 0.201 | 85120 | 0.202 | 85168 | 0.202
85287 | 0.202 | 85270 | 0.201 | 85261 | 0.201 | 85237 | 0.202 | 85285 | 0.202
85405 | 0.202 | 85388 | 0.201 | 85379 | 0.201 | 85354 | 0.202 | 85403 | 0.202
85523 | 0.202 | 85505 | 0.201 | 85495 | 0.202 | 85482 | 0.202 | 85521 | 0.202
85640 | 0.202 | 85633 | 0.201 | 85613 | 0.201 | 85600 | 0.202 | 85638 | 0.202
85757 | 0.202 | 85750 | 0.201 | 85731 | 0.201 | 85718 | 0.202 | 85755 | 0.202
85875 | 0.202 | 85868 | 0.201 | 85849 | 0.201 | 85835 | 0.202 | 85873 | 0.202
85993 | 0.202 | 85986 | 0.201 | 85965 | 0.201 | 85952 | 0.202 | 85990 | 0.203
86110 | 0.202 | 86103 | 0.201 | 86083 | 0.201 | 86070 | 0.202 | 86108 | 0.202
86227 | 0.202 | 86220 | 0.201 | 86201 | 0.201 | 86188 | 0.202 | 86225 | 0.203
86345 | 0.202 | 86338 | 0.201 | 86319 | 0.201 | 86304 | 0.202 | 86343 | 0.203
86463 | 0.202 | 86456 | 0.201 | 86435 | 0.201 | 86422 | 0.202 | 86460 | 0.202
86580 | 0.202 | 86573 | 0.201 | 86553 | 0.201 | 86540 | 0.202 | 86578 | 0.202
86697 | 0.202 | 86690 | 0.201 | 86671 | 0.201 | 86658 | 0.202 | 86695 | 0.203
86815 | 0.202 | 86808 | 0.201 | 86792 | 0.201 | 86774 | 0.202 | 86813 | 0.203
86932 | 0.202 | 86926 | 0.201 | 86909 | 0.201 | 86892 | 0.202 | 86930 | 0.202
87050 | 0.202 | 87043 | 0.201 | 87026 | 0.201 | 87010 | 0.202 | 87069 | 0.202
87171 | 0.202 | 87160 | 0.201 | 87144 | 0.201 | 87128 | 0.202 | 87190 | 0.202
87288 | 0.202 | 87278 | 0.201 | 87262 | 0.201 | 87244 | 0.202 | 87307 | 0.203
87406 | 0.202 | 87395 | 0.201 | 87379 | 0.201 | 87365 | 0.202 | 87425 | 0.202
87523 | 0.202 | 87513 | 0.201 | 87496 | 0.201 | 87483 | 0.202 | 87542 | 0.202
87641 | 0.202 | 87628 | 0.201 | 87614 | 0.201 | 87601 | 0.202 | 87660 | 0.203
87758 | 0.202 | 87745 | 0.201 | 87732 | 0.201 | 87718 | 0.202 | 87777 | 0.203
87876 | 0.202 | 87863 | 0.201 | 87849 | 0.202 | 87835 | 0.202 | 87895 | 0.203
87993 | 0.202 | 87981 | 0.201 | 87966 | 0.202 | 87953 | 0.202 | 88012 | 0.202
88111 | 0.202 | 88098 | 0.201 | 88084 | 0.202 | 88071 | 0.202 | 88130 | 0.203
88228 | 0.202 | 88215 | 0.201 | 88202 | 0.202 | 88187 | 0.202 | 88247 | 0.203
88346 | 0.202 | 88333 | 0.201 | 88318 | 0.201 | 88305 | 0.202 | 88365 | 0.203
88463 | 0.202 | 88451 | 0.201 | 88436 | 0.202 | 88423 | 0.202 | 88482 | 0.202
88581 | 0.202 | 88568 | 0.201 | 88554 | 0.202 | 88541 | 0.202 | 88600 | 0.203
88698 | 0.202 | 88685 | 0.201 | 88672 | 0.202 | 88657 | 0.202 | 88717 | 0.203
88816 | 0.202 | 88803 | 0.201 | 88788 | 0.202 | 88775 | 0.202 | 88835 | 0.202
88933 | 0.202 | 88921 | 0.201 | 88906 | 0.202 | 88912 | 0.202 | 88952 | 0.203
89051 | 0.202 | 89037 | 0.201 | 89024 | 0.202 | 89029 | 0.202 | 89069 | 0.203
89168 | 0.202 | 89155 | 0.201 | 89142 | 0.202 | 89147 | 0.202 | 89187 | 0.202
89285 | 0.202 | 89273 | 0.201 | 89259 | 0.202 | 89265 | 0.202 | 89305 | 0.203
89403 | 0.202 | 89391 | 0.201 | 89376 | 0.202 | 89382 | 0.202 | 89422 | 0.203
89521 | 0.202 | 89507 | 0.201 | 89494 | 0.202 | 89499 | 0.202 | 89539 | 0.203
89638 | 0.202 | 89625 | 0.201 | 89612 | 0.202 | 89617 | 0.202 | 89657 | 0.202
89755 | 0.202 | 89743 | 0.201 | 89729 | 0.202 | 89735 | 0.202 | 89775 | 0.203
89873 | 0.202 | 89861 | 0.201 | 89864 | 0.202 | 89852 | 0.202 | 89892 | 0.202
89991 | 0.202 | 89977 | 0.201 | 89981 | 0.202 | 89969 | 0.202 | 90009 | 0.202
90108 | 0.202 | 90095 | 0.201 | 90099 | 0.202 | 90087 | 0.202 | 90127 | 0.203
90225 | 0.202 | 90213 | 0.201 | 90217 | 0.202 | 90205 | 0.202 | 90245 | 0.203
90343 | 0.202 | 90331 | 0.201 | 90334 | 0.202 | 90321 | 0.202 | 90362 | 0.202
90461 | 0.202 | 90447 | 0.201 | 90451 | 0.202 | 90439 | 0.202 | 90479 | 0.203
90578 | 0.202 | 90565 | 0.201 | 90569 | 0.202 | 90557 | 0.202 | 90597 | 0.203
90695 | 0.202 | 90683 | 0.201 | 90687 | 0.202 | 90675 | 0.202 | 90715 | 0.203
90813 | 0.202 | 90801 | 0.201 | 90804 | 0.202 | 90791 | 0.202 | 90832 | 0.203
90931 | 0.202 | 90917 | 0.201 | 90921 | 0.202 | 90909 | 0.202 | 90949 | 0.203
91048 | 0.202 | 91035 | 0.201 | 91039 | 0.202 | 91027 | 0.202 | 91067 | 0.203
91165 | 0.202 | 91153 | 0.201 | 91157 | 0.202 | 91145 | 0.202 | 91185 | 0.203
91283 | 0.202 | 91270 | 0.201 | 91274 | 0.202 | 91261 | 0.202 | 91301 | 0.203
91401 | 0.202 | 91387 | 0.201 | 91392 | 0.202 | 91379 | 0.202 | 91419 | 0.203
91517 | 0.202 | 91505 | 0.201 | 91507 | 0.202 | 91497 | 0.202 | 91537 | 0.203
91635 | 0.202 | 91623 | 0.201 | 91624 | 0.202 | 91614 | 0.202 | 91655 | 0.203
91756 | 0.202 | 91740 | 0.201 | 91742 | 0.202 | 91731 | 0.202 | 91776 | 0.203
91874 | 0.202 | 91857 | 0.201 | 91859 | 0.202 | 91849 | 0.202 | 91892 | 0.203
91992 | 0.202 | 91992 | 0.201 | 91976 | 0.202 | 91967 | 0.202 | 92010 | 0.203
92108 | 0.202 | 92109 | 0.201 | 92094 | 0.202 | 92088 | 0.202 | 92128 | 0.203
92226 | 0.202 | 92230 | 0.201 | 92212 | 0.202 | 92205 | 0.202 | 92246 | 0.203
92344 | 0.202 | 92347 | 0.201 | 92329 | 0.202 | 92322 | 0.202 | 92362 | 0.203
92462 | 0.202 | 92465 | 0.201 | 92446 | 0.202 | 92440 | 0.202 | 92480 | 0.203
92578 | 0.202 | 92582 | 0.201 | 92564 | 0.202 | 92558 | 0.202 | 92598 | 0.203
92711 | 0.202 | 92700 | 0.201 | 92682 | 0.202 | 92675 | 0.202 | 92716 | 0.203
92829 | 0.202 | 92817 | 0.201 | 92800 | 0.202 | 92792 | 0.202 | 92832 | 0.203
92947 | 0.202 | 92935 | 0.201 | 92916 | 0.202 | 92910 | 0.202 | 92950 | 0.203
93064 | 0.202 | 93052 | 0.201 | 93034 | 0.202 | 93027 | 0.202 | 93068 | 0.203
93181 | 0.202 | 93170 | 0.201 | 93152 | 0.202 | 93145 | 0.202 | 93186 | 0.203
93299 | 0.202 | 93287 | 0.201 | 93270 | 0.202 | 93262 | 0.202 | 93302 | 0.203
93417 | 0.202 | 93405 | 0.201 | 93386 | 0.202 | 93380 | 0.202 | 93420 | 0.203
93534 | 0.202 | 93522 | 0.201 | 93504 | 0.201 | 93497 | 0.202 | 93538 | 0.203
93651 | 0.202 | 93640 | 0.201 | 93622 | 0.201 | 93615 | 0.202 | 93656 | 0.203
93769 | 0.202 | 93757 | 0.201 | 93740 | 0.201 | 93732 | 0.202 | 93772 | 0.203
93886 | 0.202 | 93875 | 0.201 | 93856 | 0.202 | 93850 | 0.202 | 93890 | 0.203
94004 | 0.202 | 93992 | 0.201 | 93974 | 0.202 | 93967 | 0.202 | 94008 | 0.203
94121 | 0.202 | 94110 | 0.201 | 94092 | 0.202 | 94085 | 0.202 | 94126 | 0.203
94239 | 0.201 | 94227 | 0.201 | 94209 | 0.202 | 94202 | 0.202 | 94242 | 0.203
94356 | 0.201 | 94345 | 0.201 | 94326 | 0.202 | 94320 | 0.202 | 94360 | 0.203
94474 | 0.202 | 94462 | 0.201 | 94444 | 0.202 | 94437 | 0.202 | 94478 | 0.203
94592 | 0.202 | 94579 | 0.201 | 94562 | 0.202 | 94555 | 0.202 | 94596 | 0.203
94709 | 0.202 | 94697 | 0.201 | 94679 | 0.202 | 94672 | 0.202 | 94712 | 0.203
94826 | 0.202 | 94815 | 0.201 | 94796 | 0.202 | 94789 | 0.202 | 94830 | 0.203
94944 | 0.202 | 94932 | 0.201 | 94914 | 0.202 | 94907 | 0.202 | 94948 | 0.203
95062 | 0.202 | 95049 | 0.201 | 95032 | 0.202 | 95025 | 0.202 | 95065 | 0.203
95179 | 0.202 | 95167 | 0.201 | 95149 | 0.202 | 95142 | 0.202 | 95182 | 0.203
95296 | 0.202 | 95285 | 0.201 | 95266 | 0.201 | 95259 | 0.202 | 95300 | 0.203
95414 | 0.202 | 95403 | 0.201 | 95384 | 0.201 | 95377 | 0.202 | 95418 | 0.203
95532 | 0.202 | 95519 | 0.201 | 95502 | 0.201 | 95495 | 0.202 | 95535 | 0.203
95648 | 0.202 | 95637 | 0.201 | 95619 | 0.201 | 95611 | 0.202 | 95652 | 0.203
95766 | 0.202 | 95755 | 0.201 | 95736 | 0.201 | 95729 | 0.202 | 95770 | 0.203
95884 | 0.202 | 95873 | 0.201 | 95854 | 0.201 | 95847 | 0.202 | 95888 | 0.203
96002 | 0.202 | 95989 | 0.201 | 95972 | 0.201 | 95965 | 0.202 | 96005 | 0.203
96118 | 0.202 | 96107 | 0.201 | 96093 | 0.201 | 96081 | 0.202 | 96122 | 0.203
96236 | 0.202 | 96225 | 0.201 | 96209 | 0.201 | 96199 | 0.202 | 96240 | 0.203
96354 | 0.202 | 96342 | 0.201 | 96327 | 0.201 | 96317 | 0.202 | 96358 | 0.203
96475 | 0.202 | 96459 | 0.201 | 96445 | 0.201 | 96435 | 0.202 | 96479 | 0.203
96592 | 0.202 | 96577 | 0.201 | 96563 | 0.201 | 96551 | 0.202 | 96595 | 0.203
96709 | 0.201 | 96695 | 0.201 | 96679 | 0.201 | 96672 | 0.202 | 96713 | 0.203
96827 | 0.202 | 96816 | 0.201 | 96797 | 0.201 | 96790 | 0.202 | 96831 | 0.203
96945 | 0.202 | 96932 | 0.201 | 96915 | 0.201 | 96908 | 0.203 | 96949 | 0.203
97062 | 0.202 | 97050 | 0.201 | 97032 | 0.201 | 97025 | 0.202 | 97065 | 0.203
97179 | 0.202 | 97168 | 0.201 | 97149 | 0.201 | 97142 | 0.202 | 97183 | 0.203
97297 | 0.202 | 97286 | 0.201 | 97267 | 0.201 | 97260 | 0.202 | 97301 | 0.203
97415 | 0.202 | 97402 | 0.201 | 97385 | 0.202 | 97378 | 0.202 | 97419 | 0.203
97532 | 0.202 | 97520 | 0.201 | 97502 | 0.201 | 97495 | 0.202 | 97536 | 0.203
97649 | 0.202 | 97638 | 0.201 | 97619 | 0.201 | 97612 | 0.202 | 97653 | 0.203
97767 | 0.202 | 97756 | 0.201 | 97737 | 0.202 | 97730 | 0.202 | 97771 | 0.203
97885 | 0.202 | 97872 | 0.201 | 97855 | 0.202 | 97848 | 0.202 | 97889 | 0.203
98002 | 0.202 | 97990 | 0.201 | 97972 | 0.201 | 97965 | 0.202 | 98006 | 0.203
98119 | 0.202 | 98108 | 0.201 | 98089 | 0.202 | 98082 | 0.202 | 98123 | 0.203
98237 | 0.202 | 98226 | 0.201 | 98207 | 0.201 | 98200 | 0.202 | 98241 | 0.203
98355 | 0.202 | 98342 | 0.201 | 98325 | 0.201 | 98317 | 0.202 | 98359 | 0.203
98471 | 0.202 | 98460 | 0.201 | 98442 | 0.202 | 98435 | 0.202 | 98476 | 0.203
98589 | 0.202 | 98578 | 0.201 | 98560 | 0.201 | 98552 | 0.202 | 98593 | 0.203
98707 | 0.202 | 98696 | 0.201 | 98677 | 0.201 | 98670 | 0.202 | 98711 | 0.203
98825 | 0.202 | 98812 | 0.201 | 98795 | 0.201 | 98787 | 0.202 | 98828 | 0.203
98941 | 0.202 | 98930 | 0.201 | 98912 | 0.201 | 98905 | 0.202 | 98946 | 0.203
99059 | 0.202 | 99048 | 0.201 | 99030 | 0.201 | 99022 | 0.202 | 99063 | 0.203
99177 | 0.201 | 99165 | 0.201 | 99147 | 0.201 | 99140 | 0.202 | 99181 | 0.203
99295 | 0.202 | 99282 | 0.201 | 99264 | 0.201 | 99257 | 0.202 | 99298 | 0.203
99411 | 0.202 | 99400 | 0.201 | 99382 | 0.201 | 99375 | 0.202 | 99416 | 0.203
99529 | 0.201 | 99518 | 0.201 | 99500 | 0.201 | 99492 | 0.202 | 99533 | 0.203
99647 | 0.201 | 99635 | 0.201 | 99617 | 0.201 | 99610 | 0.202 | 99672 | 0.203
99765 | 0.201 | 99752 | 0.201 | 99734 | 0.201 | 99727 | 0.202 | 99789 | 0.203
99881 | 0.201 | 99870 | 0.201 | 99852 | 0.201 | 99845 | 0.202 | 99907 | 0.203
99999 | 0.201 | 99988 | 0.201 | 99970 | 0.201 | 99962 | 0.202 | 100024 | 0.203
100117 | 0.202 | 100105 | 0.201 | 100087 | 0.201 | 100079 | 0.202 | 100142 | 0.203
100235 | 0.202 | 100222 | 0.201 | 100204 | 0.201 | 100197 | 0.202 | 100259 | 0.203
100352 | 0.202 | 100340 | 0.201 | 100322 | 0.201 | 100315 | 0.203 | 100377 | 0.203
100469 | 0.202 | 100458 | 0.201 | 100440 | 0.201 | 100432 | 0.202 | 100494 | 0.203
100587 | 0.202 | 100575 | 0.201 | 100557 | 0.201 | 100549 | 0.202 | 100612 | 0.203
100705 | 0.201 | 100692 | 0.201 | 100678 | 0.201 | 100667 | 0.202 | 100729 | 0.203
100822 | 0.201 | 100810 | 0.201 | 100795 | 0.201 | 100785 | 0.202 | 100847 | 0.203
100942 | 0.201 | 100928 | 0.201 | 100913 | 0.201 | 100902 | 0.202 | 100968 | 0.203
101060 | 0.201 | 101045 | 0.201 | 101030 | 0.201 | 101019 | 0.202 | 101086 | 0.203
101178 | 0.201 | 101163 | 0.201 | 101148 | 0.201 | 101137 | 0.202 | 101202 | 0.203
101296 | 0.201 | 101280 | 0.201 | 101266 | 0.201 | 101258 | 0.202 | 101320 | 0.203
101412 | 0.202 | 101401 | 0.201 | 101383 | 0.201 | 101376 | 0.202 | 101438 | 0.203
101530 | 0.202 | 101519 | 0.201 | 101500 | 0.201 | 101493 | 0.202 | 101556 | 0.203
101648 | 0.201 | 101637 | 0.201 | 101618 | 0.201 | 101611 | 0.202 | 101672 | 0.203
101766 | 0.201 | 101753 | 0.201 | 101736 | 0.202 | 101747 | 0.202 | 101790 | 0.203
101882 | 0.202 | 101871 | 0.201 | 101853 | 0.201 | 101865 | 0.202 | 101908 | 0.203
102000 | 0.202 | 101989 | 0.201 | 101970 | 0.201 | 101983 | 0.202 | 102026 | 0.203
102118 | 0.201 | 102107 | 0.201 | 102088 | 0.202 | 102100 | 0.202 | 102142 | 0.203
102236 | 0.201 | 102223 | 0.201 | 102206 | 0.202 | 102217 | 0.202 | 102260 | 0.203
102353 | 0.201 | 102341 | 0.201 | 102323 | 0.201 | 102335 | 0.202 | 102378 | 0.203
102470 | 0.201 | 102459 | 0.201 | 102440 | 0.201 | 102453 | 0.202 | 102496 | 0.203
102588 | 0.202 | 102577 | 0.201 | 102558 | 0.201 | 102571 | 0.202 | 102612 | 0.203
102706 | 0.201 | 102693 | 0.202 | 102676 | 0.201 | 102687 | 0.202 | 102730 | 0.203
102823 | 0.201 | 102811 | 0.201 | 102793 | 0.202 | 102805 | 0.202 | 102848 | 0.203
102940 | 0.201 | 102929 | 0.201 | 102928 | 0.201 | 102923 | 0.202 | 102966 | 0.203
103058 | 0.201 | 103047 | 0.202 | 103046 | 0.202 | 103041 | 0.202 | 103084 | 0.203
103176 | 0.201 | 103163 | 0.201 | 103163 | 0.202 | 103157 | 0.202 | 103200 | 0.203
103293 | 0.201 | 103281 | 0.202 | 103281 | 0.202 | 103275 | 0.202 | 103318 | 0.203
103410 | 0.201 | 103399 | 0.201 | 103399 | 0.202 | 103393 | 0.202 | 103436 | 0.203
103528 | 0.201 | 103517 | 0.201 | 103516 | 0.202 | 103511 | 0.202 | 103554 | 0.203
103646 | 0.201 | 103633 | 0.202 | 103633 | 0.201 | 103627 | 0.202 | 103670 | 0.203
103763 | 0.201 | 103751 | 0.202 | 103751 | 0.201 | 103745 | 0.202 | 103788 | 0.203
103880 | 0.201 | 103869 | 0.202 | 103869 | 0.202 | 103863 | 0.202 | 103906 | 0.203
103998 | 0.201 | 103987 | 0.201 | 103986 | 0.202 | 103981 | 0.202 | 104024 | 0.203
104116 | 0.201 | 104104 | 0.201 | 104103 | 0.202 | 104097 | 0.202 | 104140 | 0.203
104233 | 0.201 | 104221 | 0.202 | 104221 | 0.202 | 104215 | 0.202 | 104258 | 0.203
104350 | 0.201 | 104339 | 0.201 | 104339 | 0.202 | 104333 | 0.202 | 104376 | 0.203
104468 | 0.201 | 104457 | 0.201 | 104456 | 0.202 | 104451 | 0.202 | 104494 | 0.203
104586 | 0.201 | 104574 | 0.201 | 104573 | 0.202 | 104567 | 0.203 | 104610 | 0.203
104704 | 0.201 | 104691 | 0.201 | 104691 | 0.202 | 104685 | 0.203 | 104728 | 0.203
104820 | 0.201 | 104809 | 0.201 | 104809 | 0.202 | 104803 | 0.203 | 104846 | 0.203
104938 | 0.201 | 104927 | 0.202 | 104926 | 0.202 | 104921 | 0.203 | 104964 | 0.203
105056 | 0.201 | 105044 | 0.201 | 105041 | 0.201 | 105037 | 0.202 | 105080 | 0.203
105174 | 0.201 | 105178 | 0.202 | 105159 | 0.201 | 105155 | 0.203 | 105198 | 0.202
105294 | 0.201 | 105295 | 0.201 | 105276 | 0.202 | 105273 | 0.203 | 105319 | 0.203
105412 | 0.201 | 105413 | 0.201 | 105394 | 0.202 | 105391 | 0.203 | 105437 | 0.203
105529 | 0.201 | 105531 | 0.201 | 105511 | 0.201 | 105511 | 0.203 | 105554 | 0.202
105647 | 0.201 | 105652 | 0.202 | 105629 | 0.202 | 105629 | 0.203 | 105672 | 0.203
105764 | 0.201 | 105769 | 0.202 | 105746 | 0.202 | 105746 | 0.203 | 105789 | 0.203
105882 | 0.201 | 105886 | 0.202 | 105864 | 0.201 | 105864 | 0.203 | 105907 | 0.203
106014 | 0.201 | 106004 | 0.201 | 105981 | 0.202 | 105981 | 0.202 | 106024 | 0.203
106132 | 0.201 | 106122 | 0.202 | 106099 | 0.202 | 106099 | 0.203 | 106142 | 0.203
106250 | 0.201 | 106240 | 0.202 | 106217 | 0.202 | 106216 | 0.203 | 106259 | 0.203
106368 | 0.201 | 106356 | 0.202 | 106334 | 0.202 | 106334 | 0.203 | 106377 | 0.203
106484 | 0.201 | 106474 | 0.202 | 106451 | 0.202 | 106451 | 0.203 | 106494 | 0.203
106602 | 0.201 | 106592 | 0.202 | 106569 | 0.202 | 106569 | 0.203 | 106612 | 0.203
106720 | 0.201 | 106710 | 0.202 | 106687 | 0.201 | 106686 | 0.203 | 106729 | 0.203
106838 | 0.201 | 106826 | 0.202 | 106804 | 0.201 | 106804 | 0.203 | 106847 | 0.203
106954 | 0.201 | 106944 | 0.201 | 106921 | 0.202 | 106921 | 0.203 | 106964 | 0.203
107072 | 0.201 | 107062 | 0.202 | 107039 | 0.202 | 107039 | 0.203 | 107082 | 0.203
107190 | 0.201 | 107180 | 0.201 | 107157 | 0.202 | 107156 | 0.202 | 107199 | 0.203
107308 | 0.201 | 107296 | 0.201 | 107274 | 0.202 | 107274 | 0.203 | 107317 | 0.203
107424 | 0.201 | 107414 | 0.201 | 107392 | 0.202 | 107392 | 0.203 | 107435 | 0.203
107542 | 0.201 | 107532 | 0.202 | 107509 | 0.202 | 107509 | 0.203 | 107552 | 0.203
107660 | 0.201 | 107650 | 0.202 | 107627 | 0.202 | 107626 | 0.203 | 107669 | 0.203
107778 | 0.201 | 107766 | 0.202 | 107744 | 0.202 | 107744 | 0.203 | 107787 | 0.203
107895 | 0.201 | 107884 | 0.202 | 107862 | 0.202 | 107862 | 0.203 | 107905 | 0.203
108012 | 0.201 | 108002 | 0.202 | 107979 | 0.202 | 107979 | 0.203 | 108022 | 0.203
108130 | 0.201 | 108120 | 0.201 | 108097 | 0.201 | 108096 | 0.203 | 108140 | 0.203
108248 | 0.201 | 108236 | 0.202 | 108215 | 0.202 | 108214 | 0.203 | 108257 | 0.203
108365 | 0.201 | 108354 | 0.202 | 108332 | 0.202 | 108332 | 0.203 | 108375 | 0.202
108482 | 0.201 | 108472 | 0.202 | 108449 | 0.202 | 108449 | 0.203 | 108492 | 0.203
108600 | 0.201 | 108590 | 0.202 | 108567 | 0.202 | 108566 | 0.203 | 108610 | 0.203
108718 | 0.201 | 108706 | 0.202 | 108685 | 0.201 | 108684 | 0.203 | 108727 | 0.203
108835 | 0.201 | 108824 | 0.202 | 108802 | 0.201 | 108802 | 0.203 | 108845 | 0.203
108952 | 0.201 | 108942 | 0.201 | 108919 | 0.202 | 108919 | 0.203 | 108962 | 0.203
109070 | 0.201 | 109060 | 0.201 | 109037 | 0.202 | 109037 | 0.203 | 109080 | 0.202
109188 | 0.201 | 109177 | 0.201 | 109155 | 0.202 | 109154 | 0.203 | 109197 | 0.203
109305 | 0.201 | 109294 | 0.201 | 109276 | 0.202 | 109272 | 0.203 | 109315 | 0.203
109422 | 0.201 | 109412 | 0.201 | 109392 | 0.202 | 109389 | 0.203 | 109432 | 0.203
109540 | 0.201 | 109530 | 0.201 | 109510 | 0.202 | 109507 | 0.203 | 109550 | 0.203
109661 | 0.201 | 109647 | 0.201 | 109628 | 0.202 | 109624 | 0.203 | 109671 | 0.203
109778 | 0.201 | 109764 | 0.201 | 109746 | 0.202 | 109742 | 0.203 | 109789 | 0.203
109896 | 0.201 | 109882 | 0.201 | 109863 | 0.202 | 109862 | 0.203 | 109905 | 0.203
110013 | 0.201 | 110003 | 0.202 | 109980 | 0.202 | 109980 | 0.203 | 110023 | 0.203
110131 | 0.201 | 110120 | 0.201 | 110098 | 0.202 | 110098 | 0.203 | 110141 | 0.203
110248 | 0.201 | 110238 | 0.201 | 110216 | 0.202 | 110215 | 0.202 | 110259 | 0.202
110366 | 0.201 | 110355 | 0.202 | 110333 | 0.202 | 110332 | 0.203 | 110375 | 0.202
110484 | 0.201 | 110473 | 0.202 | 110450 | 0.202 | 110450 | 0.203 | 110493 | 0.202
110601 | 0.201 | 110590 | 0.201 | 110568 | 0.202 | 110568 | 0.203 | 110611 | 0.203
110718 | 0.201 | 110708 | 0.201 | 110686 | 0.202 | 110685 | 0.203 | 110729 | 0.202
110836 | 0.201 | 110825 | 0.201 | 110803 | 0.201 | 110802 | 0.203 | 110845 | 0.202
110954 | 0.201 | 110943 | 0.201 | 110920 | 0.202 | 110920 | 0.202 | 110963 | 0.202
111071 | 0.201 | 111060 | 0.201 | 111038 | 0.202 | 111038 | 0.202 | 111081 | 0.203
111188 | 0.201 | 111178 | 0.201 | 111156 | 0.202 | 111156 | 0.202 | 111199 | 0.202
111306 | 0.201 | 111296 | 0.201 | 111273 | 0.202 | 111272 | 0.202 | 111315 | 0.203
111424 | 0.201 | 111413 | 0.201 | 111390 | 0.201 | 111390 | 0.203 | 111433 | 0.202
111541 | 0.202 | 111531 | 0.201 | 111508 | 0.201 | 111508 | 0.202 | 111551 | 0.202
111659 | 0.202 | 111648 | 0.202 | 111626 | 0.202 | 111626 | 0.202 | 111669 | 0.202
111776 | 0.202 | 111766 | 0.202 | 111744 | 0.202 | 111742 | 0.202 | 111785 | 0.202
111894 | 0.202 | 111883 | 0.201 | 111860 | 0.201 | 111860 | 0.203 | 111903 | 0.202
112011 | 0.202 | 112001 | 0.201 | 111978 | 0.201 | 111978 | 0.203 | 112021 | 0.202
112129 | 0.202 | 112118 | 0.201 | 112096 | 0.201 | 112096 | 0.202 | 112159 | 0.202
112246 | 0.202 | 112236 | 0.201 | 112214 | 0.201 | 112212 | 0.203 | 112277 | 0.203
112364 | 0.202 | 112353 | 0.201 | 112331 | 0.201 | 112330 | 0.202 | 112395 | 0.202
112481 | 0.202 | 112471 | 0.201 | 112448 | 0.201 | 112448 | 0.202 | 112513 | 0.202
112599 | 0.202 | 112588 | 0.201 | 112566 | 0.201 | 112566 | 0.202 | 112629 | 0.202
112716 | 0.202 | 112706 | 0.201 | 112684 | 0.201 | 112683 | 0.202 | 112747 | 0.202
112834 | 0.202 | 112823 | 0.202 | 112801 | 0.201 | 112800 | 0.202 | 112865 | 0.202
112952 | 0.202 | 112941 | 0.201 | 112918 | 0.201 | 112918 | 0.202 | 112983 | 0.202
113069 | 0.202 | 113058 | 0.201 | 113036 | 0.201 | 113036 | 0.202 | 113099 | 0.202
113186 | 0.202 | 113176 | 0.201 | 113154 | 0.201 | 113153 | 0.202 | 113217 | 0.202
113304 | 0.202 | 113293 | 0.201 | 113271 | 0.201 | 113270 | 0.202 | 113335 | 0.202
113422 | 0.202 | 113411 | 0.201 | 113389 | 0.201 | 113388 | 0.202 | 113453 | 0.202
113539 | 0.202 | 113528 | 0.201 | 113506 | 0.201 | 113506 | 0.202 | 113570 | 0.202
113656 | 0.202 | 113646 | 0.201 | 113628 | 0.201 | 113624 | 0.202 | 113687 | 0.202
113774 | 0.202 | 113764 | 0.201 | 113756 | 0.201 | 113740 | 0.202 | 113805 | 0.202
113895 | 0.202 | 113881 | 0.201 | 113872 | 0.201 | 113858 | 0.202 | 113926 | 0.202
114023 | 0.202 | 113998 | 0.201 | 113990 | 0.201 | 113976 | 0.202 | 114054 | 0.202
114141 | 0.202 | 114116 | 0.201 | 114108 | 0.201 | 114094 | 0.202 | 114171 | 0.202
114258 | 0.202 | 114234 | 0.201 | 114226 | 0.201 | 114214 | 0.202 | 114289 | 0.202
114376 | 0.202 | 114355 | 0.201 | 114343 | 0.201 | 114342 | 0.202 | 114407 | 0.202
114493 | 0.202 | 114483 | 0.201 | 114460 | 0.201 | 114480 | 0.202 | 114525 | 0.202
114611 | 0.202 | 114600 | 0.201 | 114578 | 0.201 | 114597 | 0.202 | 114641 | 0.202
114728 | 0.202 | 114718 | 0.201 | 114696 | 0.201 | 114714 | 0.202 | 114759 | 0.202
114846 | 0.202 | 114836 | 0.201 | 114813 | 0.201 | 114832 | 0.202 | 114877 | 0.202
114964 | 0.202 | 114953 | 0.201 | 114930 | 0.201 | 114950 | 0.202 | 114995 | 0.202
115081 | 0.202 | 115070 | 0.201 | 115048 | 0.201 | 115067 | 0.202 | 115111 | 0.202
115198 | 0.202 | 115188 | 0.201 | 115166 | 0.201 | 115184 | 0.202 | 115229 | 0.202
115316 | 0.202 | 115306 | 0.201 | 115283 | 0.201 | 115302 | 0.202 | 115347 | 0.202
115434 | 0.202 | 115424 | 0.201 | 115401 | 0.201 | 115420 | 0.202 | 115465 | 0.203
115552 | 0.202 | 115540 | 0.201 | 115518 | 0.201 | 115537 | 0.202 | 115582 | 0.203
115668 | 0.202 | 115658 | 0.201 | 115636 | 0.201 | 115655 | 0.202 | 115699 | 0.203
115786 | 0.202 | 115776 | 0.201 | 115753 | 0.201 | 115772 | 0.202 | 115817 | 0.203
115904 | 0.202 | 115894 | 0.202 | 115871 | 0.201 | 115890 | 0.202 | 115935 | 0.203
116022 | 0.202 | 116011 | 0.201 | 116006 | 0.201 | 116007 | 0.202 | 116052 | 0.203
116139 | 0.202 | 116128 | 0.202 | 116124 | 0.201 | 116125 | 0.202 | 116170 | 0.203
116256 | 0.202 | 116246 | 0.201 | 116242 | 0.201 | 116242 | 0.203 | 116287 | 0.203
116374 | 0.202 | 116364 | 0.202 | 116359 | 0.201 | 116360 | 0.202 | 116405 | 0.203
116492 | 0.202 | 116482 | 0.201 | 116476 | 0.201 | 116477 | 0.202 | 116523 | 0.203
116610 | 0.202 | 116598 | 0.201 | 116594 | 0.201 | 116595 | 0.202 | 116640 | 0.202
116726 | 0.202 | 116716 | 0.201 | 116712 | 0.201 | 116713 | 0.202 | 116757 | 0.202
116844 | 0.202 | 116834 | 0.202 | 116829 | 0.201 | 116830 | 0.202 | 116875 | 0.203
116962 | 0.202 | 116952 | 0.202 | 116946 | 0.201 | 116947 | 0.202 | 116993 | 0.203
117080 | 0.202 | 117068 | 0.202 | 117064 | 0.201 | 117065 | 0.202 | 117110 | 0.203
117196 | 0.202 | 117186 | 0.202 | 117182 | 0.201 | 117183 | 0.202 | 117228 | 0.203
117314 | 0.202 | 117304 | 0.202 | 117300 | 0.201 | 117301 | 0.202 | 117345 | 0.202
117432 | 0.202 | 117422 | 0.201 | 117416 | 0.201 | 117417 | 0.202 | 117463 | 0.202
117550 | 0.202 | 117539 | 0.201 | 117534 | 0.201 | 117535 | 0.202 | 117580 | 0.202
117667 | 0.202 | 117656 | 0.202 | 117652 | 0.201 | 117653 | 0.202 | 117698 | 0.202
117784 | 0.202 | 117774 | 0.202 | 117770 | 0.201 | 117771 | 0.202 | 117816 | 0.203
117902 | 0.202 | 117892 | 0.201 | 117891 | 0.201 | 117888 | 0.202 | 117933 | 0.202
118020 | 0.202 | 118009 | 0.202 | 118008 | 0.201 | 118005 | 0.202 | 118050 | 0.202
118137 | 0.202 | 118127 | 0.202 | 118125 | 0.201 | 118123 | 0.202 | 118168 | 0.202
118258 | 0.202 | 118244 | 0.202 | 118243 | 0.201 | 118241 | 0.202 | 118289 | 0.202
118376 | 0.202 | 118378 | 0.201 | 118361 | 0.201 | 118359 | 0.202 | 118407 | 0.202
118493 | 0.202 | 118496 | 0.202 | 118478 | 0.201 | 118479 | 0.202 | 118524 | 0.202
118610 | 0.202 | 118617 | 0.202 | 118596 | 0.201 | 118597 | 0.202 | 118641 | 0.203
118728 | 0.202 | 118734 | 0.202 | 118713 | 0.201 | 118714 | 0.202 | 118759 | 0.202
118846 | 0.202 | 118852 | 0.202 | 118831 | 0.201 | 118832 | 0.202 | 118877 | 0.202
118964 | 0.202 | 118969 | 0.202 | 118948 | 0.201 | 118949 | 0.202 | 118994 | 0.202
119080 | 0.202 | 119087 | 0.201 | 119066 | 0.201 | 119067 | 0.202 | 119112 | 0.202
119198 | 0.202 | 119204 | 0.202 | 119183 | 0.201 | 119184 | 0.202 | 119229 | 0.202
119316 | 0.202 | 119322 | 0.202 | 119301 | 0.201 | 119302 | 0.202 | 119347 | 0.202
119449 | 0.202 | 119440 | 0.202 | 119418 | 0.201 | 119419 | 0.202 | 119464 | 0.202
119567 | 0.202 | 119557 | 0.202 | 119536 | 0.201 | 119537 | 0.202 | 119582 | 0.202
119684 | 0.202 | 119674 | 0.202 | 119654 | 0.202 | 119655 | 0.202 | 119699 | 0.202
119801 | 0.202 | 119792 | 0.202 | 119771 | 0.201 | 119772 | 0.202 | 119817 | 0.202
119919 | 0.202 | 119910 | 0.202 | 119888 | 0.201 | 119889 | 0.202 | 119934 | 0.202
120037 | 0.202 | 120027 | 0.202 | 120006 | 0.201 | 120007 | 0.202 | 120052 | 0.202
120154 | 0.202 | 120144 | 0.202 | 120124 | 0.201 | 120125 | 0.202 | 120170 | 0.202
120271 | 0.202 | 120262 | 0.202 | 120242 | 0.201 | 120242 | 0.202 | 120287 | 0.203
120389 | 0.202 | 120380 | 0.202 | 120358 | 0.201 | 120359 | 0.202 | 120404 | 0.202
120507 | 0.202 | 120498 | 0.202 | 120476 | 0.201 | 120477 | 0.202 | 120522 | 0.203
120625 | 0.202 | 120614 | 0.202 | 120594 | 0.202 | 120595 | 0.202 | 120640 | 0.203
120741 | 0.202 | 120732 | 0.202 | 120712 | 0.201 | 120713 | 0.202 | 120758 | 0.202
120859 | 0.202 | 120850 | 0.202 | 120828 | 0.202 | 120829 | 0.202 | 120875 | 0.202
120977 | 0.191 | 120968 | 0.208 | 120946 | 0.202 | 120947 | 0.202 | 120992 | 0.203
121095 | 0.11 | 121084 | 0.14 | 121064 | 0.204 | 121065 | 0.202 | 121110 | 0.159
121212 | 0.002 | 121202 | 0.022 | 121182 | 0.187 | 121183 | 0.202 | 121228 | 0.031
121329 | 0.002 | 121320 | 0.001 | 121299 | 0.09 | 121300 | 0.2 | 121346 | 0.003
121447 | 0.002 | 121438 | 0.001 | 121416 | 0.002 | 121417 | 0.174 | 121462 | 0.003
121565 | 0.002 | 121555 | 0.002 | 121534 | 0.002 | 121535 | 0.007 | 121580 | 0.002
121682 | 0.002 | 121672 | 0.002 | 121652 | 0.002 | 121653 | 0.002 | 121698 | 0.003
121799 | 0.002 | 121790 | 0.002 | 121769 | 0.002 | 121771 | 0.002 | 121816 | 0.003
121917 | 0.002 | 121908 | 0.001 | 121887 | 0.002 | 121887 | 0.002 | 121932 | 0.002
122035 | 0.002 | 122025 | 0.001 | 122004 | 0.002 | 122005 | 0.002 | 122050 | 0.002
122152 | 0.002 | 122142 | 0.001 | 122122 | 0.002 | 122123 | 0.002 | 122168 | 0.002
122269 | 0.002 | 122260 | 0.001 | 122243 | 0.002 | 122241 | 0.002 | 122286 | 0.002
122387 | 0.002 | 122378 | 0.001 | 122360 | 0.002 | 122358 | 0.002 | 122402 | 0.002
122508 | 0.002 | 122496 | 0.001 | 122478 | 0.002 | 122475 | 0.002 | 122524 | 0.002
122626 | 0.001 | 122613 | 0.001 | 122596 | 0.002 | 122593 | 0.002 | 122641 | 0.003
122743 | 0.002 | 122730 | 0.001 | 122713 | 0.002 | 122711 | 0.002 | 122759 | 0.003
122861 | 0.002 | 122848 | 0.001 | 122830 | 0.002 | 122831 | 0.002 | 122876 | 0.002
122978 | 0.002 | 122969 | 0.001 | 122948 | 0.002 | 122949 | 0.002 | 122994 | 0.003
123096 | 0.002 | 123086 | 0.001 | 123066 | 0.002 | 123066 | 0.002 | 123111 | 0.002
123213 | 0.002 | 123204 | 0.001 | 123183 | 0.002 | 123184 | 0.002 | 123229 | 0.002
123331 | 0.002 | 123322 | 0.001 | 123300 | 0.002 | 123301 | 0.002 | 123346 | 0.002
123449 | 0.002 | 123439 | 0.001 | 123418 | 0.002 | 123419 | 0.002 | 123464 | 0.002
123566 | 0.001 | 123556 | 0.001 | 123536 | 0.001 | 123537 | 0.002 | 123582 | 0.002
123683 | 0.002 | 123674 | 0.001 | 123653 | 0.002 | 123654 | 0.002 | 123699 | 0.002
123801 | 0.001 | 123792 | 0.001 | 123770 | 0.002 | 123771 | 0.002 | 123816 | 0.002
123919 | 0.002 | 123909 | 0.002 | 123888 | 0.002 | 123889 | 0.002 | 123934 | 0.003
124036 | 0.002 | 124026 | 0.001 | 124006 | 0.002 | 124007 | 0.002 | 124052 | 0.002
124153 | 0.002 | 124144 | 0.001 | 124124 | 0.002 | 124124 | 0.002 | 124169 | 0.003
124271 | 0.002 | 124262 | 0.001 | 124240 | 0.002 | 124241 | 0.002 | 124286 | 0.003
124389 | 0.002 | 124380 | 0.001 | 124358 | 0.002 | 124359 | 0.002 | 124404 | 0.002
124507 | 0.002 | 124496 | 0.001 | 124476 | 0.002 | 124477 | 0.002 | 124522 | 0.003
124623 | 0.002 | 124614 | 0.001 | 124594 | 0.002 | 124595 | 0.002 | 124640 | 0.002
124741 | 0.002 | 124732 | 0.001 | 124710 | 0.002 | 124711 | 0.002 | 124778 | 0.003
124859 | 0.002 | 124850 | 0.001 | 124828 | 0.002 | 124829 | 0.002 | 124896 | 0.003
124977 | 0.002 | 124966 | 0.001 | 124946 | 0.002 | 124947 | 0.002 | 125012 | 0.003
125093 | 0.002 | 125084 | 0.001 | 125064 | 0.002 | 125065 | 0.002 | 125130 | 0.002
125211 | 0.002 | 125202 | 0.002 | 125181 | 0.002 | 125181 | 0.002 | 125248 | 0.003
125329 | 0.002 | 125320 | 0.001 | 125298 | 0.002 | 125299 | 0.002 | 125366 | 0.003
125447 | 0.002 | 125437 | 0.001 | 125416 | 0.002 | 125417 | 0.002 | 125483 | 0.002
125564 | 0.002 | 125554 | 0.001 | 125534 | 0.002 | 125535 | 0.002 | 125600 | 0.002
125681 | 0.002 | 125672 | 0.001 | 125651 | 0.002 | 125652 | 0.002 | 125718 | 0.002
125799 | 0.002 | 125790 | 0.001 | 125768 | 0.002 | 125769 | 0.002 | 125836 | 0.002
125917 | 0.002 | 125907 | 0.001 | 125886 | 0.002 | 125887 | 0.002 | 125954 | 0.003
126034 | 0.002 | 126024 | 0.001 | 126004 | 0.002 | 126005 | 0.002 | 126070 | 0.003
126151 | 0.002 | 126142 | 0.002 | 126121 | 0.002 | 126123 | 0.002 | 126188 | 0.003
126269 | 0.002 | 126260 | 0.002 | 126239 | 0.002 | 126239 | 0.002 | 126306 | 0.003
126387 | 0.002 | 126377 | 0.002 | 126356 | 0.002 | 126357 | 0.002 | 126424 | 0.003
126504 | 0.002 | 126494 | 0.001 | 126477 | 0.002 | 126475 | 0.002 | 126540 | 0.003
126621 | 0.002 | 126612 | 0.001 | 126595 | 0.002 | 126593 | 0.002 | 126658 | 0.003
126739 | 0.002 | 126730 | 0.001 | 126712 | 0.002 | 126710 | 0.002 | 126776 | 0.003
126860 | 0.002 | 126848 | 0.001 | 126830 | 0.002 | 126827 | 0.002 | 126897 | 0.003
126978 | 0.002 | 126965 | 0.002 | 126948 | 0.002 | 126945 | 0.002 | 127014 | 0.002
127095 | 0.002 | 127082 | 0.002 | 127065 | 0.002 | 127066 | 0.002 | 127132 | 0.002
127213 | 0.002 | 127203 | 0.002 | 127182 | 0.002 | 127183 | 0.002 | 127249 | 0.003
127330 | 0.002 | 127321 | 0.001 | 127300 | 0.002 | 127321 | 0.003 | 127367 | 0.003
127448 | 0.002 | 127438 | 0.001 | 127418 | 0.002 | 127437 | 0.003 | 127484 | 0.003
127565 | 0.002 | 127556 | 0.002 | 127535 | 0.002 | 127555 | 0.003 | 127602 | 0.002
127683 | 0.002 | 127674 | 0.001 | 127652 | 0.002 | 127673 | 0.002 | 127720 | 0.003
127801 | 0.002 | 127791 | 0.002 | 127770 | 0.002 | 127791 | 0.002 | 127837 | 0.003
127918 | 0.002 | 127908 | 0.002 | 127888 | 0.002 | 127908 | 0.003 | 127954 | 0.002
128035 | 0.002 | 128026 | 0.002 | 128006 | 0.002 | 128025 | 0.003 | 128072 | 0.002
128153 | 0.002 | 128144 | 0.002 | 128122 | 0.002 | 128143 | 0.003 | 128190 | 0.002
128271 | 0.002 | 128261 | 0.001 | 128240 | 0.002 | 128261 | 0.003 | 128307 | 0.002
128388 | 0.002 | 128378 | 0.001 | 128358 | 0.002 | 128378 | 0.003 | 128424 | 0.003
128505 | 0.002 | 128496 | 0.001 | 128476 | 0.002 | 128495 | 0.002 | 128542 | 0.003
128623 | 0.002 | 128614 | 0.002 | 128592 | 0.002 | 128613 | 0.002 | 128660 | 0.003
128741 | 0.002 | 128732 | 0.001 | 128710 | 0.002 | 128731 | 0.002 | 128778 | 0.002
128859 | 0.002 | 128848 | 0.002 | 128828 | 0.002 | 128848 | 0.002 | 128894 | 0.002
128975 | 0.002 | 128966 | 0.001 | 128946 | 0.002 | 128965 | 0.003 | 129012 | 0.002
129093 | 0.002 | 129084 | 0.001 | 129081 | 0.002 | 129083 | 0.003 | 129130 | 0.002
129211 | 0.002 | 129202 | 0.002 | 129199 | 0.002 | 129201 | 0.002 | 129248 | 0.002
129329 | 0.002 | 129318 | 0.002 | 129315 | 0.002 | 129318 | 0.002 | 129364 | 0.002
129445 | 0.002 | 129436 | 0.002 | 129433 | 0.002 | 129436 | 0.003 | 129482 | 0.002
129563 | 0.002 | 129554 | 0.002 | 129551 | 0.002 | 129553 | 0.002 | 129600 | 0.002
129681 | 0.002 | 129672 | 0.002 | 129669 | 0.002 | 129671 | 0.002 | 129718 | 0.002
129799 | 0.002 | 129789 | 0.002 | 129786 | 0.002 | 129788 | 0.003 | 129836 | 0.002
129915 | 0.003 | 129906 | 0.002 | 129903 | 0.002 | 129906 | 0.003 | 129952 | 0.002
130033 | 0.002 | 130024 | 0.001 | 130021 | 0.003 | 130023 | 0.003 | 130070 | 0.003
130151 | 0.002 | 130142 | 0.002 | 130139 | 0.003 | 130141 | 0.003 | 130188 | 0.002
130269 | 0.002 | 130259 | 0.002 | 130256 | 0.002 | 130258 | 0.002 | 130306 | 0.002
130386 | 0.002 | 130376 | 0.002 | 130373 | 0.002 | 130376 | 0.002 | 130422 | 0.002
130503 | 0.002 | 130494 | 0.002 | 130491 | 0.002 | 130494 | 0.002 | 130540 | 0.002
130621 | 0.002 | 130612 | 0.002 | 130609 | 0.002 | 130611 | 0.002 | 130658 | 0.002
130739 | 0.002 | 130729 | 0.002 | 130726 | 0.002 | 130728 | 0.002 | 130776 | 0.002
130856 | 0.002 | 130847 | 0.002 | 130847 | 0.002 | 130846 | 0.002 | 130893 | 0.002
130973 | 0.002 | 130964 | 0.002 | 130965 | 0.002 | 130964 | 0.002 | 131010 | 0.002
131095 | 0.002 | 131082 | 0.002 | 131082 | 0.002 | 131081 | 0.003 | 131131 | 0.002
131212 | 0.002 | 131200 | 0.002 | 131200 | 0.002 | 131198 | 0.002 | 131249 | 0.002
131330 | 0.002 | 131317 | 0.002 | 131317 | 0.002 | 131316 | 0.002 | 131366 | 0.002
131447 | 0.002 | 131434 | 0.001 | 131435 | 0.003 | 131437 | 0.002 | 131484 | 0.002
131565 | 0.002 | 131572 | 0.001 | 131553 | 0.002 | 131555 | 0.003 | 131602 | 0.002
131682 | 0.002 | 131689 | 0.002 | 131670 | 0.003 | 131673 | 0.002 | 131719 | 0.002
131800 | 0.002 | 131807 | 0.002 | 131787 | 0.002 | 131789 | 0.002 | 131836 | 0.002
131917 | 0.002 | 131924 | 0.002 | 131905 | 0.003 | 131907 | 0.002 | 131954 | 0.002
132035 | 0.002 | 132042 | 0.002 | 132023 | 0.003 | 132025 | 0.002 | 132072 | 0.002
132153 | 0.002 | 132159 | 0.002 | 132140 | 0.003 | 132143 | 0.002 | 132189 | 0.002
132270 | 0.002 | 132277 | 0.001 | 132257 | 0.002 | 132260 | 0.002 | 132306 | 0.002
132387 | 0.002 | 132395 | 0.001 | 132375 | 0.002 | 132377 | 0.002 | 132424 | 0.002
132505 | 0.002 | 132512 | 0.002 | 132493 | 0.002 | 132495 | 0.002 | 132542 | 0.002
132623 | 0.002 | 132630 | 0.001 | 132611 | 0.002 | 132613 | 0.002 | 132660 | 0.002
132755 | 0.002 | 132747 | 0.001 | 132727 | 0.002 | 132730 | 0.002 | 132776 | 0.002
132873 | 0.002 | 132865 | 0.001 | 132845 | 0.002 | 132847 | 0.002 | 132894 | 0.002
132990 | 0.002 | 132982 | 0.001 | 132963 | 0.002 | 132965 | 0.002 | 133012 | 0.002
133108 | 0.002 | 133100 | 0.002 | 133081 | 0.002 | 133083 | 0.002 | 133130 | 0.002
133226 | 0.002 | 133217 | 0.002 | 133197 | 0.002 | 133200 | 0.002 | 133246 | 0.002
133343 | 0.002 | 133335 | 0.001 | 133315 | 0.002 | 133318 | 0.002 | 133364 | 0.002
133460 | 0.002 | 133452 | 0.001 | 133433 | 0.002 | 133435 | 0.003 | 133482 | 0.002
133578 | 0.002 | 133570 | 0.001 | 133551 | 0.002 | 133553 | 0.002 | 133600 | 0.002
133696 | 0.002 | 133688 | 0.001 | 133668 | 0.002 | 133670 | 0.002 | 133716 | 0.002
133813 | 0.002 | 133805 | 0.001 | 133785 | 0.002 | 133788 | 0.002 | 133834 | 0.002
133931 | 0.002 | 133922 | 0.001 | 133903 | 0.002 | 133905 | 0.002 | 133952 | 0.002
134048 | 0.002 | 134040 | 0.001 | 134021 | 0.002 | 134023 | 0.002 | 134070 | 0.002
134166 | 0.002 | 134158 | 0.001 | 134138 | 0.002 | 134140 | 0.002 | 134188 | 0.002
134284 | 0.002 | 134275 | 0.001 | 134255 | 0.002 | 134258 | 0.003 | 134304 | 0.002
134401 | 0.002 | 134392 | 0.001 | 134373 | 0.002 | 134375 | 0.002 | 134422 | 0.002
134518 | 0.002 | 134510 | 0.001 | 134491 | 0.002 | 134493 | 0.002 | 134540 | 0.002
134636 | 0.002 | 134628 | 0.001 | 134608 | 0.002 | 134610 | 0.002 | 134658 | 0.002
134754 | 0.002 | 134746 | 0.001 | 134726 | 0.002 | 134728 | 0.002 | 134774 | 0.002
134871 | 0.002 | 134862 | 0.002 | 134843 | 0.002 | 134846 | 0.002 | 134892 | 0.002
134988 | 0.002 | 134980 | 0.001 | 134961 | 0.002 | 134963 | 0.002 | 135010 | 0.002
135106 | 0.002 | 135098 | 0.001 | 135078 | 0.002 | 135080 | 0.002 | 135128 | 0.002
135224 | 0.002 | 135216 | 0.001 | 135193 | 0.002 | 135198 | 0.002 | 135245 | 0.002
135342 | 0.002 | 135332 | 0.001 | 135311 | 0.002 | 135316 | 0.002 | 135362 | 0.002
135462 | 0.002 | 135450 | 0.001 | 135429 | 0.002 | 135433 | 0.002 | 135483 | 0.002
135579 | 0.002 | 135568 | 0.001 | 135546 | 0.002 | 135550 | 0.002 | 135601 | 0.002
135697 | 0.002 | 135686 | 0.001 | 135663 | 0.002 | 135671 | 0.002 | 135718 | 0.002
135815 | 0.002 | 135806 | 0.001 | 135781 | 0.002 | 135789 | 0.002 | 135836 | 0.002
135932 | 0.002 | 135924 | 0.001 | 135899 | 0.002 | 135907 | 0.002 | 135954 | 0.002
136049 | 0.002 | 136041 | 0.001 | 136016 | 0.002 | 136025 | 0.002 | 136071 | 0.002
136167 | 0.002 | 136159 | 0.001 | 136134 | 0.002 | 136141 | 0.002 | 136188 | 0.002
136285 | 0.002 | 136276 | 0.001 | 136251 | 0.002 | 136259 | 0.002 | 136306 | 0.002
136402 | 0.002 | 136394 | 0.001 | 136369 | 0.002 | 136377 | 0.002 | 136424 | 0.002
136520 | 0.002 | 136511 | 0.001 | 136486 | 0.002 | 136495 | 0.002 | 136541 | 0.002
136637 | 0.002 | 136629 | 0.001 | 136604 | 0.002 | 136612 | 0.002 | 136658 | 0.002
136755 | 0.002 | 136747 | 0.001 | 136721 | 0.002 | 136729 | 0.002 | 136776 | 0.002
136872 | 0.002 | 136864 | 0.001 | 136839 | 0.002 | 136847 | 0.002 | 136894 | 0.002
136990 | 0.002 | 136982 | 0.001 | 136956 | 0.002 | 136965 | 0.002 | 137012 | 0.002
137107 | 0.002 | 137099 | 0.001 | 137074 | 0.002 | 137082 | 0.002 | 137128 | 0.002
137225 | 0.002 | 137217 | 0.001 | 137192 | 0.002 | 137199 | 0.002 | 137268 | 0.002
137342 | 0.002 | 137334 | 0.001 | 137309 | 0.002 | 137317 | 0.002 | 137384 | 0.002
137460 | 0.002 | 137452 | 0.001 | 137426 | 0.002 | 137435 | 0.002 | 137502 | 0.002
137578 | 0.002 | 137569 | 0.001 | 137544 | 0.002 | 137552 | 0.002 | 137620 | 0.002
137695 | 0.002 | 137687 | 0.001 | 137662 | 0.002 | 137669 | 0.002 | 137738 | 0.002
137812 | 0.002 | 137804 | 0.001 | 137779 | 0.002 | 137787 | 0.002 | 137854 | 0.002
137930 | 0.002 | 137922 | 0.001 | 137896 | 0.002 | 137905 | 0.002 | 137972 | 0.002
138048 | 0.002 | 138040 | 0.001 | 138014 | 0.002 | 138022 | 0.002 | 138090 | 0.002
138165 | 0.003 | 138157 | 0.001 | 138132 | 0.002 | 138140 | 0.002 | 138208 | 0.002
138283 | 0.002 | 138274 | 0.001 | 138250 | 0.002 | 138257 | 0.002 | 138324 | 0.002
138400 | 0.002 | 138392 | 0.001 | 138366 | 0.002 | 138375 | 0.002 | 138442 | 0.002
138518 | 0.002 | 138510 | 0.001 | 138484 | 0.002 | 138492 | 0.002 | 138560 | 0.002
138636 | 0.002 | 138627 | 0.001 | 138602 | 0.002 | 138610 | 0.002 | 138678 | 0.002
138753 | 0.002 | 138744 | 0.001 | 138720 | 0.002 | 138727 | 0.002 | 138796 | 0.002
138870 | 0.002 | 138862 | 0.001 | 138836 | 0.002 | 138845 | 0.002 | 138912 | 0.002
138988 | 0.002 | 138980 | 0.001 | 138954 | 0.002 | 138962 | 0.002 | 139030 | 0.002
139106 | 0.002 | 139098 | 0.001 | 139072 | 0.002 | 139080 | 0.002 | 139148 | 0.002
139223 | 0.002 | 139214 | 0.001 | 139190 | 0.002 | 139198 | 0.002 | 139266 | 0.002
139340 | 0.002 | 139332 | 0.001 | 139308 | 0.002 | 139315 | 0.002 | 139382 | 0.002
139458 | 0.002 | 139450 | 0.002 | 139427 | 0.002 | 139432 | 0.002 | 139500 | 0.002
139576 | 0.002 | 139568 | 0.001 | 139545 | 0.002 | 139550 | 0.002 | 139618 | 0.002
139693 | 0.002 | 139684 | 0.001 | 139663 | 0.002 | 139668 | 0.002 | 139736 | 0.002
139814 | 0.002 | 139802 | 0.001 | 139780 | 0.002 | 139785 | 0.002 | 139856 | 0.002
139931 | 0.002 | 139920 | 0.001 | 139897 | 0.002 | 139902 | 0.002 | 139973 | 0.002
140049 | 0.002 | 140038 | 0.001 | 140014 | 0.002 | 140043 | 0.002 | 140091 | 0.002
140167 | 0.002 | 140158 | 0.001 | 140132 | 0.002 | 140161 | 0.002 | 140208 | 0.002
140283 | 0.003 | 140276 | 0.001 | 140250 | 0.002 | 140278 | 0.002 | 140325 | 0.002
140401 | 0.002 | 140393 | 0.001 | 140366 | 0.002 | 140395 | 0.002 | 140443 | 0.002
140519 | 0.002 | 140511 | 0.001 | 140484 | 0.002 | 140513 | 0.002 | 140560 | 0.002
140637 | 0.002 | 140628 | 0.001 | 140601 | 0.002 | 140631 | 0.002 | 140678 | 0.002
140753 | 0.002 | 140745 | 0.001 | 140719 | 0.002 | 140748 | 0.002 | 140794 | 0.002
140871 | 0.002 | 140863 | 0.001 | 140836 | 0.002 | 140865 | 0.002 | 140912 | 0.002
140989 | 0.002 | 140981 | 0.001 | 140953 | 0.002 | 140983 | 0.002 | 141030 | 0.002
141106 | 0.002 | 141099 | 0.001 | 141071 | 0.002 | 141101 | 0.002 | 141147 | 0.002
141223 | 0.002 | 141215 | 0.001 | 141188 | 0.002 | 141217 | 0.002 | 141264 | 0.002
141341 | 0.002 | 141333 | 0.001 | 141305 | 0.002 | 141335 | 0.002 | 141382 | 0.002
141459 | 0.002 | 141451 | 0.001 | 141423 | 0.002 | 141453 | 0.002 | 141499 | 0.002
141576 | 0.002 | 141568 | 0.001 | 141540 | 0.002 | 141571 | 0.002 | 141617 | 0.002
141693 | 0.002 | 141685 | 0.001 | 141658 | 0.002 | 141687 | 0.002 | 141733 | 0.002
141811 | 0.002 | 141803 | 0.001 | 141774 | 0.002 | 141805 | 0.002 | 141851 | 0.002
141929 | 0.002 | 141921 | 0.001 | 141892 | 0.002 | 141923 | 0.002 | 141969 | 0.002
142046 | 0.002 | 142038 | 0.001 | 142010 | 0.002 | 142040 | 0.002 | 142086 | 0.002
142163 | 0.002 | 142155 | 0.001 | 142145 | 0.002 | 142157 | 0.002 | 142203 | 0.002
142281 | 0.002 | 142273 | 0.001 | 142262 | 0.002 | 142275 | 0.002 | 142320 | 0.002
142399 | 0.002 | 142391 | 0.001 | 142380 | 0.002 | 142393 | 0.002 | 142438 | 0.002
142516 | 0.002 | 142508 | 0.001 | 142497 | 0.002 | 142510 | 0.002 | 142556 | 0.002
142633 | 0.002 | 142625 | 0.001 | 142625 | 0.002 | 142627 | 0.002 | 142672 | 0.002
142751 | 0.002 | 142743 | 0.001 | 142742 | 0.002 | 142745 | 0.002 | 142790 | 0.002
142879 | 0.002 | 142861 | 0.001 | 142860 | 0.002 | 142863 | 0.002 | 142918 | 0.002
142997 | 0.002 | 142978 | 0.001 | 142977 | 0.002 | 142981 | 0.002 | 143036 | 0.002
143114 | 0.002 | 143095 | 0.001 | 143094 | 0.002 | 143097 | 0.002 | 143153 | 0.002
143232 | 0.002 | 143213 | 0.001 | 143212 | 0.002 | 143225 | 0.002 | 143270 | 0.002
143349 | 0.002 | 143331 | 0.001 | 143329 | 0.002 | 143343 | 0.002 | 143387 | 0.002
143467 | 0.002 | 143459 | 0.001 | 143447 | 0.002 | 143461 | 0.002 | 143505 | 0.002
143584 | 0.002 | 143576 | 0.001 | 143563 | 0.002 | 143579 | 0.002 | 143623 | 0.002
143702 | 0.002 | 143693 | 0.001 | 143681 | 0.002 | 143695 | 0.002 | 143739 | 0.002
143819 | 0.002 | 143811 | 0.001 | 143802 | 0.002 | 143813 | 0.002 | 143857 | 0.002
143936 | 0.002 | 143929 | 0.001 | 143919 | 0.002 | 143931 | 0.002 | 143974 | 0.002
144057 | 0.002 | 144046 | 0.001 | 144036 | 0.002 | 144048 | 0.002 | 144095 | 0.002
144175 | 0.002 | 144163 | 0.001 | 144154 | 0.002 | 144165 | 0.002 | 144212 | 0.002
144291 | 0.002 | 144281 | 0.001 | 144271 | 0.002 | 144283 | 0.002 | 144329 | 0.002
144409 | 0.002 | 144398 | 0.001 | 144389 | 0.002 | 144404 | 0.002 | 144447 | 0.002
144527 | 0.002 | 144519 | 0.001 | 144505 | 0.002 | 144521 | 0.002 | 144565 | 0.002
144644 | 0.002 | 144636 | 0.001 | 144623 | 0.002 | 144638 | 0.002 | 144681 | 0.002
144761 | 0.002 | 144771 | 0.001 | 144741 | 0.002 | 144756 | 0.002 | 144799 | 0.002
144879 | 0.002 | 144888 | 0.001 | 144858 | 0.002 | 144873 | 0.002 | 144916 | 0.002
144996 | 0.002 | 145005 | 0.001 | 144975 | 0.002 | 144991 | 0.002 | 145034 | 0.002
145114 | 0.002 | 145123 | 0.001 | 145092 | 0.002 | 145107 | 0.002 | 145152 | 0.002
145230 | 0.002 | 145241 | 0.001 | 145210 | 0.002 | 145225 | 0.002 | 145268 | 0.002
145348 | 0.002 | 145357 | 0.001 | 145328 | 0.002 | 145343 | 0.002 | 145386 | 0.002
145465 | 0.002 | 145475 | 0.001 | 145444 | 0.002 | 145460 | 0.002 | 145503 | 0.002
145583 | 0.002 | 145592 | 0.001 | 145562 | 0.002 | 145577 | 0.002 | 145620 | 0.002
145700 | 0.002 | 145710 | 0.001 | 145679 | 0.002 | 145694 | 0.002 | 145738 | 0.002
145817 | 0.002 | 145827 | 0.001 | 145797 | 0.002 | 145812 | 0.002 | 145855 | 0.002
145935 | 0.002 | 145944 | 0.001 | 145914 | 0.002 | 145930 | 0.002 | 145973 | 0.002
146052 | 0.002 | 146062 | 0.001 | 146031 | 0.002 | 146046 | 0.002 | 146089 | 0.002
146185 | 0.002 | 146179 | 0.001 | 146149 | 0.002 | 146164 | 0.002 | 146207 | 0.002
146302 | 0.002 | 146296 | 0.001 | 146266 | 0.002 | 146281 | 0.002 | 146325 | 0.002
146420 | 0.002 | 146414 | 0.001 | 146383 | 0.002 | 146399 | 0.002 | 146442 | 0.002
146537 | 0.002 | 146531 | 0.001 | 146501 | 0.002 | 146516 | 0.002 | 146560 | 0.002
146655 | 0.002 | 146649 | 0.001 | 146618 | 0.002 | 146633 | 0.002 | 146676 | 0.002
146771 | 0.002 | 146765 | 0.001 | 146736 | 0.002 | 146751 | 0.002 | 146794 | 0.002
146889 | 0.002 | 146883 | 0.001 | 146852 | 0.002 | 146868 | 0.002 | 146912 | 0.002
147007 | 0.002 | 147000 | 0.001 | 146970 | 0.002 | 146985 | 0.002 | 147028 | 0.002
147124 | 0.002 | 147118 | 0.001 | 147088 | 0.002 | 147103 | 0.002 | 147146 | 0.002
147241 | 0.002 | 147235 | 0.001 | 147205 | 0.002 | 147220 | 0.002 | 147263 | 0.002
147358 | 0.002 | 147352 | 0.001 | 147322 | 0.002 | 147338 | 0.002 | 147381 | 0.002
147476 | 0.002 | 147470 | 0.001 | 147439 | 0.002 | 147454 | 0.002 | 147499 | 0.002
147594 | 0.002 | 147588 | 0.001 | 147557 | 0.003 | 147572 | 0.002 | 147615 | 0.002
147710 | 0.002 | 147704 | 0.001 | 147675 | 0.003 | 147690 | 0.002 | 147733 | 0.002
147828 | 0.002 | 147822 | 0.001 | 147791 | 0.003 | 147807 | 0.002 | 147850 | 0.002
147945 | 0.002 | 147939 | 0.001 | 147909 | 0.003 | 147924 | 0.002 | 147968 | 0.002
148063 | 0.002 | 148057 | 0.001 | 148030 | 0.002 | 148041 | 0.002 | 148085 | 0.002
148180 | 0.002 | 148174 | 0.001 | 148147 | 0.002 | 148159 | 0.002 | 148202 | 0.002
148297 | 0.002 | 148291 | 0.001 | 148264 | 0.002 | 148277 | 0.002 | 148320 | 0.002
148418 | 0.002 | 148409 | 0.001 | 148382 | 0.002 | 148393 | 0.002 | 148441 | 0.002
148536 | 0.002 | 148526 | 0.001 | 148499 | 0.002 | 148511 | 0.002 | 148557 | 0.002
148653 | 0.002 | 148643 | 0.001 | 148617 | 0.002 | 148632 | 0.002 | 148675 | 0.002
148770 | 0.002 | 148764 | 0.001 | 148734 | 0.002 | 148749 | 0.002 | 148793 | 0.002
148888 | 0.002 | 148881 | 0.001 | 148852 | 0.003 | 148867 | 0.002 | 148910 | 0.002
149005 | 0.002 | 148999 | 0.001 | 148969 | 0.002 | 148984 | 0.002 | 149027 | 0.002
149123 | 0.002 | 149116 | 0.001 | 149087 | 0.002 | 149101 | 0.002 | 149145 | 0.002
149240 | 0.002 | 149234 | 0.001 | 149204 | 0.002 | 149219 | 0.002 | 149263 | 0.002
149358 | 0.002 | 149351 | 0.002 | 149321 | 0.002 | 149337 | 0.002 | 149380 | 0.002
149475 | 0.002 | 149469 | 0.001 | 149439 | 0.002 | 149453 | 0.002 | 149497 | 0.002
149593 | 0.002 | 149586 | 0.001 | 149557 | 0.002 | 149571 | 0.002 | 149615 | 0.002
149710 | 0.002 | 149703 | 0.001 | 149675 | 0.002 | 149689 | 0.002 | 149732 | 0.002
149827 | 0.002 | 149821 | 0.001 | 149791 | 0.002 | 149807 | 0.002 | 149871 | 0.002
149945 | 0.002 | 149939 | 0.001 | 149909 | 0.002 | 149923 | 0.002 | 149988 | 0.002
150063 | 0.002 | 150056 | 0.001 | 150027 | 0.002 | 150041 | 0.002 | 150106 | 0.002
150180 | 0.002 | 150173 | 0.001 | 150144 | 0.002 | 150159 | 0.002 | 150223 | 0.002
150297 | 0.002 | 150291 | 0.001 | 150261 | 0.002 | 150276 | 0.002 | 150341 | 0.002
150415 | 0.002 | 150409 | 0.001 | 150379 | 0.002 | 150393 | 0.002 | 150458 | 0.002
150533 | 0.002 | 150525 | 0.001 | 150496 | 0.002 | 150511 | 0.002 | 150576 | 0.002
150649 | 0.002 | 150643 | 0.001 | 150614 | 0.002 | 150629 | 0.002 | 150693 | 0.002
150767 | 0.002 | 150761 | 0.001 | 150731 | 0.002 | 150746 | 0.002 | 150810 | 0.002
150885 | 0.002 | 150878 | 0.001 | 150849 | 0.002 | 150863 | 0.002 | 150928 | 0.002
151002 | 0.002 | 150995 | 0.001 | 150966 | 0.002 | 150981 | 0.002 | 151046 | 0.002
151119 | 0.002 | 151113 | 0.001 | 151084 | 0.002 | 151098 | 0.002 | 151163 | 0.002
151237 | 0.002 | 151231 | 0.001 | 151201 | 0.002 | 151216 | 0.002 | 151280 | 0.002
151355 | 0.002 | 151348 | 0.001 | 151318 | 0.002 | 151333 | 0.002 | 151398 | 0.002
151472 | 0.002 | 151465 | 0.001 | 151436 | 0.002 | 151450 | 0.002 | 151516 | 0.002
151589 | 0.002 | 151583 | 0.001 | 151554 | 0.002 | 151568 | 0.002 | 151632 | 0.002
151707 | 0.001 | 151700 | 0.001 | 151670 | 0.002 | 151686 | 0.002 | 151750 | 0.002
151824 | 0.002 | 151818 | 0.001 | 151788 | 0.002 | 151803 | 0.002 | 151868 | 0.002
151942 | 0.002 | 151936 | 0.001 | 151906 | 0.002 | 151920 | 0.002 | 151985 | 0.002
152059 | 0.002 | 152052 | 0.001 | 152024 | 0.002 | 152038 | 0.002 | 152102 | 0.002
152176 | 0.002 | 152170 | 0.001 | 152140 | 0.002 | 152156 | 0.002 | 152220 | 0.003
152294 | 0.002 | 152288 | 0.001 | 152258 | 0.002 | 152272 | 0.002 | 152338 | 0.003
152412 | 0.002 | 152406 | 0.001 | 152379 | 0.002 | 152390 | 0.002 | 152455 | 0.003
152529 | 0.002 | 152522 | 0.001 | 152497 | 0.002 | 152508 | 0.002 | 152572 | 0.002
152649 | 0.002 | 152640 | 0.001 | 152613 | 0.002 | 152625 | 0.002 | 152693 | 0.003
152767 | 0.002 | 152758 | 0.001 | 152731 | 0.002 | 152742 | 0.002 | 152811 | 0.003
152885 | 0.002 | 152875 | 0.001 | 152849 | 0.002 | 152880 | 0.002 | 152928 | 0.003
153003 | 0.002 | 152992 | 0.001 | 152966 | 0.002 | 153000 | 0.002 | 153046 | 0.003
153119 | 0.002 | 153113 | 0.001 | 153083 | 0.002 | 153118 | 0.002 | 153163 | 0.002
153237 | 0.002 | 153231 | 0.001 | 153201 | 0.002 | 153235 | 0.002 | 153280 | 0.002
153355 | 0.002 | 153348 | 0.001 | 153319 | 0.002 | 153353 | 0.002 | 153398 | 0.002
153472 | 0.002 | 153465 | 0.001 | 153436 | 0.002 | 153470 | 0.002 | 153516 | 0.002
153589 | 0.002 | 153583 | 0.001 | 153553 | 0.002 | 153588 | 0.002 | 153632 | 0.002
153707 | 0.002 | 153701 | 0.001 | 153671 | 0.002 | 153705 | 0.002 | 153750 | 0.002
153825 | 0.002 | 153818 | 0.001 | 153788 | 0.002 | 153823 | 0.002 | 153868 | 0.002
153942 | 0.002 | 153935 | 0.001 | 153906 | 0.002 | 153940 | 0.002 | 153986 | 0.002
154059 | 0.002 | 154053 | 0.001 | 154023 | 0.002 | 154057 | 0.002 | 154102 | 0.002
154177 | 0.002 | 154170 | 0.001 | 154140 | 0.002 | 154175 | 0.002 | 154220 | 0.002
154294 | 0.002 | 154288 | 0.001 | 154258 | 0.002 | 154293 | 0.002 | 154338 | 0.002
154412 | 0.002 | 154405 | 0.001 | 154376 | 0.002 | 154409 | 0.002 | 154455 | 0.002
154529 | 0.002 | 154522 | 0.001 | 154493 | 0.002 | 154527 | 0.002 | 154572 | 0.002
154646 | 0.002 | 154640 | 0.001 | 154610 | 0.002 | 154645 | 0.002 | 154690 | 0.002
154764 | 0.002 | 154758 | 0.001 | 154728 | 0.002 | 154763 | 0.002 | 154808 | 0.003
154882 | 0.002 | 154875 | 0.001 | 154846 | 0.002 | 154880 | 0.002 | 154925 | 0.002
154999 | 0.002 | 154992 | 0.001 | 154962 | 0.002 | 154997 | 0.002 | 155042 | 0.002
155116 | 0.002 | 155110 | 0.001 | 155099 | 0.002 | 155115 | 0.002 | 155160 | 0.003
155234 | 0.002 | 155228 | 0.001 | 155216 | 0.002 | 155232 | 0.002 | 155277 | 0.003
155352 | 0.002 | 155344 | 0.001 | 155333 | 0.002 | 155350 | 0.002 | 155395 | 0.002
155468 | 0.002 | 155462 | 0.001 | 155451 | 0.002 | 155467 | 0.002 | 155512 | 0.002
155586 | 0.002 | 155580 | 0.001 | 155569 | 0.002 | 155584 | 0.002 | 155629 | 0.002
155704 | 0.002 | 155698 | 0.001 | 155685 | 0.002 | 155702 | 0.002 | 155747 | 0.003
155821 | 0.002 | 155814 | 0.001 | 155803 | 0.002 | 155820 | 0.002 | 155865 | 0.002
155938 | 0.002 | 155932 | 0.001 | 155921 | 0.002 | 155937 | 0.002 | 155982 | 0.002
156056 | 0.002 | 156050 | 0.001 | 156038 | 0.002 | 156054 | 0.002 | 156099 | 0.002
156174 | 0.002 | 156167 | 0.001 | 156155 | 0.002 | 156172 | 0.002 | 156217 | 0.002
156291 | 0.002 | 156284 | 0.001 | 156273 | 0.002 | 156290 | 0.002 | 156335 | 0.002
156408 | 0.002 | 156402 | 0.001 | 156391 | 0.002 | 156406 | 0.002 | 156451 | 0.002
156526 | 0.002 | 156519 | 0.001 | 156508 | 0.002 | 156524 | 0.002 | 156569 | 0.002
156643 | 0.002 | 156637 | 0.001 | 156628 | 0.002 | 156642 | 0.002 | 156687 | 0.002
156761 | 0.002 | 156754 | 0.001 | 156746 | 0.002 | 156759 | 0.002 | 156804 | 0.002
156878 | 0.002 | 156872 | 0.001 | 156864 | 0.002 | 156876 | 0.002 | 156921 | 0.002
156999 | 0.002 | 156989 | 0.001 | 156981 | 0.002 | 156994 | 0.002 | 157042 | 0.002
157116 | 0.002 | 157107 | 0.001 | 157099 | 0.002 | 157112 | 0.002 | 157160 | 0.002
157234 | 0.002 | 157224 | 0.001 | 157216 | 0.002 | 157233 | 0.002 | 157278 | 0.002
157351 | 0.002 | 157345 | 0.001 | 157333 | 0.002 | 157349 | 0.002 | 157394 | 0.002
157469 | 0.002 | 157462 | 0.001 | 157451 | 0.002 | 157467 | 0.002 | 157512 | 0.002
157586 | 0.002 | 157580 | 0.001 | 157569 | 0.002 | 157585 | 0.002 | 157630 | 0.002
157704 | 0.002 | 157698 | 0.001 | 157685 | 0.002 | 157702 | 0.002 | 157747 | 0.002
157821 | 0.002 | 157832 | 0.001 | 157803 | 0.002 | 157819 | 0.002 | 157864 | 0.002
157938 | 0.002 | 157949 | 0.001 | 157921 | 0.002 | 157937 | 0.002 | 157982 | 0.002
158056 | 0.002 | 158067 | 0.001 | 158039 | 0.002 | 158054 | 0.002 | 158099 | 0.003
158174 | 0.002 | 158184 | 0.001 | 158155 | 0.002 | 158172 | 0.002 | 158217 | 0.003
158290 | 0.002 | 158302 | 0.001 | 158273 | 0.002 | 158289 | 0.002 | 158334 | 0.002
158408 | 0.002 | 158419 | 0.001 | 158391 | 0.002 | 158407 | 0.002 | 158452 | 0.002
158526 | 0.002 | 158536 | 0.001 | 158508 | 0.002 | 158524 | 0.002 | 158569 | 0.002
158644 | 0.002 | 158654 | 0.001 | 158625 | 0.002 | 158642 | 0.002 | 158687 | 0.002
158760 | 0.002 | 158772 | 0.001 | 158743 | 0.002 | 158759 | 0.002 | 158804 | 0.002
158878 | 0.002 | 158889 | 0.001 | 158861 | 0.002 | 158876 | 0.002 | 158921 | 0.003
158996 | 0.002 | 159006 | 0.001 | 158978 | 0.002 | 158994 | 0.002 | 159039 | 0.002
159113 | 0.002 | 159124 | 0.001 | 159095 | 0.002 | 159112 | 0.002 | 159157 | 0.002
159230 | 0.002 | 159242 | 0.001 | 159213 | 0.002 | 159228 | 0.002 | 159273 | 0.002
159348 | 0.002 | 159358 | 0.001 | 159330 | 0.002 | 159346 | 0.002 | 159391 | 0.003
159481 | 0.002 | 159476 | 0.001 | 159448 | 0.002 | 159464 | 0.002 | 159509 | 0.002
159598 | 0.002 | 159594 | 0.001 | 159565 | 0.002 | 159582 | 0.002 | 159627 | 0.003
159716 | 0.002 | 159711 | 0.001 | 159682 | 0.002 | 159698 | 0.002 | 159744 | 0.002
159834 | 0.002 | 159828 | 0.001 | 159800 | 0.002 | 159816 | 0.002 | 159861 | 0.002
159951 | 0.002 | 159946 | 0.001 | 159918 | 0.002 | 159934 | 0.002 | 159979 | 0.002
160068 | 0.002 | 160064 | 0.001 | 160035 | 0.002 | 160051 | 0.002 | 160096 | 0.002
160186 | 0.002 | 160181 | 0.001 | 160152 | 0.002 | 160168 | 0.002 | 160214 | 0.002
160304 | 0.002 | 160299 | 0.001 | 160270 | 0.002 | 160286 | 0.002 | 160331 | 0.003
160420 | 0.002 | 160416 | 0.001 | 160388 | 0.002 | 160403 | 0.002 | 160448 | 0.003
160538 | 0.002 | 160533 | 0.001 | 160504 | 0.002 | 160521 | 0.002 | 160566 | 0.002
160656 | 0.002 | 160651 | 0.001 | 160622 | 0.002 | 160638 | 0.002 | 160684 | 0.003
160773 | 0.002 | 160769 | 0.001 | 160740 | 0.002 | 160756 | 0.002 | 160801 | 0.003
160890 | 0.002 | 160886 | 0.001 | 160857 | 0.002 | 160873 | 0.002 | 160918 | 0.003
161008 | 0.002 | 161003 | 0.001 | 160977 | 0.002 | 160991 | 0.002 | 161036 | 0.002
161126 | 0.002 | 161121 | 0.001 | 161095 | 0.002 | 161108 | 0.002 | 161154 | 0.003
161246 | 0.002 | 161239 | 0.001 | 161213 | 0.002 | 161225 | 0.002 | 161274 | 0.003
161363 | 0.002 | 161355 | 0.001 | 161330 | 0.002 | 161343 | 0.002 | 161391 | 0.003
161481 | 0.002 | 161473 | 0.001 | 161447 | 0.002 | 161461 | 0.002 | 161509 | 0.003
161599 | 0.002 | 161591 | 0.001 | 161565 | 0.002 | 161582 | 0.002 | 161627 | 0.003
161716 | 0.002 | 161712 | 0.001 | 161683 | 0.002 | 161698 | 0.002 | 161743 | 0.003
161833 | 0.002 | 161828 | 0.001 | 161800 | 0.002 | 161816 | 0.002 | 161861 | 0.003
161951 | 0.002 | 161946 | 0.001 | 161917 | 0.002 | 161934 | 0.002 | 161979 | 0.002
162068 | 0.002 | 162064 | 0.001 | 162035 | 0.002 | 162051 | 0.002 | 162097 | 0.003
162186 | 0.002 | 162182 | 0.001 | 162152 | 0.002 | 162168 | 0.002 | 162213 | 0.003
162304 | 0.002 | 162298 | 0.001 | 162270 | 0.002 | 162286 | 0.002 | 162353 | 0.003
162420 | 0.002 | 162416 | 0.001 | 162387 | 0.002 | 162404 | 0.002 | 162469 | 0.003
162538 | 0.002 | 162534 | 0.001 | 162505 | 0.002 | 162521 | 0.002 | 162587 | 0.003
162656 | 0.002 | 162651 | 0.001 | 162622 | 0.002 | 162638 | 0.002 | 162705 | 0.003
162774 | 0.002 | 162768 | 0.001 | 162740 | 0.002 | 162756 | 0.002 | 162823 | 0.003
162890 | 0.002 | 162886 | 0.001 | 162857 | 0.002 | 162873 | 0.002 | 162939 | 0.003
163008 | 0.002 | 163003 | 0.001 | 162974 | 0.002 | 162991 | 0.002 | 163057 | 0.003
163126 | 0.002 | 163121 | 0.001 | 163092 | 0.002 | 163108 | 0.002 | 163175 | 0.003
163243 | 0.002 | 163238 | 0.001 | 163210 | 0.002 | 163226 | 0.002 | 163292 | 0.003
163360 | 0.002 | 163356 | 0.001 | 163326 | 0.002 | 163343 | 0.002 | 163409 | 0.003
163478 | 0.002 | 163473 | 0.001 | 163444 | 0.002 | 163461 | 0.002 | 163527 | 0.003
163596 | 0.002 | 163591 | 0.001 | 163562 | 0.002 | 163578 | 0.002 | 163644 | 0.003
163713 | 0.002 | 163708 | 0.001 | 163680 | 0.002 | 163695 | 0.002 | 163762 | 0.003
163830 | 0.002 | 163825 | 0.001 | 163796 | 0.002 | 163813 | 0.002 | 163879 | 0.002
163948 | 0.002 | 163943 | 0.001 | 163914 | 0.002 | 163931 | 0.002 | 163997 | 0.003
164065 | 0.002 | 164061 | 0.001 | 164032 | 0.002 | 164047 | 0.002 | 164114 | 0.003
164183 | 0.002 | 164178 | 0.001 | 164149 | 0.002 | 164165 | 0.002 | 164232 | 0.003
164300 | 0.002 | 164295 | 0.001 | 164267 | 0.002 | 164283 | 0.002 | 164349 | 0.003
164417 | 0.002 | 164413 | 0.001 | 164384 | 0.002 | 164401 | 0.002 | 164466 | 0.002
164535 | 0.002 | 164531 | 0.001 | 164501 | 0.002 | 164517 | 0.002 | 164584 | 0.003
164653 | 0.002 | 164647 | 0.001 | 164619 | 0.002 | 164635 | 0.002 | 164702 | 0.003
164770 | 0.002 | 164765 | 0.001 | 164737 | 0.002 | 164753 | 0.002 | 164819 | 0.003
164887 | 0.002 | 164883 | 0.001 | 164854 | 0.002 | 164870 | 0.002 | 164936 | 0.003
165005 | 0.002 | 165001 | 0.001 | 164971 | 0.002 | 164987 | 0.002 | 165054 | 0.003
165123 | 0.002 | 165117 | 0.001 | 165089 | 0.002 | 165105 | 0.002 | 165172 | 0.003
165239 | 0.002 | 165235 | 0.001 | 165207 | 0.002 | 165222 | 0.002 | 165288 | 0.003
165357 | 0.002 | 165353 | 0.001 | 165321 | 0.002 | 165340 | 0.002 | 165406 | 0.003
165475 | 0.002 | 165470 | 0.001 | 165439 | 0.002 | 165457 | 0.002 | 165524 | 0.003
165596 | 0.002 | 165587 | 0.001 | 165556 | 0.002 | 165595 | 0.002 | 165645 | 0.003
165712 | 0.002 | 165705 | 0.001 | 165673 | 0.002 | 165711 | 0.002 | 165762 | 0.003
165830 | 0.002 | 165822 | 0.001 | 165791 | 0.002 | 165832 | 0.002 | 165879 | 0.003
165948 | 0.002 | 165943 | 0.001 | 165909 | 0.002 | 165950 | 0.002 | 165997 | 0.003
166066 | 0.002 | 166061 | 0.001 | 166026 | 0.002 | 166068 | 0.002 | 166115 | 0.003
166182 | 0.002 | 166178 | 0.001 | 166144 | 0.002 | 166184 | 0.002 | 166232 | 0.003
166300 | 0.002 | 166296 | 0.001 | 166261 | 0.002 | 166302 | 0.002 | 166349 | 0.003
166418 | 0.002 | 166413 | 0.001 | 166378 | 0.002 | 166420 | 0.002 | 166467 | 0.003
166535 | 0.002 | 166531 | 0.001 | 166496 | 0.002 | 166538 | 0.002 | 166584 | 0.003
166652 | 0.002 | 166648 | 0.001 | 166614 | 0.002 | 166654 | 0.002 | 166702 | 0.003
166770 | 0.002 | 166765 | 0.001 | 166730 | 0.002 | 166772 | 0.002 | 166819 | 0.003
166887 | 0.002 | 166883 | 0.001 | 166848 | 0.002 | 166890 | 0.002 | 166936 | 0.003
167005 | 0.002 | 167001 | 0.001 | 166966 | 0.002 | 167007 | 0.002 | 167054 | 0.003
167122 | 0.002 | 167117 | 0.001 | 167084 | 0.002 | 167124 | 0.002 | 167172 | 0.003
167240 | 0.002 | 167235 | 0.001 | 167200 | 0.002 | 167242 | 0.002 | 167289 | 0.003
167357 | 0.002 | 167353 | 0.001 | 167318 | 0.002 | 167359 | 0.002 | 167406 | 0.003
167475 | 0.002 | 167471 | 0.001 | 167436 | 0.002 | 167477 | 0.002 | 167524 | 0.003
167592 | 0.002 | 167587 | 0.001 | 167553 | 0.002 | 167595 | 0.002 | 167642 | 0.003
167709 | 0.002 | 167705 | 0.001 | 167670 | 0.002 | 167712 | 0.002 | 167758 | 0.003
167827 | 0.002 | 167823 | 0.001 | 167788 | 0.002 | 167829 | 0.002 | 167876 | 0.003
167945 | 0.002 | 167940 | 0.001 | 167905 | 0.002 | 167947 | 0.002 | 167994 | 0.003
168061 | 0.002 | 168057 | 0.001 | 168023 | 0.002 | 168065 | 0.002 | 168111 | 0.003
168179 | 0.002 | 168175 | 0.001 | 168159 | 0.002 | 168181 | 0.002 | 168228 | 0.003
168297 | 0.002 | 168292 | 0.001 | 168275 | 0.002 | 168299 | 0.002 | 168346 | 0.003
168415 | 0.002 | 168410 | 0.001 | 168393 | 0.002 | 168417 | 0.002 | 168464 | 0.003
168531 | 0.002 | 168527 | 0.001 | 168511 | 0.002 | 168534 | 0.002 | 168581 | 0.003
168649 | 0.002 | 168645 | 0.001 | 168628 | 0.002 | 168651 | 0.003 | 168698 | 0.002
168767 | 0.002 | 168762 | 0.001 | 168745 | 0.002 | 168769 | 0.002 | 168816 | 0.003
168884 | 0.002 | 168880 | 0.001 | 168863 | 0.002 | 168887 | 0.002 | 168933 | 0.003
169001 | 0.002 | 168997 | 0.001 | 168980 | 0.002 | 169004 | 0.002 | 169051 | 0.003
169119 | 0.002 | 169114 | 0.001 | 169098 | 0.002 | 169121 | 0.003 | 169168 | 0.003
169236 | 0.002 | 169232 | 0.001 | 169216 | 0.002 | 169239 | 0.002 | 169286 | 0.003
169354 | 0.002 | 169350 | 0.001 | 169333 | 0.002 | 169356 | 0.002 | 169403 | 0.003
169472 | 0.002 | 169466 | 0.001 | 169450 | 0.002 | 169474 | 0.002 | 169521 | 0.003
169589 | 0.002 | 169584 | 0.001 | 169571 | 0.002 | 169591 | 0.002 | 169638 | 0.003
169706 | 0.002 | 169702 | 0.001 | 169689 | 0.002 | 169708 | 0.002 | 169755 | 0.003
169824 | 0.002 | 169820 | 0.001 | 169806 | 0.002 | 169826 | 0.002 | 169873 | 0.002
169945 | 0.002 | 169936 | 0.001 | 169923 | 0.002 | 169944 | 0.003 | 169994 | 0.003
170062 | 0.002 | 170054 | 0.001 | 170041 | 0.002 | 170061 | 0.002 | 170112 | 0.003
170179 | 0.002 | 170172 | 0.001 | 170159 | 0.002 | 170181 | 0.003 | 170228 | 0.003
170297 | 0.002 | 170293 | 0.001 | 170275 | 0.002 | 170299 | 0.002 | 170346 | 0.002
170415 | 0.002 | 170409 | 0.001 | 170393 | 0.002 | 170417 | 0.002 | 170464 | 0.002
170531 | 0.002 | 170527 | 0.001 | 170511 | 0.002 | 170534 | 0.002 | 170582 | 0.003
170649 | 0.002 | 170645 | 0.001 | 170628 | 0.002 | 170651 | 0.003 | 170698 | 0.003
170767 | 0.002 | 170762 | 0.001 | 170745 | 0.002 | 170769 | 0.002 | 170816 | 0.003
170885 | 0.002 | 170880 | 0.001 | 170863 | 0.002 | 170887 | 0.002 | 170934 | 0.003
171001 | 0.002 | 171014 | 0.001 | 170981 | 0.002 | 171003 | 0.002 | 171051 | 0.003
171119 | 0.002 | 171131 | 0.001 | 171098 | 0.002 | 171121 | 0.002 | 171168 | 0.002
171237 | 0.002 | 171249 | 0.001 | 171215 | 0.002 | 171239 | 0.002 | 171286 | 0.002
171354 | 0.002 | 171367 | 0.001 | 171333 | 0.002 | 171357 | 0.002 | 171404 | 0.003
171471 | 0.002 | 171484 | 0.001 | 171451 | 0.002 | 171473 | 0.002 | 171521 | 0.002
171589 | 0.002 | 171601 | 0.001 | 171579 | 0.002 | 171591 | 0.002 | 171638 | 0.002
171717 | 0.002 | 171719 | 0.001 | 171696 | 0.002 | 171709 | 0.002 | 171766 | 0.002
171835 | 0.002 | 171837 | 0.001 | 171813 | 0.002 | 171827 | 0.002 | 171884 | 0.002
171952 | 0.002 | 171954 | 0.001 | 171931 | 0.002 | 171943 | 0.002 | 172002 | 0.002
172069 | 0.002 | 172071 | 0.001 | 172049 | 0.002 | 172071 | 0.002 | 172119 | 0.002
172187 | 0.002 | 172199 | 0.001 | 172166 | 0.002 | 172189 | 0.002 | 172236 | 0.002
172305 | 0.002 | 172317 | 0.001 | 172283 | 0.002 | 172307 | 0.002 | 172354 | 0.003
172422 | 0.002 | 172435 | 0.001 | 172401 | 0.002 | 172424 | 0.002 | 172471 | 0.003
172539 | 0.002 | 172552 | 0.001 | 172518 | 0.002 | 172541 | 0.002 | 172589 | 0.002
172657 | 0.002 | 172669 | 0.001 | 172636 | 0.002 | 172659 | 0.002 | 172706 | 0.002
172790 | 0.002 | 172787 | 0.001 | 172753 | 0.002 | 172777 | 0.002 | 172823 | 0.002
172907 | 0.002 | 172904 | 0.001 | 172870 | 0.002 | 172894 | 0.002 | 172941 | 0.002
173025 | 0.002 | 173022 | 0.001 | 172988 | 0.002 | 173011 | 0.002 | 173059 | 0.002
173143 | 0.002 | 173139 | 0.001 | 173106 | 0.002 | 173129 | 0.002 | 173176 | 0.002
173259 | 0.002 | 173256 | 0.001 | 173223 | 0.002 | 173246 | 0.002 | 173293 | 0.002
173377 | 0.002 | 173374 | 0.001 | 173340 | 0.002 | 173364 | 0.002 | 173411 | 0.002
173495 | 0.002 | 173492 | 0.001 | 173458 | 0.002 | 173481 | 0.002 | 173529 | 0.002
173613 | 0.002 | 173609 | 0.001 | 173576 | 0.002 | 173598 | 0.002 | 173645 | 0.002
173729 | 0.002 | 173726 | 0.002 | 173692 | 0.002 | 173716 | 0.002 | 173763 | 0.002
173847 | 0.002 | 173844 | 0.002 | 173810 | 0.002 | 173834 | 0.002 | 173881 | 0.002
173965 | 0.002 | 173962 | 0.001 | 173931 | 0.002 | 173952 | 0.002 | 173999 | 0.002
174082 | 0.002 | 174078 | 0.001 | 174049 | 0.002 | 174068 | 0.002 | 174115 | 0.002
174202 | 0.002 | 174196 | 0.002 | 174166 | 0.002 | 174186 | 0.002 | 174236 | 0.002
174320 | 0.002 | 174314 | 0.002 | 174283 | 0.002 | 174304 | 0.002 | 174354 | 0.002
174438 | 0.002 | 174431 | 0.001 | 174401 | 0.002 | 174420 | 0.002 | 174472 | 0.002
174555 | 0.002 | 174548 | 0.001 | 174519 | 0.002 | 174541 | 0.002 | 174588 | 0.002
174672 | 0.002 | 174669 | 0.001 | 174636 | 0.002 | 174659 | 0.002 | 174706 | 0.002
174790 | 0.002 | 174787 | 0.002 | 174753 | 0.002 | 174777 | 0.002 | 174844 | 0.002
174908 | 0.002 | 174905 | 0.002 | 174871 | 0.002 | 174894 | 0.002 | 174962 | 0.003
175025 | 0.002 | 175022 | 0.002 | 174988 | 0.002 | 175011 | 0.002 | 175080 | 0.003
175142 | 0.002 | 175139 | 0.002 | 175106 | 0.002 | 175129 | 0.002 | 175198 | 0.002
175260 | 0.002 | 175257 | 0.002 | 175223 | 0.002 | 175246 | 0.002 | 175314 | 0.002
175377 | 0.002 | 175374 | 0.002 | 175340 | 0.002 | 175364 | 0.002 | 175432 | 0.002
175495 | 0.003 | 175492 | 0.001 | 175458 | 0.002 | 175481 | 0.002 | 175550 | 0.002
175612 | 0.002 | 175609 | 0.001 | 175576 | 0.002 | 175599 | 0.002 | 175667 | 0.002
175729 | 0.002 | 175726 | 0.001 | 175693 | 0.002 | 175716 | 0.002 | 175784 | 0.002
175847 | 0.002 | 175844 | 0.001 | 175810 | 0.002 | 175834 | 0.002 | 175902 | 0.003
175965 | 0.002 | 175962 | 0.002 | 175928 | 0.002 | 175951 | 0.002 | 176019 | 0.003
176082 | 0.003 | 176079 | 0.002 | 176046 | 0.002 | 176068 | 0.002 | 176137 | 0.003
176199 | 0.003 | 176196 | 0.002 | 176162 | 0.002 | 176186 | 0.002 | 176254 | 0.003
176317 | 0.003 | 176314 | 0.001 | 176280 | 0.002 | 176304 | 0.002 | 176372 | 0.003
176435 | 0.002 | 176432 | 0.001 | 176398 | 0.002 | 176421 | 0.002 | 176489 | 0.003
176551 | 0.003 | 176548 | 0.002 | 176516 | 0.002 | 176538 | 0.002 | 176607 | 0.003
176669 | 0.002 | 176666 | 0.002 | 176632 | 0.002 | 176656 | 0.002 | 176724 | 0.003
176787 | 0.002 | 176784 | 0.001 | 176750 | 0.002 | 176774 | 0.002 | 176841 | 0.003
176904 | 0.003 | 176901 | 0.001 | 176868 | 0.002 | 176890 | 0.002 | 176959 | 0.003
177021 | 0.003 | 177018 | 0.001 | 176985 | 0.002 | 177008 | 0.002 | 177077 | 0.003
177139 | 0.003 | 177136 | 0.001 | 177102 | 0.002 | 177126 | 0.002 | 177195 | 0.003
177257 | 0.003 | 177254 | 0.001 | 177220 | 0.002 | 177243 | 0.002 | 177311 | 0.003
177374 | 0.002 | 177371 | 0.001 | 177337 | 0.002 | 177360 | 0.002 | 177429 | 0.003
177491 | 0.002 | 177488 | 0.001 | 177455 | 0.002 | 177478 | 0.002 | 177547 | 0.003
177609 | 0.002 | 177606 | 0.002 | 177572 | 0.002 | 177596 | 0.002 | 177664 | 0.003
177726 | 0.002 | 177723 | 0.002 | 177690 | 0.002 | 177713 | 0.002 | 177781 | 0.003
177844 | 0.002 | 177841 | 0.001 | 177807 | 0.002 | 177830 | 0.002 | 177899 | 0.003
177961 | 0.002 | 177958 | 0.001 | 177925 | 0.002 | 177948 | 0.002 | 178016 | 0.003
178078 | 0.003 | 178075 | 0.002 | 178042 | 0.002 | 178065 | 0.002 | 178134 | 0.003
178196 | 0.003 | 178193 | 0.002 | 178163 | 0.002 | 178183 | 0.002 | 178251 | 0.003
178314 | 0.002 | 178311 | 0.002 | 178280 | 0.002 | 178320 | 0.002 | 178369 | 0.003
178432 | 0.002 | 178428 | 0.002 | 178398 | 0.002 | 178438 | 0.002 | 178486 | 0.003
178551 | 0.003 | 178545 | 0.002 | 178515 | 0.002 | 178554 | 0.002 | 178607 | 0.003
178669 | 0.002 | 178663 | 0.002 | 178632 | 0.002 | 178672 | 0.002 | 178724 | 0.003
178787 | 0.002 | 178781 | 0.001 | 178750 | 0.002 | 178793 | 0.002 | 178842 | 0.003
178905 | 0.003 | 178902 | 0.002 | 178868 | 0.002 | 178911 | 0.002 | 178959 | 0.003
179021 | 0.003 | 179018 | 0.001 | 178984 | 0.002 | 179027 | 0.002 | 179077 | 0.003
179139 | 0.003 | 179136 | 0.001 | 179102 | 0.002 | 179145 | 0.002 | 179194 | 0.003
179257 | 0.003 | 179254 | 0.001 | 179220 | 0.002 | 179263 | 0.002 | 179311 | 0.003
179374 | 0.003 | 179371 | 0.002 | 179338 | 0.002 | 179380 | 0.002 | 179429 | 0.003
179491 | 0.003 | 179488 | 0.001 | 179454 | 0.002 | 179497 | 0.002 | 179547 | 0.003
179609 | 0.002 | 179606 | 0.002 | 179572 | 0.002 | 179615 | 0.002 | 179664 | 0.003
179726 | 0.003 | 179724 | 0.002 | 179690 | 0.002 | 179733 | 0.002 | 179781 | 0.003
179844 | 0.003 | 179841 | 0.001 | 179807 | 0.002 | 179850 | 0.002 | 179899 | 0.003
179961 | 0.002 | 179958 | 0.001 | 179924 | 0.002 | 179968 | 0.002 | 180017 | 0.003
180079 | 0.002 | 180076 | 0.002 | 180042 | 0.002 | 180085 | 0.002 | 180133 | 0.003
180196 | 0.002 | 180193 | 0.001 | 180160 | 0.002 | 180202 | 0.002 | 180251 | 0.003
180314 | 0.003 | 180311 | 0.001 | 180277 | 0.002 | 180320 | 0.002 | 180369 | 0.003
180431 | 0.002 | 180428 | 0.001 | 180395 | 0.002 | 180438 | 0.003 | 180486 | 0.003
180548 | 0.002 | 180545 | 0.002 | 180512 | 0.002 | 180555 | 0.002 | 180603 | 0.003
180666 | 0.002 | 180663 | 0.001 | 180629 | 0.002 | 180672 | 0.002 | 180721 | 0.003
180784 | 0.002 | 180781 | 0.001 | 180747 | 0.002 | 180790 | 0.002 | 180839 | 0.01
180901 | 0.002 | 180898 | 0.001 | 180865 | 0.002 | 180908 | 0.002 | 180956 | 0.054
181018 | 0.004 | 181015 | 0.006 | 180981 | 0.002 | 181024 | 0.002 | 181073 | 0.161
181136 | 0.018 | 181133 | 0.048 | 181117 | 0.053 | 181142 | 0.003 | 181191 | 0.415
181254 | 0.091 | 181251 | 0.166 | 181234 | 0.144 | 181260 | 0.178 | 181308 | 0.404
181370 | 0.374 | 181367 | 0.307 | 181352 | 0.279 | 181377 | 0.38 | 181426 | 0.403
181488 | 0.405 | 181485 | 0.403 | 181470 | 0.43 | 181494 | 0.419 | 181543 | 0.403
181606 | 0.402 | 181603 | 0.401 | 181587 | 0.403 | 181612 | 0.402 | 181661 | 0.403
181723 | 0.402 | 181720 | 0.401 | 181704 | 0.402 | 181730 | 0.402 | 181778 | 0.403
181840 | 0.402 | 181837 | 0.401 | 181822 | 0.402 | 181847 | 0.402 | 181896 | 0.403
181958 | 0.402 | 181955 | 0.401 | 181940 | 0.402 | 181964 | 0.402 | 182013 | 0.403
182076 | 0.402 | 182073 | 0.401 | 182056 | 0.402 | 182082 | 0.402 | 182130 | 0.403
182193 | 0.402 | 182190 | 0.401 | 182174 | 0.402 | 182199 | 0.402 | 182248 | 0.403
182310 | 0.402 | 182307 | 0.401 | 182292 | 0.402 | 182317 | 0.402 | 182366 | 0.403
182428 | 0.402 | 182425 | 0.401 | 182410 | 0.402 | 182434 | 0.402 | 182483 | 0.403
182545 | 0.402 | 182542 | 0.401 | 182531 | 0.402 | 182552 | 0.402 | 182600 | 0.403
182663 | 0.402 | 182660 | 0.401 | 182647 | 0.402 | 182669 | 0.402 | 182718 | 0.403
182784 | 0.402 | 182777 | 0.401 | 182765 | 0.402 | 182787 | 0.402 | 182839 | 0.403
182901 | 0.402 | 182895 | 0.401 | 182883 | 0.402 | 182904 | 0.402 | 182957 | 0.403
183019 | 0.402 | 183012 | 0.401 | 183001 | 0.402 | 183021 | 0.402 | 183073 | 0.403
183136 | 0.402 | 183130 | 0.401 | 183117 | 0.402 | 183142 | 0.402 | 183191 | 0.403
183254 | 0.402 | 183250 | 0.401 | 183235 | 0.402 | 183260 | 0.402 | 183309 | 0.403
183371 | 0.402 | 183368 | 0.401 | 183353 | 0.402 | 183377 | 0.402 | 183427 | 0.403
183489 | 0.402 | 183485 | 0.401 | 183471 | 0.402 | 183495 | 0.402 | 183543 | 0.403
183606 | 0.402 | 183603 | 0.401 | 183587 | 0.402 | 183612 | 0.402 | 183661 | 0.403
183724 | 0.402 | 183720 | 0.401 | 183705 | 0.402 | 183730 | 0.402 | 183779 | 0.403
183841 | 0.402 | 183838 | 0.401 | 183823 | 0.402 | 183847 | 0.402 | 183897 | 0.403
183958 | 0.402 | 183955 | 0.401 | 183940 | 0.402 | 183965 | 0.402 | 184013 | 0.403
184076 | 0.402 | 184090 | 0.401 | 184057 | 0.402 | 184082 | 0.402 | 184131 | 0.403
184194 | 0.402 | 184207 | 0.401 | 184175 | 0.402 | 184200 | 0.402 | 184249 | 0.403
184311 | 0.402 | 184325 | 0.401 | 184293 | 0.402 | 184317 | 0.402 | 184367 | 0.403
184428 | 0.402 | 184442 | 0.401 | 184410 | 0.402 | 184434 | 0.402 | 184483 | 0.403
184546 | 0.402 | 184560 | 0.401 | 184527 | 0.402 | 184552 | 0.402 | 184601 | 0.403
184664 | 0.402 | 184677 | 0.401 | 184645 | 0.402 | 184670 | 0.402 | 184719 | 0.403
184781 | 0.402 | 184795 | 0.401 | 184763 | 0.402 | 184787 | 0.402 | 184836 | 0.403
184898 | 0.402 | 184912 | 0.401 | 184880 | 0.402 | 184904 | 0.402 | 184953 | 0.403
185016 | 0.402 | 185030 | 0.401 | 184997 | 0.403 | 185022 | 0.402 | 185071 | 0.403
185134 | 0.402 | 185147 | 0.401 | 185115 | 0.402 | 185140 | 0.402 | 185189 | 0.403
185251 | 0.402 | 185265 | 0.401 | 185233 | 0.402 | 185257 | 0.402 | 185306 | 0.403
185368 | 0.402 | 185382 | 0.401 | 185350 | 0.402 | 185374 | 0.402 | 185423 | 0.403
185486 | 0.402 | 185500 | 0.401 | 185467 | 0.402 | 185492 | 0.402 | 185541 | 0.403
185604 | 0.402 | 185617 | 0.401 | 185585 | 0.402 | 185610 | 0.403 | 185659 | 0.403
185721 | 0.402 | 185735 | 0.401 | 185703 | 0.403 | 185728 | 0.402 | 185776 | 0.403
185838 | 0.402 | 185852 | 0.401 | 185820 | 0.403 | 185844 | 0.402 | 185893 | 0.403
185956 | 0.402 | 185970 | 0.401 | 185937 | 0.403 | 185962 | 0.402 | 186011 | 0.403
186089 | 0.402 | 186087 | 0.401 | 186055 | 0.403 | 186080 | 0.402 | 186129 | 0.403
186207 | 0.402 | 186205 | 0.401 | 186172 | 0.403 | 186198 | 0.402 | 186246 | 0.403
186323 | 0.402 | 186322 | 0.401 | 186290 | 0.403 | 186314 | 0.402 | 186363 | 0.403
186441 | 0.402 | 186439 | 0.401 | 186407 | 0.403 | 186432 | 0.403 | 186481 | 0.403
186559 | 0.402 | 186557 | 0.401 | 186525 | 0.403 | 186550 | 0.402 | 186599 | 0.403
186677 | 0.403 | 186675 | 0.401 | 186642 | 0.402 | 186667 | 0.402 | 186716 | 0.403
186793 | 0.402 | 186792 | 0.401 | 186763 | 0.402 | 186784 | 0.402 | 186833 | 0.403
186911 | 0.402 | 186909 | 0.401 | 186881 | 0.403 | 186902 | 0.402 | 186951 | 0.403
187029 | 0.402 | 187027 | 0.401 | 186998 | 0.403 | 187020 | 0.402 | 187068 | 0.403
187150 | 0.402 | 187145 | 0.401 | 187116 | 0.403 | 187137 | 0.402 | 187189 | 0.403
187268 | 0.402 | 187262 | 0.401 | 187233 | 0.403 | 187254 | 0.402 | 187328 | 0.403
187384 | 0.402 | 187379 | 0.401 | 187351 | 0.403 | 187375 | 0.402 | 187446 | 0.403
187502 | 0.402 | 187500 | 0.401 | 187468 | 0.403 | 187493 | 0.402 | 187563 | 0.403
187620 | 0.402 | 187618 | 0.401 | 187586 | 0.403 | 187611 | 0.402 | 187680 | 0.403
187738 | 0.402 | 187736 | 0.401 | 187703 | 0.403 | 187727 | 0.402 | 187798 | 0.403
187854 | 0.402 | 187853 | 0.401 | 187821 | 0.403 | 187845 | 0.403 | 187916 | 0.403
187972 | 0.402 | 187970 | 0.401 | 187938 | 0.403 | 187963 | 0.402 | 188033 | 0.403
188090 | 0.403 | 188088 | 0.401 | 188055 | 0.403 | 188081 | 0.403 | 188150 | 0.403
188207 | 0.403 | 188206 | 0.401 | 188173 | 0.403 | 188197 | 0.403 | 188268 | 0.403
188324 | 0.402 | 188322 | 0.401 | 188291 | 0.403 | 188315 | 0.402 | 188386 | 0.403
188442 | 0.402 | 188440 | 0.401 | 188408 | 0.403 | 188433 | 0.402 | 188503 | 0.403
188560 | 0.402 | 188558 | 0.401 | 188525 | 0.403 | 188551 | 0.402 | 188620 | 0.403
188677 | 0.403 | 188676 | 0.401 | 188643 | 0.403 | 188667 | 0.402 | 188738 | 0.403
188794 | 0.403 | 188792 | 0.401 | 188761 | 0.403 | 188785 | 0.402 | 188856 | 0.403
188912 | 0.402 | 188910 | 0.401 | 188878 | 0.403 | 188903 | 0.402 | 188973 | 0.403
189030 | 0.402 | 189028 | 0.401 | 188995 | 0.403 | 189020 | 0.402 | 189090 | 0.403
189147 | 0.402 | 189146 | 0.401 | 189113 | 0.403 | 189137 | 0.402 | 189208 | 0.403
189264 | 0.402 | 189262 | 0.401 | 189231 | 0.403 | 189255 | 0.402 | 189326 | 0.403
189382 | 0.403 | 189380 | 0.401 | 189348 | 0.403 | 189373 | 0.402 | 189443 | 0.403
189500 | 0.403 | 189498 | 0.401 | 189465 | 0.403 | 189490 | 0.402 | 189560 | 0.403
189617 | 0.403 | 189616 | 0.401 | 189583 | 0.403 | 189607 | 0.402 | 189678 | 0.403
189734 | 0.403 | 189732 | 0.401 | 189701 | 0.403 | 189725 | 0.402 | 189795 | 0.403
189852 | 0.403 | 189850 | 0.401 | 189818 | 0.403 | 189843 | 0.402 | 189913 | 0.403
189970 | 0.402 | 189968 | 0.4 | 189935 | 0.403 | 189960 | 0.402 | 190030 | 0.403
190087 | 0.402 | 190086 | 0.4 | 190053 | 0.403 | 190077 | 0.402 | 190148 | 0.403
190204 | 0.403 | 190202 | 0.401 | 190171 | 0.403 | 190195 | 0.402 | 190265 | 0.403
190322 | 0.402 | 190320 | 0.401 | 190287 | 0.403 | 190313 | 0.402 | 190383 | 0.403
190439 | 0.402 | 190438 | 0.401 | 190405 | 0.403 | 190430 | 0.402 | 190500 | 0.403
190557 | 0.403 | 190556 | 0.401 | 190523 | 0.403 | 190547 | 0.402 | 190618 | 0.403
190674 | 0.402 | 190672 | 0.401 | 190641 | 0.403 | 190665 | 0.402 | 190735 | 0.403
190792 | 0.402 | 190790 | 0.401 | 190757 | 0.403 | 190782 | 0.402 | 190853 | 0.403
190909 | 0.403 | 190908 | 0.401 | 190875 | 0.403 | 190900 | 0.402 | 190970 | 0.403
191027 | 0.402 | 191026 | 0.401 | 190993 | 0.403 | 191037 | 0.402 | 191088 | 0.403
191144 | 0.403 | 191142 | 0.4 | 191114 | 0.403 | 191155 | 0.402 | 191205 | 0.403
191262 | 0.403 | 191260 | 0.4 | 191232 | 0.403 | 191272 | 0.402 | 191323 | 0.403
191383 | 0.403 | 191378 | 0.401 | 191348 | 0.403 | 191389 | 0.402 | 191443 | 0.403
191500 | 0.403 | 191495 | 0.401 | 191466 | 0.403 | 191507 | 0.402 | 191561 | 0.403
191618 | 0.403 | 191612 | 0.4 | 191584 | 0.403 | 191625 | 0.402 | 191679 | 0.403
191735 | 0.403 | 191730 | 0.401 | 191702 | 0.403 | 191746 | 0.402 | 191796 | 0.403
191853 | 0.403 | 191851 | 0.401 | 191818 | 0.403 | 191863 | 0.402 | 191913 | 0.403
191970 | 0.403 | 191969 | 0.401 | 191936 | 0.403 | 191980 | 0.402 | 192031 | 0.403
192088 | 0.403 | 192086 | 0.401 | 192054 | 0.403 | 192098 | 0.402 | 192149 | 0.403
192205 | 0.403 | 192203 | 0.401 | 192172 | 0.403 | 192216 | 0.402 | 192266 | 0.403
192323 | 0.403 | 192321 | 0.401 | 192288 | 0.403 | 192333 | 0.402 | 192384 | 0.403
192440 | 0.403 | 192439 | 0.401 | 192406 | 0.403 | 192450 | 0.402 | 192501 | 0.403
192558 | 0.403 | 192556 | 0.401 | 192524 | 0.403 | 192568 | 0.402 | 192619 | 0.403
192675 | 0.403 | 192673 | 0.401 | 192641 | 0.403 | 192686 | 0.402 | 192736 | 0.403
192792 | 0.403 | 192791 | 0.401 | 192758 | 0.403 | 192802 | 0.402 | 192854 | 0.403
192910 | 0.402 | 192909 | 0.401 | 192876 | 0.403 | 192920 | 0.402 | 192971 | 0.403
193028 | 0.402 | 193026 | 0.401 | 192994 | 0.403 | 193038 | 0.402 | 193088 | 0.403
193145 | 0.402 | 193143 | 0.401 | 193111 | 0.403 | 193156 | 0.402 | 193206 | 0.403
193262 | 0.402 | 193261 | 0.401 | 193228 | 0.403 | 193272 | 0.402 | 193324 | 0.403
193380 | 0.402 | 193379 | 0.401 | 193346 | 0.403 | 193390 | 0.402 | 193441 | 0.403
193498 | 0.403 | 193496 | 0.401 | 193464 | 0.403 | 193508 | 0.402 | 193558 | 0.403
193615 | 0.403 | 193613 | 0.401 | 193581 | 0.403 | 193626 | 0.402 | 193676 | 0.403
193732 | 0.403 | 193731 | 0.401 | 193698 | 0.403 | 193742 | 0.402 | 193794 | 0.403
193850 | 0.402 | 193848 | 0.401 | 193816 | 0.403 | 193860 | 0.402 | 193911 | 0.403
193968 | 0.402 | 193966 | 0.401 | 193934 | 0.403 | 193978 | 0.402 | 194028 | 0.403
194085 | 0.402 | 194083 | 0.401 | 194069 | 0.403 | 194096 | 0.402 | 194146 | 0.403
194202 | 0.402 | 194201 | 0.401 | 194187 | 0.403 | 194212 | 0.402 | 194264 | 0.403
194320 | 0.402 | 194318 | 0.401 | 194304 | 0.403 | 194330 | 0.402 | 194381 | 0.403
194438 | 0.402 | 194436 | 0.401 | 194421 | 0.403 | 194448 | 0.402 | 194498 | 0.403
194554 | 0.402 | 194553 | 0.401 | 194539 | 0.403 | 194565 | 0.402 | 194616 | 0.403
194672 | 0.402 | 194671 | 0.401 | 194657 | 0.403 | 194682 | 0.402 | 194734 | 0.403
194790 | 0.402 | 194788 | 0.401 | 194774 | 0.403 | 194800 | 0.402 | 194851 | 0.403
194908 | 0.402 | 194906 | 0.401 | 194891 | 0.403 | 194918 | 0.402 | 194968 | 0.403
195024 | 0.402 | 195023 | 0.401 | 195009 | 0.403 | 195035 | 0.402 | 195086 | 0.403
195142 | 0.402 | 195141 | 0.401 | 195127 | 0.403 | 195152 | 0.402 | 195204 | 0.403
195260 | 0.402 | 195258 | 0.401 | 195244 | 0.403 | 195270 | 0.402 | 195320 | 0.403
195378 | 0.402 | 195376 | 0.401 | 195361 | 0.403 | 195388 | 0.402 | 195438 | 0.403
195494 | 0.402 | 195493 | 0.401 | 195477 | 0.403 | 195505 | 0.402 | 195556 | 0.403
195612 | 0.402 | 195611 | 0.401 | 195594 | 0.403 | 195622 | 0.402 | 195674 | 0.403
195733 | 0.402 | 195728 | 0.401 | 195712 | 0.403 | 195740 | 0.402 | 195794 | 0.403
195851 | 0.402 | 195846 | 0.401 | 195829 | 0.403 | 195858 | 0.402 | 195911 | 0.403
195969 | 0.402 | 195963 | 0.401 | 195946 | 0.403 | 195979 | 0.402 | 196029 | 0.403
196085 | 0.402 | 196084 | 0.401 | 196064 | 0.403 | 196096 | 0.402 | 196147 | 0.403
196203 | 0.402 | 196202 | 0.401 | 196182 | 0.403 | 196213 | 0.402 | 196264 | 0.403
196321 | 0.402 | 196319 | 0.401 | 196299 | 0.403 | 196331 | 0.402 | 196381 | 0.403
196439 | 0.402 | 196436 | 0.401 | 196416 | 0.403 | 196449 | 0.402 | 196499 | 0.403
196555 | 0.402 | 196554 | 0.401 | 196534 | 0.403 | 196566 | 0.402 | 196617 | 0.403
196673 | 0.402 | 196671 | 0.401 | 196652 | 0.403 | 196683 | 0.402 | 196734 | 0.403
196791 | 0.402 | 196789 | 0.401 | 196769 | 0.403 | 196801 | 0.402 | 196851 | 0.403
196908 | 0.402 | 196907 | 0.401 | 196886 | 0.403 | 196918 | 0.402 | 196969 | 0.403
197025 | 0.402 | 197024 | 0.401 | 197004 | 0.403 | 197036 | 0.402 | 197087 | 0.403
197143 | 0.402 | 197141 | 0.401 | 197122 | 0.403 | 197153 | 0.402 | 197204 | 0.403
197261 | 0.402 | 197276 | 0.401 | 197239 | 0.403 | 197271 | 0.402 | 197321 | 0.403
197378 | 0.402 | 197393 | 0.401 | 197356 | 0.403 | 197388 | 0.402 | 197439 | 0.403
197495 | 0.402 | 197511 | 0.401 | 197474 | 0.403 | 197506 | 0.402 | 197557 | 0.403
197613 | 0.402 | 197628 | 0.401 | 197592 | 0.403 | 197623 | 0.402 | 197675 | 0.403
197731 | 0.402 | 197746 | 0.401 | 197709 | 0.403 | 197741 | 0.402 | 197791 | 0.403
197848 | 0.402 | 197863 | 0.401 | 197826 | 0.403 | 197858 | 0.402 | 197909 | 0.403
197965 | 0.402 | 197981 | 0.401 | 197944 | 0.403 | 197976 | 0.402 | 198027 | 0.403
198083 | 0.402 | 198098 | 0.401 | 198062 | 0.403 | 198093 | 0.402 | 198144 | 0.403
198201 | 0.402 | 198216 | 0.401 | 198178 | 0.403 | 198211 | 0.402 | 198261 | 0.403
198318 | 0.402 | 198333 | 0.401 | 198296 | 0.403 | 198328 | 0.402 | 198379 | 0.403
198435 | 0.402 | 198451 | 0.401 | 198414 | 0.403 | 198446 | 0.402 | 198497 | 0.403
198553 | 0.402 | 198568 | 0.401 | 198532 | 0.403 | 198563 | 0.402 | 198614 | 0.403
198671 | 0.402 | 198686 | 0.401 | 198648 | 0.403 | 198681 | 0.402 | 198731 | 0.403
198788 | 0.402 | 198803 | 0.401 | 198766 | 0.403 | 198798 | 0.402 | 198849 | 0.403
198905 | 0.402 | 198921 | 0.401 | 198884 | 0.403 | 198916 | 0.402 | 198967 | 0.403
199023 | 0.402 | 199038 | 0.401 | 199002 | 0.403 | 199033 | 0.402 | 199084 | 0.403
199140 | 0.402 | 199156 | 0.401 | 199118 | 0.403 | 199150 | 0.402 | 199201 | 0.403
199258 | 0.402 | 199273 | 0.401 | 199236 | 0.403 | 199268 | 0.402 | 199319 | 0.403
199391 | 0.402 | 199391 | 0.401 | 199354 | 0.403 | 199386 | 0.402 | 199437 | 0.403
199508 | 0.402 | 199508 | 0.401 | 199472 | 0.403 | 199503 | 0.402 | 199554 | 0.403
199626 | 0.402 | 199625 | 0.401 | 199588 | 0.403 | 199620 | 0.402 | 199671 | 0.403
199743 | 0.402 | 199743 | 0.401 | 199709 | 0.403 | 199738 | 0.402 | 199811 | 0.402
199861 | 0.402 | 199861 | 0.401 | 199827 | 0.403 | 199856 | 0.402 | 199927 | 0.403
199979 | 0.402 | 199979 | 0.401 | 199945 | 0.403 | 199973 | 0.402 | 200045 | 0.403
200099 | 0.402 | 200095 | 0.401 | 200062 | 0.403 | 200090 | 0.402 | 200166 | 0.403
200217 | 0.402 | 200213 | 0.401 | 200179 | 0.403 | 200208 | 0.402 | 200284 | 0.403
200335 | 0.402 | 200331 | 0.401 | 200297 | 0.403 | 200329 | 0.402 | 200401 | 0.403
200452 | 0.402 | 200452 | 0.401 | 200425 | 0.403 | 200446 | 0.402 | 200519 | 0.403
200580 | 0.402 | 200569 | 0.401 | 200543 | 0.403 | 200564 | 0.402 | 200647 | 0.403
200697 | 0.402 | 200687 | 0.401 | 200660 | 0.403 | 200682 | 0.402 | 200764 | 0.403
200815 | 0.402 | 200804 | 0.401 | 200778 | 0.403 | 200799 | 0.402 | 200882 | 0.403
200933 | 0.402 | 200922 | 0.401 | 200895 | 0.403 | 200927 | 0.402 | 200999 | 0.403
201050 | 0.402 | 201050 | 0.401 | 201013 | 0.403 | 201044 | 0.402 | 201117 | 0.402
201167 | 0.402 | 201167 | 0.401 | 201131 | 0.403 | 201162 | 0.402 | 201234 | 0.403
201285 | 0.402 | 201285 | 0.401 | 201247 | 0.403 | 201280 | 0.402 | 201352 | 0.403
201403 | 0.402 | 201402 | 0.401 | 201365 | 0.403 | 201397 | 0.402 | 201469 | 0.403
201520 | 0.402 | 201520 | 0.401 | 201483 | 0.403 | 201515 | 0.402 | 201587 | 0.403
201637 | 0.402 | 201637 | 0.401 | 201601 | 0.403 | 201632 | 0.402 | 201704 | 0.403
201755 | 0.402 | 201755 | 0.401 | 201717 | 0.403 | 201750 | 0.402 | 201822 | 0.402
201873 | 0.402 | 201872 | 0.401 | 201835 | 0.403 | 201867 | 0.402 | 201939 | 0.402
201990 | 0.402 | 201990 | 0.401 | 201953 | 0.403 | 201985 | 0.402 | 202056 | 0.402
202107 | 0.402 | 202107 | 0.401 | 202071 | 0.403 | 202102 | 0.402 | 202174 | 0.403
202225 | 0.402 | 202225 | 0.401 | 202187 | 0.403 | 202219 | 0.402 | 202292 | 0.403
202343 | 0.402 | 202342 | 0.401 | 202305 | 0.403 | 202337 | 0.402 | 202409 | 0.403
202460 | 0.402 | 202460 | 0.401 | 202423 | 0.403 | 202455 | 0.402 | 202526 | 0.402
202577 | 0.402 | 202577 | 0.401 | 202541 | 0.403 | 202572 | 0.402 | 202644 | 0.403
202695 | 0.402 | 202694 | 0.401 | 202657 | 0.403 | 202689 | 0.402 | 202762 | 0.403
202812 | 0.402 | 202812 | 0.401 | 202775 | 0.403 | 202807 | 0.402 | 202880 | 0.403
202930 | 0.402 | 202930 | 0.401 | 202893 | 0.403 | 202925 | 0.402 | 202996 | 0.403
203047 | 0.402 | 203047 | 0.401 | 203011 | 0.403 | 203042 | 0.402 | 203114 | 0.403
203165 | 0.402 | 203164 | 0.401 | 203127 | 0.403 | 203159 | 0.402 | 203232 | 0.403
203282 | 0.402 | 203282 | 0.401 | 203245 | 0.403 | 203277 | 0.402 | 203350 | 0.403
203400 | 0.402 | 203400 | 0.401 | 203363 | 0.403 | 203395 | 0.402 | 203466 | 0.403
203517 | 0.402 | 203517 | 0.401 | 203481 | 0.403 | 203512 | 0.402 | 203584 | 0.403
203635 | 0.402 | 203634 | 0.401 | 203597 | 0.403 | 203629 | 0.402 | 203702 | 0.402
203752 | 0.402 | 203752 | 0.401 | 203715 | 0.403 | 203766 | 0.402 | 203820 | 0.402
203870 | 0.402 | 203870 | 0.401 | 203833 | 0.403 | 203884 | 0.402 | 203936 | 0.403
203987 | 0.402 | 203987 | 0.401 | 203950 | 0.403 | 204002 | 0.402 | 204054 | 0.403
204105 | 0.402 | 204104 | 0.401 | 204070 | 0.403 | 204119 | 0.402 | 204172 | 0.403
204222 | 0.402 | 204222 | 0.401 | 204188 | 0.403 | 204236 | 0.402 | 204290 | 0.403
204343 | 0.402 | 204340 | 0.401 | 204306 | 0.403 | 204354 | 0.402 | 204410 | 0.403
204461 | 0.402 | 204457 | 0.401 | 204424 | 0.403 | 204472 | 0.402 | 204527 | 0.403
204578 | 0.402 | 204574 | 0.401 | 204540 | 0.403 | 204589 | 0.402 | 204645 | 0.403
204695 | 0.402 | 204692 | 0.401 | 204658 | 0.403 | 204709 | 0.402 | 204763 | 0.403
204813 | 0.402 | 204813 | 0.401 | 204776 | 0.403 | 204827 | 0.402 | 204879 | 0.403
204931 | 0.402 | 204931 | 0.401 | 204894 | 0.403 | 204945 | 0.402 | 204997 | 0.403
205048 | 0.402 | 205047 | 0.401 | 205010 | 0.403 | 205063 | 0.402 | 205115 | 0.403
205165 | 0.402 | 205165 | 0.401 | 205128 | 0.403 | 205179 | 0.402 | 205233 | 0.403
205283 | 0.402 | 205283 | 0.401 | 205246 | 0.403 | 205297 | 0.402 | 205349 | 0.403
205401 | 0.402 | 205401 | 0.401 | 205364 | 0.403 | 205415 | 0.402 | 205467 | 0.403
205518 | 0.402 | 205517 | 0.401 | 205480 | 0.403 | 205533 | 0.402 | 205585 | 0.403
205635 | 0.402 | 205635 | 0.401 | 205598 | 0.403 | 205649 | 0.402 | 205703 | 0.403
205753 | 0.402 | 205753 | 0.401 | 205716 | 0.403 | 205767 | 0.402 | 205819 | 0.403
205871 | 0.402 | 205871 | 0.401 | 205834 | 0.403 | 205885 | 0.402 | 205937 | 0.403
205988 | 0.402 | 205987 | 0.401 | 205950 | 0.403 | 206003 | 0.402 | 206055 | 0.403
206105 | 0.402 | 206105 | 0.401 | 206068 | 0.403 | 206119 | 0.402 | 206173 | 0.402
206223 | 0.402 | 206223 | 0.401 | 206186 | 0.403 | 206237 | 0.402 | 206289 | 0.403
206341 | 0.402 | 206341 | 0.401 | 206303 | 0.403 | 206355 | 0.402 | 206407 | 0.403
206457 | 0.402 | 206457 | 0.401 | 206420 | 0.403 | 206472 | 0.402 | 206525 | 0.402
206575 | 0.402 | 206575 | 0.401 | 206538 | 0.403 | 206589 | 0.402 | 206642 | 0.402
206693 | 0.402 | 206693 | 0.401 | 206656 | 0.403 | 206707 | 0.402 | 206759 | 0.402
206811 | 0.402 | 206811 | 0.401 | 206773 | 0.403 | 206825 | 0.402 | 206877 | 0.403
206927 | 0.402 | 206927 | 0.401 | 206891 | 0.403 | 206942 | 0.402 | 206995 | 0.403
207045 | 0.402 | 207045 | 0.401 | 207027 | 0.403 | 207059 | 0.402 | 207112 | 0.402
207163 | 0.402 | 207163 | 0.401 | 207143 | 0.403 | 207177 | 0.402 | 207229 | 0.402
207281 | 0.402 | 207280 | 0.401 | 207261 | 0.403 | 207295 | 0.402 | 207347 | 0.403
207397 | 0.402 | 207397 | 0.401 | 207379 | 0.403 | 207412 | 0.402 | 207465 | 0.402
207515 | 0.402 | 207515 | 0.401 | 207497 | 0.403 | 207529 | 0.402 | 207582 | 0.402
207633 | 0.402 | 207633 | 0.401 | 207613 | 0.403 | 207647 | 0.402 | 207699 | 0.402
207751 | 0.402 | 207750 | 0.401 | 207731 | 0.403 | 207765 | 0.402 | 207817 | 0.402
207867 | 0.402 | 207867 | 0.401 | 207849 | 0.403 | 207882 | 0.402 | 207935 | 0.402
207985 | 0.402 | 207985 | 0.401 | 207966 | 0.403 | 208000 | 0.402 | 208052 | 0.402
208103 | 0.402 | 208103 | 0.401 | 208083 | 0.403 | 208117 | 0.402 | 208169 | 0.402
208221 | 0.402 | 208220 | 0.401 | 208201 | 0.403 | 208234 | 0.402 | 208287 | 0.403
208337 | 0.402 | 208337 | 0.401 | 208322 | 0.403 | 208352 | 0.402 | 208405 | 0.403
208455 | 0.402 | 208455 | 0.401 | 208440 | 0.402 | 208470 | 0.402 | 208522 | 0.403
208573 | 0.402 | 208573 | 0.401 | 208557 | 0.403 | 208587 | 0.402 | 208640 | 0.402
208694 | 0.402 | 208690 | 0.401 | 208674 | 0.402 | 208704 | 0.402 | 208760 | 0.402
208811 | 0.402 | 208807 | 0.401 | 208792 | 0.403 | 208822 | 0.402 | 208878 | 0.402
208928 | 0.402 | 208925 | 0.401 | 208910 | 0.403 | 208943 | 0.402 | 208996 | 0.402
209046 | 0.402 | 209046 | 0.401 | 209027 | 0.403 | 209060 | 0.402 | 209113 | 0.402
209164 | 0.402 | 209163 | 0.401 | 209144 | 0.403 | 209178 | 0.402 | 209230 | 0.402
209281 | 0.402 | 209281 | 0.401 | 209262 | 0.402 | 209295 | 0.402 | 209348 | 0.402
209398 | 0.402 | 209398 | 0.401 | 209380 | 0.403 | 209413 | 0.402 | 209465 | 0.402
209516 | 0.402 | 209516 | 0.401 | 209497 | 0.403 | 209530 | 0.402 | 209583 | 0.402
209634 | 0.402 | 209633 | 0.401 | 209614 | 0.403 | 209648 | 0.402 | 209700 | 0.403
209751 | 0.402 | 209751 | 0.401 | 209732 | 0.403 | 209765 | 0.402 | 209818 | 0.402
209868 | 0.402 | 209868 | 0.401 | 209850 | 0.402 | 209883 | 0.402 | 209935 | 0.402
209986 | 0.402 | 209986 | 0.401 | 209967 | 0.402 | 210000 | 0.402 | 210053 | 0.402
210104 | 0.402 | 210103 | 0.401 | 210084 | 0.403 | 210118 | 0.402 | 210170 | 0.402
210221 | 0.402 | 210221 | 0.401 | 210202 | 0.403 | 210235 | 0.402 | 210288 | 0.402
210338 | 0.402 | 210355 | 0.401 | 210320 | 0.402 | 210353 | 0.402 | 210405 | 0.402
210456 | 0.402 | 210472 | 0.401 | 210437 | 0.402 | 210470 | 0.402 | 210523 | 0.402
210574 | 0.402 | 210590 | 0.401 | 210554 | 0.402 | 210587 | 0.402 | 210640 | 0.403
210691 | 0.402 | 210708 | 0.401 | 210672 | 0.402 | 210705 | 0.402 | 210758 | 0.402
210808 | 0.402 | 210825 | 0.401 | 210789 | 0.402 | 210823 | 0.402 | 210875 | 0.402
210926 | 0.402 | 210942 | 0.401 | 210907 | 0.403 | 210940 | 0.402 | 210993 | 0.403
211043 | 0.402 | 211060 | 0.401 | 211024 | 0.403 | 211057 | 0.402 | 211110 | 0.403
211161 | 0.402 | 211178 | 0.401 | 211142 | 0.402 | 211175 | 0.402 | 211228 | 0.403
211278 | 0.402 | 211295 | 0.401 | 211259 | 0.402 | 211293 | 0.402 | 211345 | 0.402
211396 | 0.402 | 211412 | 0.401 | 211377 | 0.402 | 211410 | 0.402 | 211463 | 0.402
211513 | 0.402 | 211530 | 0.401 | 211494 | 0.402 | 211527 | 0.402 | 211580 | 0.402
211631 | 0.402 | 211647 | 0.401 | 211612 | 0.403 | 211645 | 0.402 | 211697 | 0.402
211748 | 0.402 | 211765 | 0.401 | 211729 | 0.402 | 211763 | 0.402 | 211815 | 0.402
211866 | 0.402 | 211883 | 0.401 | 211847 | 0.402 | 211880 | 0.402 | 211933 | 0.402
211983 | 0.402 | 212000 | 0.401 | 211964 | 0.402 | 211997 | 0.402 | 212050 | 0.402
212101 | 0.402 | 212117 | 0.401 | 212082 | 0.402 | 212115 | 0.402 | 212167 | 0.402
212218 | 0.402 | 212235 | 0.401 | 212199 | 0.402 | 212233 | 0.402 | 212306 | 0.402
212336 | 0.402 | 212353 | 0.401 | 212317 | 0.403 | 212350 | 0.402 | 212424 | 0.403
212453 | 0.402 | 212470 | 0.401 | 212434 | 0.403 | 212467 | 0.402 | 212541 | 0.402
212571 | 0.402 | 212587 | 0.401 | 212552 | 0.402 | 212585 | 0.402 | 212659 | 0.402
212704 | 0.402 | 212705 | 0.401 | 212673 | 0.402 | 212703 | 0.402 | 212776 | 0.402
212822 | 0.402 | 212823 | 0.401 | 212790 | 0.402 | 212819 | 0.402 | 212894 | 0.403
212942 | 0.402 | 212940 | 0.401 | 212908 | 0.402 | 212937 | 0.402 | 213015 | 0.403
213060 | 0.402 | 213057 | 0.401 | 213025 | 0.402 | 213055 | 0.402 | 213132 | 0.403
213177 | 0.402 | 213175 | 0.401 | 213142 | 0.402 | 213173 | 0.402 | 213250 | 0.403
213295 | 0.402 | 213293 | 0.401 | 213260 | 0.402 | 213294 | 0.402 | 213367 | 0.403
213412 | 0.402 | 213413 | 0.401 | 213378 | 0.402 | 213410 | 0.402 | 213485 | 0.403
213530 | 0.402 | 213531 | 0.401 | 213495 | 0.402 | 213528 | 0.402 | 213602 | 0.403
213647 | 0.402 | 213648 | 0.401 | 213612 | 0.402 | 213646 | 0.402 | 213720 | 0.403
213765 | 0.401 | 213766 | 0.401 | 213730 | 0.402 | 213764 | 0.402 | 213837 | 0.403
213882 | 0.402 | 213883 | 0.401 | 213848 | 0.402 | 213880 | 0.402 | 213955 | 0.403
213999 | 0.402 | 214000 | 0.401 | 213965 | 0.402 | 213998 | 0.402 | 214072 | 0.403
214117 | 0.402 | 214118 | 0.401 | 214082 | 0.402 | 214116 | 0.402 | 214190 | 0.403
214235 | 0.402 | 214236 | 0.401 | 214200 | 0.402 | 214234 | 0.402 | 214307 | 0.403
214352 | 0.402 | 214353 | 0.401 | 214318 | 0.402 | 214350 | 0.402 | 214424 | 0.403
214469 | 0.402 | 214470 | 0.401 | 214435 | 0.402 | 214468 | 0.402 | 214542 | 0.403
214587 | 0.402 | 214588 | 0.401 | 214552 | 0.402 | 214586 | 0.402 | 214660 | 0.403
214705 | 0.402 | 214706 | 0.401 | 214670 | 0.402 | 214704 | 0.402 | 214777 | 0.403
214822 | 0.402 | 214823 | 0.401 | 214788 | 0.402 | 214820 | 0.402 | 214894 | 0.402
214939 | 0.402 | 214940 | 0.401 | 214905 | 0.402 | 214938 | 0.402 | 215012 | 0.403
215057 | 0.402 | 215058 | 0.401 | 215022 | 0.402 | 215056 | 0.402 | 215130 | 0.403
215175 | 0.402 | 215176 | 0.401 | 215140 | 0.402 | 215173 | 0.402 | 215247 | 0.403
215292 | 0.402 | 215293 | 0.401 | 215258 | 0.402 | 215290 | 0.402 | 215364 | 0.403
215409 | 0.402 | 215410 | 0.401 | 215374 | 0.402 | 215408 | 0.402 | 215482 | 0.402
215527 | 0.402 | 215528 | 0.401 | 215492 | 0.402 | 215526 | 0.402 | 215600 | 0.402
215645 | 0.402 | 215646 | 0.401 | 215610 | 0.402 | 215643 | 0.402 | 215717 | 0.403
215762 | 0.402 | 215762 | 0.401 | 215728 | 0.402 | 215760 | 0.402 | 215834 | 0.402
215879 | 0.402 | 215880 | 0.401 | 215844 | 0.402 | 215878 | 0.402 | 215952 | 0.402
215997 | 0.402 | 215998 | 0.401 | 215962 | 0.402 | 215996 | 0.402 | 216070 | 0.402
216115 | 0.402 | 216116 | 0.401 | 216080 | 0.402 | 216113 | 0.402 | 216186 | 0.403
216231 | 0.402 | 216232 | 0.401 | 216198 | 0.402 | 216230 | 0.402 | 216304 | 0.403
216349 | 0.402 | 216350 | 0.401 | 216314 | 0.402 | 216348 | 0.402 | 216422 | 0.403
216467 | 0.402 | 216468 | 0.401 | 216432 | 0.402 | 216485 | 0.402 | 216540 | 0.402
216585 | 0.402 | 216586 | 0.401 | 216550 | 0.402 | 216602 | 0.402 | 216656 | 0.402
216701 | 0.402 | 216702 | 0.401 | 216668 | 0.402 | 216720 | 0.402 | 216774 | 0.402
216819 | 0.402 | 216820 | 0.401 | 216784 | 0.402 | 216838 | 0.402 | 216892 | 0.402
216937 | 0.402 | 216938 | 0.401 | 216905 | 0.402 | 216955 | 0.402 | 217010 | 0.403
217055 | 0.402 | 217056 | 0.401 | 217023 | 0.402 | 217072 | 0.402 | 217126 | 0.402
217171 | 0.402 | 217172 | 0.401 | 217141 | 0.402 | 217190 | 0.402 | 217244 | 0.402
217292 | 0.402 | 217290 | 0.401 | 217259 | 0.402 | 217308 | 0.402 | 217365 | 0.402
217410 | 0.402 | 217408 | 0.401 | 217375 | 0.402 | 217425 | 0.402 | 217483 | 0.402
217528 | 0.402 | 217526 | 0.401 | 217493 | 0.402 | 217546 | 0.402 | 217601 | 0.402
217646 | 0.402 | 217647 | 0.401 | 217611 | 0.402 | 217663 | 0.402 | 217717 | 0.402
217762 | 0.402 | 217763 | 0.401 | 217729 | 0.402 | 217781 | 0.402 | 217835 | 0.402
217880 | 0.402 | 217881 | 0.401 | 217845 | 0.402 | 217899 | 0.402 | 217953 | 0.402
217998 | 0.402 | 217999 | 0.401 | 217963 | 0.402 | 218016 | 0.402 | 218071 | 0.402
218116 | 0.402 | 218117 | 0.401 | 218081 | 0.402 | 218133 | 0.402 | 218188 | 0.402
218232 | 0.402 | 218233 | 0.401 | 218199 | 0.402 | 218251 | 0.402 | 218305 | 0.402
218350 | 0.401 | 218351 | 0.401 | 218315 | 0.402 | 218369 | 0.402 | 218423 | 0.402
218468 | 0.401 | 218469 | 0.401 | 218433 | 0.402 | 218486 | 0.402 | 218541 | 0.402
218586 | 0.402 | 218587 | 0.401 | 218551 | 0.402 | 218603 | 0.402 | 218658 | 0.402
218703 | 0.402 | 218703 | 0.401 | 218669 | 0.402 | 218721 | 0.402 | 218775 | 0.402
218820 | 0.402 | 218821 | 0.401 | 218786 | 0.402 | 218839 | 0.402 | 218893 | 0.402
218938 | 0.402 | 218939 | 0.401 | 218903 | 0.402 | 218956 | 0.402 | 219011 | 0.402
219056 | 0.402 | 219057 | 0.401 | 219021 | 0.402 | 219074 | 0.402 | 219128 | 0.402
219173 | 0.401 | 219173 | 0.401 | 219139 | 0.402 | 219191 | 0.402 | 219245 | 0.402
219290 | 0.402 | 219291 | 0.401 | 219256 | 0.402 | 219309 | 0.402 | 219363 | 0.402
219408 | 0.402 | 219409 | 0.401 | 219373 | 0.402 | 219426 | 0.402 | 219481 | 0.402
219526 | 0.401 | 219527 | 0.401 | 219491 | 0.402 | 219544 | 0.402 | 219598 | 0.402
219643 | 0.402 | 219643 | 0.401 | 219609 | 0.402 | 219661 | 0.402 | 219715 | 0.402
219760 | 0.402 | 219761 | 0.401 | 219726 | 0.402 | 219779 | 0.402 | 219833 | 0.402
219878 | 0.402 | 219879 | 0.401 | 219843 | 0.402 | 219896 | 0.402 | 219951 | 0.402
219996 | 0.402 | 219997 | 0.401 | 219979 | 0.402 | 220014 | 0.402 | 220068 | 0.402
220113 | 0.402 | 220113 | 0.401 | 220097 | 0.402 | 220131 | 0.402 | 220185 | 0.402
220230 | 0.402 | 220231 | 0.401 | 220214 | 0.402 | 220249 | 0.402 | 220303 | 0.402
220348 | 0.402 | 220349 | 0.401 | 220332 | 0.402 | 220366 | 0.402 | 220421 | 0.402
220466 | 0.401 | 220467 | 0.401 | 220449 | 0.402 | 220484 | 0.402 | 220538 | 0.402
220583 | 0.401 | 220583 | 0.401 | 220567 | 0.402 | 220601 | 0.402 | 220656 | 0.402
220700 | 0.402 | 220701 | 0.401 | 220684 | 0.402 | 220719 | 0.402 | 220773 | 0.402
220818 | 0.402 | 220819 | 0.401 | 220802 | 0.402 | 220836 | 0.402 | 220891 | 0.402
220936 | 0.402 | 220937 | 0.401 | 220919 | 0.402 | 220954 | 0.402 | 221008 | 0.402
221053 | 0.402 | 221054 | 0.401 | 221037 | 0.402 | 221071 | 0.402 | 221126 | 0.402
221171 | 0.402 | 221171 | 0.401 | 221154 | 0.402 | 221189 | 0.402 | 221243 | 0.402
221288 | 0.401 | 221289 | 0.401 | 221275 | 0.402 | 221306 | 0.402 | 221361 | 0.402
221406 | 0.402 | 221407 | 0.401 | 221392 | 0.402 | 221424 | 0.402 | 221478 | 0.402
221527 | 0.402 | 221524 | 0.401 | 221510 | 0.402 | 221542 | 0.402 | 221599 | 0.402
221644 | 0.402 | 221641 | 0.401 | 221628 | 0.402 | 221659 | 0.402 | 221717 | 0.402
221762 | 0.402 | 221759 | 0.401 | 221745 | 0.402 | 221776 | 0.402 | 221834 | 0.402
221879 | 0.402 | 221877 | 0.401 | 221862 | 0.402 | 221897 | 0.402 | 221952 | 0.402
221997 | 0.402 | 221998 | 0.401 | 221980 | 0.402 | 222015 | 0.402 | 222069 | 0.402
222114 | 0.401 | 222115 | 0.401 | 222098 | 0.402 | 222133 | 0.402 | 222187 | 0.402
222232 | 0.401 | 222232 | 0.401 | 222216 | 0.402 | 222249 | 0.402 | 222305 | 0.402
222349 | 0.401 | 222350 | 0.401 | 222332 | 0.402 | 222367 | 0.402 | 222422 | 0.402
222467 | 0.402 | 222468 | 0.401 | 222450 | 0.402 | 222485 | 0.402 | 222539 | 0.402
222584 | 0.401 | 222585 | 0.401 | 222568 | 0.402 | 222603 | 0.402 | 222657 | 0.402
222702 | 0.401 | 222702 | 0.401 | 222686 | 0.402 | 222720 | 0.402 | 222775 | 0.402
222819 | 0.402 | 222820 | 0.401 | 222802 | 0.402 | 222837 | 0.402 | 222892 | 0.402
222937 | 0.402 | 222938 | 0.401 | 222920 | 0.402 | 222955 | 0.402 | 223009 | 0.402
223054 | 0.402 | 223055 | 0.401 | 223038 | 0.402 | 223073 | 0.402 | 223127 | 0.402
223172 | 0.401 | 223172 | 0.401 | 223156 | 0.402 | 223190 | 0.402 | 223245 | 0.402
223290 | 0.402 | 223290 | 0.401 | 223272 | 0.402 | 223307 | 0.402 | 223362 | 0.402
223407 | 0.402 | 223408 | 0.401 | 223390 | 0.402 | 223425 | 0.402 | 223479 | 0.402
223524 | 0.401 | 223542 | 0.401 | 223508 | 0.402 | 223543 | 0.402 | 223597 | 0.402
223642 | 0.401 | 223660 | 0.401 | 223626 | 0.402 | 223660 | 0.402 | 223715 | 0.402
223760 | 0.402 | 223777 | 0.401 | 223742 | 0.402 | 223777 | 0.402 | 223832 | 0.402
223877 | 0.402 | 223895 | 0.401 | 223860 | 0.402 | 223895 | 0.402 | 223949 | 0.402
223994 | 0.402 | 224012 | 0.401 | 223978 | 0.402 | 224013 | 0.402 | 224067 | 0.402
224112 | 0.402 | 224130 | 0.401 | 224096 | 0.402 | 224131 | 0.402 | 224185 | 0.402
224230 | 0.401 | 224247 | 0.401 | 224214 | 0.402 | 224247 | 0.402 | 224302 | 0.402
224347 | 0.402 | 224365 | 0.401 | 224330 | 0.402 | 224365 | 0.402 | 224419 | 0.402
224464 | 0.402 | 224482 | 0.401 | 224448 | 0.402 | 224483 | 0.402 | 224537 | 0.402
224582 | 0.402 | 224600 | 0.401 | 224566 | 0.402 | 224601 | 0.402 | 224655 | 0.402
224700 | 0.402 | 224717 | 0.401 | 224684 | 0.402 | 224717 | 0.402 | 224793 | 0.402
224817 | 0.402 | 224835 | 0.401 | 224800 | 0.402 | 224835 | 0.402 | 224911 | 0.402
224934 | 0.402 | 224952 | 0.401 | 224918 | 0.402 | 224953 | 0.402 | 225029 | 0.402
225052 | 0.402 | 225070 | 0.401 | 225036 | 0.402 | 225071 | 0.402 | 225146 | 0.402
225170 | 0.402 | 225187 | 0.401 | 225154 | 0.402 | 225187 | 0.402 | 225263 | 0.402
225287 | 0.402 | 225305 | 0.401 | 225270 | 0.402 | 225305 | 0.402 | 225381 | 0.402
225404 | 0.402 | 225422 | 0.401 | 225388 | 0.402 | 225423 | 0.402 | 225499 | 0.402
225522 | 0.402 | 225540 | 0.401 | 225506 | 0.402 | 225541 | 0.402 | 225616 | 0.402
225640 | 0.402 | 225657 | 0.401 | 225620 | 0.402 | 225657 | 0.402 | 225733 | 0.402
225758 | 0.402 | 225775 | 0.401 | 225738 | 0.402 | 225775 | 0.402 | 225851 | 0.402
225894 | 0.402 | 225892 | 0.401 | 225856 | 0.402 | 225893 | 0.402 | 225972 | 0.402
226012 | 0.402 | 226010 | 0.401 | 225974 | 0.402 | 226011 | 0.402 | 226089 | 0.402
226128 | 0.402 | 226127 | 0.401 | 226091 | 0.402 | 226131 | 0.402 | 226207 | 0.402
226246 | 0.402 | 226248 | 0.401 | 226208 | 0.402 | 226248 | 0.402 | 226324 | 0.402
226364 | 0.402 | 226366 | 0.401 | 226326 | 0.402 | 226366 | 0.402 | 226442 | 0.402
226482 | 0.402 | 226484 | 0.401 | 226444 | 0.402 | 226484 | 0.402 | 226560 | 0.402
226598 | 0.402 | 226600 | 0.401 | 226561 | 0.402 | 226601 | 0.402 | 226677 | 0.402
226716 | 0.402 | 226718 | 0.401 | 226678 | 0.402 | 226719 | 0.402 | 226795 | 0.402
226834 | 0.401 | 226836 | 0.401 | 226796 | 0.402 | 226836 | 0.402 | 226912 | 0.402
226952 | 0.402 | 226954 | 0.401 | 226914 | 0.402 | 226954 | 0.402 | 227030 | 0.402
227068 | 0.402 | 227070 | 0.401 | 227032 | 0.402 | 227072 | 0.402 | 227147 | 0.402
227186 | 0.402 | 227188 | 0.401 | 227148 | 0.402 | 227189 | 0.402 | 227265 | 0.402
227304 | 0.401 | 227306 | 0.401 | 227266 | 0.402 | 227306 | 0.402 | 227382 | 0.402
227422 | 0.402 | 227424 | 0.401 | 227384 | 0.402 | 227424 | 0.402 | 227500 | 0.402
227538 | 0.401 | 227540 | 0.401 | 227502 | 0.402 | 227542 | 0.402 | 227617 | 0.402
227656 | 0.401 | 227658 | 0.401 | 227618 | 0.402 | 227659 | 0.402 | 227735 | 0.402
227774 | 0.401 | 227776 | 0.401 | 227736 | 0.402 | 227776 | 0.402 | 227852 | 0.402
227892 | 0.401 | 227894 | 0.401 | 227854 | 0.402 | 227894 | 0.402 | 227970 | 0.402
228009 | 0.401 | 228011 | 0.401 | 227972 | 0.402 | 228012 | 0.402 | 228087 | 0.402
228126 | 0.402 | 228128 | 0.401 | 228088 | 0.402 | 228129 | 0.402 | 228205 | 0.402
228244 | 0.401 | 228246 | 0.401 | 228206 | 0.402 | 228246 | 0.402 | 228322 | 0.402
228362 | 0.402 | 228364 | 0.401 | 228324 | 0.402 | 228364 | 0.402 | 228440 | 0.402
228480 | 0.401 | 228482 | 0.401 | 228442 | 0.402 | 228482 | 0.402 | 228557 | 0.402
228596 | 0.401 | 228598 | 0.401 | 228558 | 0.402 | 228599 | 0.402 | 228675 | 0.402
228714 | 0.401 | 228716 | 0.401 | 228676 | 0.402 | 228716 | 0.402 | 228793 | 0.402
228832 | 0.402 | 228834 | 0.401 | 228794 | 0.402 | 228834 | 0.402 | 228910 | 0.402
228950 | 0.401 | 228952 | 0.401 | 228912 | 0.402 | 228952 | 0.402 | 229027 | 0.402
229067 | 0.401 | 229068 | 0.401 | 229029 | 0.402 | 229069 | 0.402 | 229145 | 0.402
229184 | 0.402 | 229186 | 0.401 | 229146 | 0.402 | 229207 | 0.402 | 229263 | 0.402
229302 | 0.401 | 229304 | 0.401 | 229264 | 0.402 | 229324 | 0.402 | 229381 | 0.402
229430 | 0.401 | 229422 | 0.401 | 229392 | 0.402 | 229441 | 0.402 | 229509 | 0.402
229548 | 0.401 | 229539 | 0.401 | 229510 | 0.402 | 229559 | 0.402 | 229625 | 0.402
229665 | 0.401 | 229656 | 0.401 | 229627 | 0.402 | 229677 | 0.402 | 229743 | 0.402
229783 | 0.401 | 229774 | 0.401 | 229745 | 0.402 | 229805 | 0.402 | 229861 | 0.402
229900 | 0.401 | 229892 | 0.401 | 229866 | 0.402 | 229922 | 0.402 | 229979 | 0.402
230018 | 0.402 | 230020 | 0.401 | 229984 | 0.402 | 230040 | 0.402 | 230095 | 0.402
230135 | 0.402 | 230137 | 0.401 | 230100 | 0.402 | 230157 | 0.402 | 230213 | 0.402
230256 | 0.401 | 230255 | 0.401 | 230218 | 0.402 | 230275 | 0.402 | 230334 | 0.402
230374 | 0.401 | 230372 | 0.401 | 230336 | 0.402 | 230392 | 0.402 | 230452 | 0.402
230492 | 0.401 | 230490 | 0.401 | 230454 | 0.403 | 230513 | 0.402 | 230569 | 0.402
230608 | 0.402 | 230610 | 0.401 | 230571 | 0.402 | 230631 | 0.402 | 230687 | 0.402
230726 | 0.401 | 230728 | 0.401 | 230688 | 0.402 | 230748 | 0.402 | 230804 | 0.402
230844 | 0.401 | 230846 | 0.401 | 230806 | 0.402 | 230865 | 0.402 | 230922 | 0.402
230962 | 0.401 | 230963 | 0.401 | 230924 | 0.402 | 230983 | 0.403 | 231040 | 0.402
231078 | 0.401 | 231080 | 0.401 | 231041 | 0.402 | 231101 | 0.403 | 231157 | 0.402
231196 | 0.401 | 231198 | 0.401 | 231158 | 0.402 | 231218 | 0.402 | 231274 | 0.402
231314 | 0.402 | 231316 | 0.401 | 231276 | 0.402 | 231335 | 0.402 | 231392 | 0.402
231432 | 0.402 | 231433 | 0.401 | 231394 | 0.402 | 231453 | 0.402 | 231510 | 0.402
231548 | 0.402 | 231550 | 0.401 | 231511 | 0.402 | 231571 | 0.402 | 231627 | 0.402
231666 | 0.402 | 231668 | 0.401 | 231628 | 0.402 | 231688 | 0.403 | 231744 | 0.402
231784 | 0.401 | 231786 | 0.401 | 231746 | 0.402 | 231805 | 0.403 | 231862 | 0.402
231902 | 0.401 | 231904 | 0.401 | 231864 | 0.402 | 231923 | 0.403 | 231980 | 0.402
232018 | 0.401 | 232020 | 0.401 | 231981 | 0.402 | 232041 | 0.403 | 232097 | 0.402
232136 | 0.401 | 232138 | 0.401 | 232098 | 0.402 | 232159 | 0.403 | 232214 | 0.402
232254 | 0.401 | 232256 | 0.401 | 232216 | 0.402 | 232275 | 0.403 | 232332 | 0.402
232372 | 0.401 | 232374 | 0.401 | 232334 | 0.402 | 232393 | 0.402 | 232450 | 0.402
232488 | 0.402 | 232491 | 0.401 | 232452 | 0.402 | 232511 | 0.403 | 232567 | 0.402
232606 | 0.402 | 232608 | 0.401 | 232568 | 0.402 | 232629 | 0.403 | 232685 | 0.402
232724 | 0.402 | 232726 | 0.401 | 232686 | 0.402 | 232745 | 0.403 | 232802 | 0.402
232842 | 0.402 | 232844 | 0.401 | 232804 | 0.402 | 232863 | 0.402 | 232920 | 0.402
232960 | 0.401 | 232961 | 0.401 | 232922 | 0.402 | 232981 | 0.403 | 233037 | 0.402
233076 | 0.401 | 233078 | 0.401 | 233057 | 0.402 | 233099 | 0.402 | 233155 | 0.402
233194 | 0.402 | 233196 | 0.401 | 233175 | 0.402 | 233216 | 0.403 | 233272 | 0.402
233312 | 0.402 | 233314 | 0.401 | 233292 | 0.402 | 233333 | 0.403 | 233390 | 0.402
233430 | 0.401 | 233431 | 0.401 | 233409 | 0.402 | 233451 | 0.403 | 233507 | 0.402
233546 | 0.401 | 233548 | 0.401 | 233527 | 0.402 | 233569 | 0.403 | 233625 | 0.402
233664 | 0.401 | 233666 | 0.401 | 233645 | 0.402 | 233686 | 0.403 | 233742 | 0.402
233782 | 0.402 | 233784 | 0.401 | 233762 | 0.402 | 233803 | 0.403 | 233860 | 0.402
233900 | 0.402 | 233901 | 0.401 | 233879 | 0.402 | 233921 | 0.402 | 233977 | 0.402
234016 | 0.401 | 234018 | 0.401 | 233997 | 0.402 | 234039 | 0.403 | 234095 | 0.402
234134 | 0.401 | 234136 | 0.401 | 234115 | 0.402 | 234156 | 0.402 | 234212 | 0.402
234252 | 0.401 | 234254 | 0.401 | 234235 | 0.402 | 234273 | 0.403 | 234330 | 0.402
234370 | 0.402 | 234372 | 0.401 | 234353 | 0.402 | 234391 | 0.403 | 234447 | 0.402
234490 | 0.402 | 234488 | 0.401 | 234470 | 0.402 | 234509 | 0.403 | 234568 | 0.402
234607 | 0.401 | 234606 | 0.401 | 234588 | 0.402 | 234626 | 0.403 | 234686 | 0.402
234725 | 0.401 | 234724 | 0.401 | 234705 | 0.402 | 234743 | 0.403 | 234804 | 0.402
234843 | 0.401 | 234842 | 0.401 | 234823 | 0.402 | 234864 | 0.403 | 234920 | 0.402
234960 | 0.401 | 234962 | 0.401 | 234941 | 0.402 | 234982 | 0.403 | 235038 | 0.402
235077 | 0.401 | 235079 | 0.401 | 235058 | 0.402 | 235100 | 0.403 | 235156 | 0.402
235195 | 0.401 | 235197 | 0.401 | 235175 | 0.402 | 235217 | 0.402 | 235274 | 0.402
235313 | 0.401 | 235315 | 0.401 | 235293 | 0.402 | 235334 | 0.402 | 235390 | 0.402
235430 | 0.401 | 235432 | 0.401 | 235411 | 0.402 | 235452 | 0.402 | 235508 | 0.402
235548 | 0.401 | 235549 | 0.401 | 235528 | 0.402 | 235570 | 0.403 | 235626 | 0.402
235665 | 0.401 | 235667 | 0.401 | 235645 | 0.402 | 235687 | 0.402 | 235744 | 0.402
235783 | 0.401 | 235785 | 0.401 | 235763 | 0.402 | 235804 | 0.403 | 235860 | 0.402
235900 | 0.401 | 235902 | 0.401 | 235881 | 0.402 | 235922 | 0.402 | 235978 | 0.402
236018 | 0.401 | 236019 | 0.401 | 235998 | 0.402 | 236040 | 0.402 | 236096 | 0.402
236135 | 0.401 | 236137 | 0.401 | 236115 | 0.402 | 236157 | 0.402 | 236214 | 0.402
236253 | 0.401 | 236255 | 0.401 | 236233 | 0.402 | 236274 | 0.402 | 236331 | 0.402
236370 | 0.401 | 236372 | 0.401 | 236351 | 0.402 | 236392 | 0.403 | 236448 | 0.402
236488 | 0.401 | 236490 | 0.401 | 236468 | 0.402 | 236510 | 0.402 | 236566 | 0.402
236605 | 0.401 | 236624 | 0.401 | 236585 | 0.402 | 236627 | 0.402 | 236684 | 0.402
236723 | 0.401 | 236742 | 0.401 | 236703 | 0.402 | 236744 | 0.403 | 236801 | 0.402
236840 | 0.401 | 236859 | 0.401 | 236821 | 0.402 | 236862 | 0.403 | 236918 | 0.402
236958 | 0.401 | 236977 | 0.401 | 236938 | 0.402 | 236980 | 0.402 | 237036 | 0.402
237075 | 0.401 | 237094 | 0.401 | 237056 | 0.402 | 237097 | 0.402 | 237154 | 0.402
237193 | 0.401 | 237212 | 0.401 | 237173 | 0.402 | 237214 | 0.402 | 237292 | 0.402
237310 | 0.401 | 237329 | 0.401 | 237291 | 0.402 | 237332 | 0.402 | 237410 | 0.402
237428 | 0.401 | 237447 | 0.401 | 237409 | 0.402 | 237450 | 0.402 | 237528 | 0.402
237545 | 0.401 | 237564 | 0.401 | 237526 | 0.402 | 237568 | 0.402 | 237644 | 0.402
237663 | 0.401 | 237682 | 0.401 | 237643 | 0.402 | 237684 | 0.402 | 237762 | 0.402
237780 | 0.401 | 237799 | 0.401 | 237761 | 0.402 | 237802 | 0.402 | 237880 | 0.402
237898 | 0.401 | 237917 | 0.401 | 237879 | 0.402 | 237920 | 0.402 | 237998 | 0.402
238015 | 0.401 | 238034 | 0.401 | 237996 | 0.402 | 238038 | 0.402 | 238114 | 0.402
238133 | 0.401 | 238152 | 0.401 | 238113 | 0.402 | 238155 | 0.402 | 238232 | 0.402
238251 | 0.401 | 238270 | 0.401 | 238231 | 0.402 | 238272 | 0.402 | 238350 | 0.402
238368 | 0.401 | 238387 | 0.401 | 238349 | 0.402 | 238390 | 0.402 | 238468 | 0.402
238485 | 0.401 | 238504 | 0.401 | 238470 | 0.402 | 238508 | 0.402 | 238585 | 0.402
238603 | 0.401 | 238622 | 0.401 | 238586 | 0.402 | 238625 | 0.402 | 238702 | 0.402
238721 | 0.401 | 238740 | 0.401 | 238704 | 0.402 | 238742 | 0.402 | 238820 | 0.402
238841 | 0.401 | 238857 | 0.401 | 238822 | 0.402 | 238860 | 0.402 | 238941 | 0.402
238959 | 0.401 | 238974 | 0.401 | 238940 | 0.402 | 238978 | 0.402 | 239058 | 0.402
239077 | 0.401 | 239092 | 0.401 | 239057 | 0.402 | 239098 | 0.402 | 239176 | 0.402
239210 | 0.401 | 239213 | 0.401 | 239174 | 0.402 | 239216 | 0.402 | 239293 | 0.402
239327 | 0.401 | 239331 | 0.401 | 239292 | 0.402 | 239333 | 0.402 | 239411 | 0.402
239444 | 0.401 | 239448 | 0.401 | 239410 | 0.402 | 239451 | 0.402 | 239528 | 0.402
239562 | 0.401 | 239565 | 0.401 | 239527 | 0.402 | 239568 | 0.402 | 239646 | 0.402
239680 | 0.401 | 239683 | 0.401 | 239644 | 0.402 | 239686 | 0.402 | 239763 | 0.402
239797 | 0.401 | 239801 | 0.401 | 239762 | 0.402 | 239803 | 0.402 | 239881 | 0.402
239915 | 0.401 | 239919 | 0.401 | 239880 | 0.402 | 239921 | 0.402 | 239998 | 0.402
240032 | 0.401 | 240035 | 0.401 | 239997 | 0.402 | 240038 | 0.402 | 240116 | 0.402
240150 | 0.401 | 240153 | 0.401 | 240114 | 0.402 | 240156 | 0.402 | 240233 | 0.402
240267 | 0.401 | 240271 | 0.401 | 240232 | 0.402 | 240274 | 0.402 | 240351 | 0.402
240385 | 0.401 | 240389 | 0.401 | 240350 | 0.402 | 240391 | 0.402 | 240468 | 0.402
240502 | 0.401 | 240505 | 0.401 | 240467 | 0.402 | 240508 | 0.402 | 240586 | 0.402
240620 | 0.401 | 240623 | 0.401 | 240584 | 0.402 | 240626 | 0.402 | 240704 | 0.402
240737 | 0.401 | 240741 | 0.401 | 240702 | 0.402 | 240744 | 0.402 | 240821 | 0.402
240855 | 0.402 | 240859 | 0.401 | 240820 | 0.402 | 240861 | 0.402 | 240939 | 0.402
240972 | 0.146 | 240975 | 0.401 | 240937 | 0.402 | 240978 | 0.402 | 241056 | 0.406
241090 | 0.012 | 241093 | 0.417 | 241055 | 0.407 | 241096 | 0.402 | 241174 | 0.213
241208 | 0.001 | 241211 | 0.386 | 241172 | 0.389 | 241214 | 0.402 | 241291 | 0.03
241325 | 0.001 | 241329 | 0.181 | 241290 | 0.274 | 241331 | 0.41 | 241409 | 0.002
241442 | 0.002 | 241445 | 0.016 | 241408 | 0.067 | 241448 | 0.239 | 241526 | 0.002
241560 | 0.002 | 241563 | 0.001 | 241525 | 0.002 | 241566 | 0.028 | 241644 | 0.002
241678 | 0.002 | 241681 | 0.001 | 241642 | 0.002 | 241684 | 0.002 | 241761 | 0.002
241795 | 0.002 | 241799 | 0.001 | 241760 | 0.002 | 241801 | 0.002 | 241879 | 0.002
241912 | 0.002 | 241915 | 0.001 | 241878 | 0.002 | 241938 | 0.002 | 241996 | 0.002
242030 | 0.001 | 242033 | 0.001 | 241995 | 0.002 | 242056 | 0.002 | 242114 | 0.002
242148 | 0.001 | 242151 | 0.001 | 242112 | 0.002 | 242173 | 0.002 | 242231 | 0.002
242265 | 0.002 | 242269 | 0.001 | 242230 | 0.002 | 242291 | 0.002 | 242349 | 0.002
242382 | 0.002 | 242385 | 0.001 | 242348 | 0.002 | 242408 | 0.002 | 242466 | 0.002
242500 | 0.002 | 242503 | 0.001 | 242465 | 0.002 | 242526 | 0.002 | 242584 | 0.002
242618 | 0.002 | 242621 | 0.001 | 242582 | 0.002 | 242643 | 0.002 | 242701 | 0.002
242735 | 0.002 | 242739 | 0.001 | 242700 | 0.002 | 242761 | 0.002 | 242819 | 0.002
242852 | 0.002 | 242855 | 0.001 | 242821 | 0.002 | 242878 | 0.002 | 242936 | 0.002
242970 | 0.002 | 242973 | 0.001 | 242938 | 0.002 | 242996 | 0.002 | 243054 | 0.002
243091 | 0.002 | 243091 | 0.001 | 243056 | 0.002 | 243113 | 0.002 | 243175 | 0.002
243209 | 0.002 | 243209 | 0.001 | 243173 | 0.002 | 243231 | 0.002 | 243292 | 0.002
243325 | 0.002 | 243325 | 0.001 | 243291 | 0.002 | 243348 | 0.002 | 243409 | 0.002
243443 | 0.002 | 243443 | 0.001 | 243408 | 0.002 | 243469 | 0.002 | 243527 | 0.002
243561 | 0.002 | 243564 | 0.001 | 243526 | 0.002 | 243587 | 0.002 | 243645 | 0.002
243679 | 0.002 | 243682 | 0.001 | 243643 | 0.002 | 243704 | 0.002 | 243762 | 0.002
243795 | 0.002 | 243800 | 0.001 | 243761 | 0.002 | 243822 | 0.002 | 243879 | 0.002
243913 | 0.002 | 243917 | 0.001 | 243878 | 0.002 | 243939 | 0.002 | 243997 | 0.002
244031 | 0.002 | 244034 | 0.001 | 243996 | 0.002 | 244057 | 0.003 | 244115 | 0.002
244149 | 0.002 | 244152 | 0.001 | 244114 | 0.002 | 244174 | 0.002 | 244232 | 0.002
244265 | 0.002 | 244270 | 0.001 | 244231 | 0.002 | 244292 | 0.002 | 244349 | 0.002
244383 | 0.002 | 244387 | 0.001 | 244348 | 0.002 | 244409 | 0.002 | 244467 | 0.002
244501 | 0.002 | 244504 | 0.001 | 244466 | 0.002 | 244527 | 0.002 | 244585 | 0.002
244619 | 0.002 | 244622 | 0.001 | 244584 | 0.002 | 244644 | 0.002 | 244702 | 0.002
244736 | 0.002 | 244740 | 0.001 | 244701 | 0.002 | 244762 | 0.002 | 244819 | 0.002
244853 | 0.002 | 244857 | 0.001 | 244818 | 0.002 | 244879 | 0.002 | 244937 | 0.002
244971 | 0.002 | 244974 | 0.001 | 244936 | 0.002 | 244997 | 0.002 | 245055 | 0.002
245089 | 0.002 | 245092 | 0.001 | 245054 | 0.002 | 245114 | 0.002 | 245172 | 0.002
245206 | 0.002 | 245210 | 0.001 | 245171 | 0.002 | 245232 | 0.002 | 245289 | 0.002
245323 | 0.002 | 245327 | 0.001 | 245288 | 0.003 | 245349 | 0.002 | 245407 | 0.002
245441 | 0.002 | 245444 | 0.001 | 245406 | 0.002 | 245467 | 0.002 | 245525 | 0.002
245559 | 0.002 | 245562 | 0.001 | 245524 | 0.002 | 245584 | 0.002 | 245642 | 0.002
245676 | 0.002 | 245680 | 0.001 | 245641 | 0.002 | 245702 | 0.002 | 245760 | 0.002
245793 | 0.002 | 245797 | 0.001 | 245758 | 0.003 | 245819 | 0.002 | 245877 | 0.002
245911 | 0.002 | 245914 | 0.001 | 245876 | 0.002 | 245937 | 0.002 | 245995 | 0.002
246029 | 0.002 | 246032 | 0.001 | 246011 | 0.002 | 246054 | 0.002 | 246112 | 0.002
246146 | 0.002 | 246150 | 0.001 | 246129 | 0.002 | 246172 | 0.002 | 246230 | 0.002
246263 | 0.002 | 246267 | 0.001 | 246247 | 0.003 | 246289 | 0.002 | 246347 | 0.002
246381 | 0.002 | 246384 | 0.001 | 246364 | 0.002 | 246407 | 0.002 | 246465 | 0.002
246499 | 0.002 | 246502 | 0.001 | 246481 | 0.002 | 246524 | 0.002 | 246582 | 0.002
246616 | 0.002 | 246620 | 0.001 | 246599 | 0.002 | 246642 | 0.002 | 246700 | 0.002
246733 | 0.002 | 246737 | 0.001 | 246717 | 0.002 | 246759 | 0.002 | 246817 | 0.002
246851 | 0.002 | 246854 | 0.001 | 246834 | 0.002 | 246877 | 0.002 | 246935 | 0.002
246969 | 0.002 | 246972 | 0.001 | 246951 | 0.002 | 246994 | 0.002 | 247052 | 0.002
247086 | 0.002 | 247090 | 0.001 | 247072 | 0.003 | 247112 | 0.002 | 247170 | 0.002
247203 | 0.002 | 247207 | 0.001 | 247190 | 0.002 | 247229 | 0.002 | 247287 | 0.002
247321 | 0.002 | 247324 | 0.001 | 247308 | 0.003 | 247347 | 0.002 | 247405 | 0.002
247442 | 0.002 | 247442 | 0.001 | 247426 | 0.002 | 247464 | 0.002 | 247526 | 0.002
247560 | 0.002 | 247560 | 0.001 | 247542 | 0.002 | 247582 | 0.002 | 247643 | 0.002
247677 | 0.002 | 247677 | 0.001 | 247660 | 0.002 | 247703 | 0.002 | 247760 | 0.002
247794 | 0.002 | 247798 | 0.001 | 247778 | 0.002 | 247820 | 0.002 | 247878 | 0.002
247912 | 0.002 | 247916 | 0.001 | 247896 | 0.002 | 247938 | 0.002 | 247996 | 0.002
248030 | 0.002 | 248033 | 0.001 | 248012 | 0.002 | 248055 | 0.002 | 248113 | 0.002
248147 | 0.002 | 248151 | 0.001 | 248130 | 0.002 | 248173 | 0.002 | 248230 | 0.002
248264 | 0.002 | 248268 | 0.001 | 248248 | 0.002 | 248290 | 0.002 | 248348 | 0.002
248382 | 0.002 | 248386 | 0.001 | 248366 | 0.002 | 248408 | 0.002 | 248466 | 0.002
248500 | 0.002 | 248503 | 0.001 | 248482 | 0.002 | 248525 | 0.002 | 248583 | 0.002
248617 | 0.002 | 248621 | 0.001 | 248600 | 0.002 | 248643 | 0.002 | 248700 | 0.002
248734 | 0.003 | 248738 | 0.001 | 248718 | 0.002 | 248760 | 0.002 | 248818 | 0.002
248852 | 0.002 | 248856 | 0.001 | 248836 | 0.002 | 248878 | 0.002 | 248936 | 0.002
248970 | 0.002 | 248973 | 0.001 | 248952 | 0.002 | 248995 | 0.002 | 249053 | 0.002
249088 | 0.002 | 249091 | 0.001 | 249070 | 0.002 | 249113 | 0.002 | 249170 | 0.002
249204 | 0.002 | 249208 | 0.001 | 249188 | 0.002 | 249230 | 0.002 | 249288 | 0.002
249322 | 0.002 | 249326 | 0.001 | 249306 | 0.002 | 249348 | 0.002 | 249406 | 0.002
249440 | 0.002 | 249443 | 0.001 | 249422 | 0.002 | 249465 | 0.002 | 249523 | 0.002
249558 | 0.002 | 249561 | 0.001 | 249540 | 0.002 | 249583 | 0.002 | 249640 | 0.002
249674 | 0.003 | 249678 | 0.001 | 249658 | 0.002 | 249700 | 0.002 | 249780 | 0.002
249792 | 0.002 | 249812 | 0.001 | 249776 | 0.002 | 249818 | 0.002 | 249896 | 0.002
249910 | 0.002 | 249930 | 0.001 | 249892 | 0.002 | 249935 | 0.002 | 250014 | 0.002
250028 | 0.002 | 250048 | 0.001 | 250010 | 0.002 | 250053 | 0.002 | 250132 | 0.002
250144 | 0.002 | 250165 | 0.001 | 250128 | 0.002 | 250170 | 0.002 | 250250 | 0.002
250262 | 0.002 | 250283 | 0.001 | 250246 | 0.002 | 250288 | 0.002 | 250366 | 0.002
250380 | 0.002 | 250400 | 0.001 | 250362 | 0.002 | 250405 | 0.002 | 250484 | 0.002
250498 | 0.002 | 250518 | 0.001 | 250480 | 0.003 | 250523 | 0.002 | 250602 | 0.002
250614 | 0.002 | 250635 | 0.001 | 250598 | 0.003 | 250640 | 0.002 | 250720 | 0.002
250732 | 0.002 | 250753 | 0.001 | 250716 | 0.002 | 250758 | 0.002 | 250836 | 0.002
250850 | 0.002 | 250870 | 0.001 | 250832 | 0.002 | 250875 | 0.002 | 250954 | 0.002
250968 | 0.002 | 250988 | 0.001 | 250950 | 0.002 | 250993 | 0.002 | 251072 | 0.002
251084 | 0.002 | 251105 | 0.001 | 251068 | 0.002 | 251110 | 0.002 | 251190 | 0.002
251202 | 0.002 | 251223 | 0.001 | 251186 | 0.002 | 251228 | 0.002 | 251307 | 0.002
251320 | 0.002 | 251340 | 0.001 | 251302 | 0.002 | 251345 | 0.002 | 251424 | 0.002
251438 | 0.002 | 251458 | 0.001 | 251423 | 0.002 | 251463 | 0.002 | 251542 | 0.002
251554 | 0.002 | 251575 | 0.001 | 251541 | 0.002 | 251580 | 0.002 | 251660 | 0.002
251675 | 0.002 | 251693 | 0.001 | 251659 | 0.002 | 251698 | 0.002 | 251781 | 0.002
251793 | 0.002 | 251810 | 0.001 | 251777 | 0.002 | 251815 | 0.002 | 251897 | 0.002
251911 | 0.002 | 251928 | 0.001 | 251894 | 0.002 | 251933 | 0.002 | 252015 | 0.002
252028 | 0.002 | 252045 | 0.001 | 252011 | 0.002 | 252054 | 0.002 | 252133 | 0.002
252145 | 0.002 | 252166 | 0.001 | 252129 | 0.002 | 252171 | 0.002 | 252251 | 0.002
252263 | 0.002 | 252283 | 0.001 | 252247 | 0.002 | 252289 | 0.002 | 252368 | 0.002
252381 | 0.002 | 252401 | 0.001 | 252364 | 0.002 | 252406 | 0.002 | 252485 | 0.002
252514 | 0.002 | 252519 | 0.001 | 252481 | 0.002 | 252524 | 0.002 | 252603 | 0.002
252632 | 0.002 | 252637 | 0.001 | 252599 | 0.002 | 252642 | 0.002 | 252721 | 0.002
252749 | 0.002 | 252753 | 0.001 | 252717 | 0.002 | 252759 | 0.002 | 252838 | 0.002
252866 | 0.002 | 252871 | 0.001 | 252834 | 0.002 | 252876 | 0.002 | 252955 | 0.002
252984 | 0.002 | 252989 | 0.001 | 252952 | 0.002 | 252994 | 0.002 | 253073 | 0.002
253102 | 0.002 | 253107 | 0.001 | 253069 | 0.002 | 253112 | 0.002 | 253191 | 0.002
253219 | 0.002 | 253223 | 0.001 | 253187 | 0.002 | 253229 | 0.002 | 253308 | 0.002
253337 | 0.002 | 253341 | 0.001 | 253304 | 0.003 | 253346 | 0.002 | 253426 | 0.002
253454 | 0.002 | 253459 | 0.001 | 253422 | 0.003 | 253464 | 0.002 | 253543 | 0.002
253572 | 0.002 | 253577 | 0.001 | 253539 | 0.002 | 253582 | 0.002 | 253661 | 0.002
253689 | 0.002 | 253693 | 0.001 | 253657 | 0.002 | 253700 | 0.002 | 253778 | 0.002
253807 | 0.002 | 253811 | 0.001 | 253774 | 0.002 | 253816 | 0.002 | 253896 | 0.002
253924 | 0.002 | 253929 | 0.001 | 253892 | 0.002 | 253934 | 0.002 | 254013 | 0.002
254042 | 0.002 | 254047 | 0.001 | 254010 | 0.002 | 254052 | 0.002 | 254131 | 0.002
254159 | 0.002 | 254164 | 0.001 | 254127 | 0.002 | 254170 | 0.002 | 254248 | 0.002
254277 | 0.002 | 254281 | 0.001 | 254244 | 0.002 | 254286 | 0.002 | 254366 | 0.002
254395 | 0.002 | 254399 | 0.001 | 254362 | 0.002 | 254404 | 0.002 | 254484 | 0.002
254512 | 0.002 | 254517 | 0.001 | 254480 | 0.002 | 254522 | 0.002 | 254601 | 0.002
254629 | 0.002 | 254635 | 0.001 | 254597 | 0.002 | 254659 | 0.002 | 254718 | 0.002
254747 | 0.002 | 254751 | 0.001 | 254714 | 0.002 | 254777 | 0.002 | 254836 | 0.002
254865 | 0.002 | 254869 | 0.001 | 254832 | 0.002 | 254894 | 0.002 | 254954 | 0.002
254982 | 0.002 | 254987 | 0.001 | 254950 | 0.002 | 255011 | 0.002 | 255071 | 0.002
255099 | 0.002 | 255105 | 0.001 | 255068 | 0.002 | 255129 | 0.002 | 255188 | 0.002
255217 | 0.002 | 255222 | 0.001 | 255184 | 0.002 | 255247 | 0.002 | 255306 | 0.002
255335 | 0.002 | 255339 | 0.001 | 255302 | 0.002 | 255365 | 0.002 | 255424 | 0.002
255453 | 0.002 | 255457 | 0.001 | 255420 | 0.002 | 255481 | 0.002 | 255542 | 0.002
255569 | 0.002 | 255575 | 0.001 | 255538 | 0.002 | 255599 | 0.002 | 255659 | 0.002
255687 | 0.002 | 255692 | 0.001 | 255654 | 0.002 | 255717 | 0.002 | 255776 | 0.002
255805 | 0.002 | 255809 | 0.001 | 255770 | 0.002 | 255835 | 0.002 | 255894 | 0.002
255923 | 0.002 | 255927 | 0.001 | 255888 | 0.002 | 255951 | 0.002 | 256012 | 0.002
256044 | 0.002 | 256045 | 0.001 | 256005 | 0.002 | 256069 | 0.002 | 256133 | 0.002
256161 | 0.002 | 256162 | 0.001 | 256122 | 0.002 | 256187 | 0.002 | 256249 | 0.002
256278 | 0.002 | 256280 | 0.001 | 256240 | 0.002 | 256308 | 0.002 | 256367 | 0.002
256396 | 0.002 | 256401 | 0.001 | 256358 | 0.002 | 256425 | 0.002 | 256485 | 0.002
256514 | 0.002 | 256518 | 0.001 | 256476 | 0.002 | 256543 | 0.002 | 256603 | 0.002
256631 | 0.002 | 256635 | 0.001 | 256592 | 0.002 | 256660 | 0.002 | 256720 | 0.002
256748 | 0.002 | 256753 | 0.001 | 256710 | 0.002 | 256778 | 0.002 | 256837 | 0.002
256866 | 0.002 | 256871 | 0.001 | 256828 | 0.002 | 256896 | 0.002 | 256955 | 0.002
256984 | 0.002 | 256989 | 0.001 | 256946 | 0.002 | 257013 | 0.002 | 257073 | 0.002
257101 | 0.002 | 257105 | 0.001 | 257062 | 0.002 | 257131 | 0.002 | 257190 | 0.002
257219 | 0.002 | 257223 | 0.001 | 257180 | 0.002 | 257248 | 0.002 | 257307 | 0.002
257336 | 0.002 | 257341 | 0.001 | 257298 | 0.002 | 257366 | 0.002 | 257425 | 0.002
257454 | 0.002 | 257459 | 0.001 | 257416 | 0.002 | 257483 | 0.002 | 257543 | 0.002
257571 | 0.002 | 257575 | 0.001 | 257533 | 0.002 | 257601 | 0.002 | 257660 | 0.002
257689 | 0.002 | 257693 | 0.001 | 257650 | 0.002 | 257718 | 0.002 | 257778 | 0.002
257807 | 0.002 | 257811 | 0.001 | 257768 | 0.002 | 257836 | 0.002 | 257895 | 0.002
257924 | 0.002 | 257929 | 0.001 | 257886 | 0.002 | 257953 | 0.002 | 258013 | 0.002
258041 | 0.002 | 258046 | 0.001 | 258003 | 0.002 | 258071 | 0.002 | 258130 | 0.002
258159 | 0.002 | 258163 | 0.001 | 258121 | 0.002 | 258189 | 0.002 | 258248 | 0.002
258287 | 0.002 | 258281 | 0.001 | 258239 | 0.002 | 258306 | 0.002 | 258376 | 0.002
258405 | 0.002 | 258399 | 0.001 | 258367 | 0.002 | 258423 | 0.002 | 258494 | 0.002
258523 | 0.002 | 258516 | 0.001 | 258484 | 0.002 | 258541 | 0.002 | 258612 | 0.002
258640 | 0.002 | 258634 | 0.001 | 258601 | 0.002 | 258669 | 0.002 | 258728 | 0.002
258757 | 0.002 | 258752 | 0.001 | 258719 | 0.002 | 258787 | 0.002 | 258846 | 0.002
258875 | 0.002 | 258880 | 0.001 | 258837 | 0.002 | 258905 | 0.002 | 258964 | 0.002
258993 | 0.002 | 258997 | 0.001 | 258972 | 0.002 | 259022 | 0.002 | 259082 | 0.002
259110 | 0.002 | 259115 | 0.001 | 259090 | 0.002 | 259139 | 0.002 | 259200 | 0.002
259227 | 0.002 | 259232 | 0.001 | 259208 | 0.002 | 259257 | 0.002 | 259316 | 0.002
259345 | 0.002 | 259350 | 0.001 | 259325 | 0.002 | 259375 | 0.002 | 259434 | 0.002
259463 | 0.002 | 259468 | 0.001 | 259442 | 0.002 | 259492 | 0.002 | 259552 | 0.002
259580 | 0.002 | 259585 | 0.001 | 259560 | 0.002 | 259609 | 0.002 | 259670 | 0.002
259698 | 0.002 | 259702 | 0.001 | 259678 | 0.002 | 259727 | 0.002 | 259786 | 0.002
259815 | 0.002 | 259820 | 0.001 | 259795 | 0.002 | 259845 | 0.002 | 259904 | 0.002
259933 | 0.002 | 259938 | 0.001 | 259912 | 0.002 | 259962 | 0.002 | 260022 | 0.002
260050 | 0.002 | 260055 | 0.001 | 260033 | 0.002 | 260080 | 0.002 | 260140 | 0.002
260168 | 0.002 | 260172 | 0.001 | 260151 | 0.002 | 260197 | 0.002 | 260257 | 0.002
260285 | 0.002 | 260290 | 0.001 | 260269 | 0.002 | 260315 | 0.002 | 260374 | 0.002
260407 | 0.002 | 260408 | 0.001 | 260386 | 0.002 | 260432 | 0.002 | 260495 | 0.002
260524 | 0.002 | 260526 | 0.001 | 260504 | 0.002 | 260550 | 0.002 | 260613 | 0.002
260641 | 0.002 | 260642 | 0.001 | 260621 | 0.002 | 260671 | 0.002 | 260730 | 0.002
260759 | 0.002 | 260763 | 0.002 | 260739 | 0.002 | 260789 | 0.002 | 260848 | 0.002
260877 | 0.002 | 260881 | 0.001 | 260856 | 0.002 | 260905 | 0.002 | 260966 | 0.002
260994 | 0.002 | 260999 | 0.001 | 260974 | 0.002 | 261023 | 0.002 | 261083 | 0.002
261111 | 0.002 | 261116 | 0.001 | 261091 | 0.002 | 261141 | 0.002 | 261200 | 0.002
261229 | 0.002 | 261233 | 0.001 | 261209 | 0.002 | 261259 | 0.002 | 261318 | 0.002
261347 | 0.002 | 261351 | 0.001 | 261326 | 0.002 | 261376 | 0.002 | 261436 | 0.002
261465 | 0.002 | 261469 | 0.001 | 261444 | 0.002 | 261493 | 0.002 | 261553 | 0.002
261581 | 0.002 | 261586 | 0.001 | 261562 | 0.002 | 261611 | 0.002 | 261670 | 0.002
261699 | 0.002 | 261704 | 0.001 | 261679 | 0.002 | 261729 | 0.002 | 261788 | 0.002
261817 | 0.002 | 261821 | 0.001 | 261796 | 0.002 | 261847 | 0.002 | 261906 | 0.002
261935 | 0.002 | 261939 | 0.001 | 261914 | 0.002 | 261963 | 0.002 | 262023 | 0.002
262051 | 0.002 | 262056 | 0.001 | 262032 | 0.002 | 262081 | 0.002 | 262140 | 0.002
262169 | 0.002 | 262174 | 0.001 | 262149 | 0.002 | 262199 | 0.002 | 262280 | 0.002
262287 | 0.002 | 262291 | 0.001 | 262266 | 0.002 | 262317 | 0.002 | 262397 | 0.002
262405 | 0.002 | 262409 | 0.001 | 262384 | 0.002 | 262433 | 0.002 | 262514 | 0.002
262522 | 0.002 | 262526 | 0.001 | 262502 | 0.002 | 262551 | 0.002 | 262632 | 0.002
262639 | 0.002 | 262644 | 0.001 | 262620 | 0.002 | 262669 | 0.002 | 262750 | 0.002
262757 | 0.002 | 262762 | 0.001 | 262736 | 0.002 | 262787 | 0.002 | 262867 | 0.002
262875 | 0.002 | 262896 | 0.001 | 262854 | 0.002 | 262904 | 0.002 | 262984 | 0.002
262992 | 0.002 | 263014 | 0.001 | 262972 | 0.002 | 263021 | 0.002 | 263102 | 0.002
263109 | 0.002 | 263131 | 0.001 | 263090 | 0.002 | 263139 | 0.002 | 263220 | 0.002
263227 | 0.002 | 263248 | 0.001 | 263206 | 0.002 | 263257 | 0.002 | 263337 | 0.002
263345 | 0.002 | 263366 | 0.001 | 263324 | 0.002 | 263374 | 0.002 | 263455 | 0.002
263462 | 0.002 | 263484 | 0.001 | 263442 | 0.002 | 263491 | 0.002 | 263572 | 0.002
263580 | 0.002 | 263602 | 0.001 | 263560 | 0.002 | 263609 | 0.002 | 263690 | 0.002
263697 | 0.002 | 263718 | 0.001 | 263677 | 0.002 | 263727 | 0.002 | 263808 | 0.002
263815 | 0.002 | 263836 | 0.001 | 263794 | 0.002 | 263844 | 0.002 | 263925 | 0.002
263932 | 0.002 | 263954 | 0.001 | 263912 | 0.002 | 263961 | 0.002 | 264042 | 0.002
264050 | 0.002 | 264072 | 0.001 | 264030 | 0.002 | 264079 | 0.002 | 264160 | 0.002
264167 | 0.002 | 264188 | 0.001 | 264147 | 0.002 | 264197 | 0.002 | 264278 | 0.002
264285 | 0.002 | 264306 | 0.001 | 264264 | 0.001 | 264314 | 0.002 | 264395 | 0.002
264402 | 0.002 | 264424 | 0.001 | 264386 | 0.002 | 264432 | 0.002 | 264512 | 0.002
264520 | 0.002 | 264542 | 0.001 | 264503 | 0.001 | 264549 | 0.002 | 264630 | 0.002
264641 | 0.002 | 264659 | 0.001 | 264621 | 0.001 | 264667 | 0.002 | 264751 | 0.002
264759 | 0.002 | 264776 | 0.001 | 264738 | 0.001 | 264784 | 0.002 | 264868 | 0.002
264876 | 0.002 | 264894 | 0.001 | 264856 | 0.001 | 264902 | 0.002 | 264986 | 0.002
264993 | 0.002 | 265012 | 0.001 | 264973 | 0.002 | 265023 | 0.002 | 265104 | 0.002
265111 | 0.002 | 265133 | 0.001 | 265091 | 0.002 | 265141 | 0.002 | 265221 | 0.002
265229 | 0.002 | 265250 | 0.001 | 265208 | 0.002 | 265257 | 0.002 | 265338 | 0.002
265347 | 0.002 | 265368 | 0.001 | 265326 | 0.001 | 265375 | 0.002 | 265456 | 0.002
265463 | 0.002 | 265485 | 0.001 | 265444 | 0.001 | 265493 | 0.002 | 265574 | 0.003
265581 | 0.002 | 265603 | 0.001 | 265561 | 0.002 | 265611 | 0.002 | 265692 | 0.002
265699 | 0.002 | 265720 | 0.001 | 265678 | 0.001 | 265728 | 0.002 | 265808 | 0.002
265832 | 0.002 | 265838 | 0.001 | 265796 | 0.002 | 265845 | 0.002 | 265926 | 0.002
265950 | 0.002 | 265955 | 0.001 | 265914 | 0.001 | 265963 | 0.002 | 266044 | 0.002
266067 | 0.002 | 266073 | 0.001 | 266031 | 0.001 | 266081 | 0.002 | 266162 | 0.002
266184 | 0.002 | 266190 | 0.001 | 266148 | 0.001 | 266199 | 0.002 | 266278 | 0.002
266302 | 0.002 | 266308 | 0.001 | 266266 | 0.001 | 266315 | 0.002 | 266396 | 0.002
266420 | 0.002 | 266426 | 0.001 | 266384 | 0.002 | 266433 | 0.002 | 266514 | 0.002
266537 | 0.002 | 266543 | 0.001 | 266502 | 0.001 | 266551 | 0.002 | 266632 | 0.002
266654 | 0.002 | 266660 | 0.001 | 266618 | 0.001 | 266669 | 0.002 | 266748 | 0.002
266772 | 0.002 | 266778 | 0.001 | 266736 | 0.001 | 266785 | 0.002 | 266866 | 0.002
266890 | 0.002 | 266896 | 0.001 | 266854 | 0.001 | 266903 | 0.002 | 266984 | 0.002
267008 | 0.001 | 267013 | 0.001 | 266972 | 0.001 | 267021 | 0.002 | 267102 | 0.002
267125 | 0.002 | 267130 | 0.002 | 267088 | 0.001 | 267139 | 0.002 | 267219 | 0.003
267242 | 0.002 | 267248 | 0.002 | 267206 | 0.001 | 267256 | 0.002 | 267336 | 0.002
267360 | 0.001 | 267366 | 0.002 | 267324 | 0.001 | 267393 | 0.002 | 267454 | 0.002
267478 | 0.001 | 267483 | 0.001 | 267442 | 0.001 | 267510 | 0.002 | 267572 | 0.002
267595 | 0.002 | 267600 | 0.001 | 267559 | 0.001 | 267628 | 0.002 | 267689 | 0.002
267712 | 0.002 | 267718 | 0.002 | 267676 | 0.001 | 267746 | 0.002 | 267806 | 0.002
267830 | 0.002 | 267836 | 0.001 | 267794 | 0.001 | 267864 | 0.002 | 267924 | 0.002
267948 | 0.002 | 267954 | 0.001 | 267912 | 0.001 | 267981 | 0.002 | 268042 | 0.002
268065 | 0.002 | 268070 | 0.001 | 268029 | 0.001 | 268098 | 0.002 | 268160 | 0.002
268183 | 0.002 | 268188 | 0.001 | 268146 | 0.001 | 268216 | 0.002 | 268277 | 0.002
268300 | 0.002 | 268306 | 0.001 | 268264 | 0.001 | 268334 | 0.002 | 268394 | 0.002
268418 | 0.002 | 268424 | 0.001 | 268382 | 0.001 | 268451 | 0.002 | 268512 | 0.002
268535 | 0.002 | 268540 | 0.001 | 268499 | 0.001 | 268568 | 0.002 | 268630 | 0.002
268653 | 0.002 | 268658 | 0.001 | 268620 | 0.001 | 268686 | 0.002 | 268747 | 0.002
268770 | 0.002 | 268776 | 0.001 | 268738 | 0.001 | 268804 | 0.002 | 268864 | 0.002
268888 | 0.001 | 268894 | 0.001 | 268855 | 0.001 | 268921 | 0.002 | 268982 | 0.002
269008 | 0.001 | 269011 | 0.001 | 268973 | 0.001 | 269038 | 0.002 | 269103 | 0.002
269126 | 0.001 | 269128 | 0.001 | 269090 | 0.001 | 269156 | 0.002 | 269220 | 0.002
269244 | 0.001 | 269246 | 0.001 | 269208 | 0.001 | 269277 | 0.002 | 269338 | 0.002
269362 | 0.001 | 269367 | 0.001 | 269326 | 0.001 | 269394 | 0.002 | 269456 | 0.002
269478 | 0.001 | 269485 | 0.001 | 269443 | 0.001 | 269512 | 0.002 | 269573 | 0.002
269596 | 0.001 | 269602 | 0.001 | 269560 | 0.001 | 269630 | 0.002 | 269690 | 0.002
269714 | 0.001 | 269720 | 0.001 | 269678 | 0.001 | 269747 | 0.002 | 269808 | 0.002
269832 | 0.002 | 269837 | 0.001 | 269796 | 0.001 | 269864 | 0.002 | 269926 | 0.002
269949 | 0.001 | 269955 | 0.001 | 269913 | 0.001 | 269982 | 0.002 | 270043 | 0.002
270066 | 0.001 | 270072 | 0.001 | 270030 | 0.001 | 270100 | 0.002 | 270160 | 0.002
270184 | 0.001 | 270190 | 0.001 | 270148 | 0.001 | 270218 | 0.002 | 270278 | 0.002
270302 | 0.001 | 270307 | 0.001 | 270266 | 0.001 | 270334 | 0.002 | 270396 | 0.002
270419 | 0.001 | 270425 | 0.001 | 270384 | 0.001 | 270452 | 0.002 | 270514 | 0.002
270536 | 0.001 | 270542 | 0.001 | 270500 | 0.001 | 270570 | 0.002 | 270630 | 0.002
270654 | 0.001 | 270660 | 0.001 | 270618 | 0.002 | 270688 | 0.002 | 270748 | 0.002
270772 | 0.001 | 270778 | 0.001 | 270736 | 0.001 | 270804 | 0.002 | 270866 | 0.002
270889 | 0.001 | 270895 | 0.001 | 270854 | 0.001 | 270922 | 0.002 | 270984 | 0.002
271007 | 0.001 | 271012 | 0.001 | 270970 | 0.001 | 271040 | 0.002 | 271100 | 0.002
271124 | 0.001 | 271130 | 0.001 | 271088 | 0.001 | 271158 | 0.002 | 271218 | 0.002
271242 | 0.001 | 271248 | 0.001 | 271206 | 0.001 | 271276 | 0.002 | 271336 | 0.002
271360 | 0.001 | 271365 | 0.001 | 271324 | 0.001 | 271392 | 0.002 | 271454 | 0.002
271477 | 0.001 | 271482 | 0.001 | 271441 | 0.001 | 271510 | 0.002 | 271571 | 0.002
271594 | 0.001 | 271600 | 0.001 | 271558 | 0.001 | 271628 | 0.002 | 271688 | 0.002
271712 | 0.001 | 271718 | 0.001 | 271676 | 0.001 | 271746 | 0.002 | 271806 | 0.002
271830 | 0.001 | 271836 | 0.001 | 271794 | 0.001 | 271862 | 0.002 | 271924 | 0.002
271947 | 0.001 | 271952 | 0.001 | 271929 | 0.001 | 271980 | 0.002 | 272041 | 0.002
272065 | 0.001 | 272070 | 0.001 | 272047 | 0.001 | 272098 | 0.002 | 272158 | 0.002
272182 | 0.001 | 272188 | 0.001 | 272164 | 0.001 | 272216 | 0.002 | 272276 | 0.002
272300 | 0.001 | 272306 | 0.001 | 272282 | 0.001 | 272332 | 0.002 | 272394 | 0.002
272417 | 0.001 | 272422 | 0.001 | 272399 | 0.001 | 272450 | 0.002 | 272512 | 0.002
272535 | 0.001 | 272540 | 0.001 | 272517 | 0.001 | 272568 | 0.002 | 272629 | 0.002
272652 | 0.001 | 272658 | 0.001 | 272634 | 0.001 | 272686 | 0.002 | 272746 | 0.002
272770 | 0.001 | 272776 | 0.001 | 272752 | 0.001 | 272803 | 0.002 | 272864 | 0.002
272887 | 0.001 | 272892 | 0.001 | 272870 | 0.001 | 272920 | 0.002 | 272982 | 0.002
273005 | 0.001 | 273010 | 0.001 | 272991 | 0.001 | 273038 | 0.002 | 273099 | 0.002
273123 | 0.001 | 273128 | 0.001 | 273107 | 0.001 | 273156 | 0.002 | 273216 | 0.002
273244 | 0.001 | 273246 | 0.001 | 273225 | 0.001 | 273273 | 0.002 | 273338 | 0.002
273360 | 0.001 | 273363 | 0.001 | 273343 | 0.001 | 273390 | 0.002 | 273455 | 0.002
273478 | 0.001 | 273480 | 0.001 | 273461 | 0.001 | 273508 | 0.002 | 273572 | 0.002
273596 | 0.001 | 273598 | 0.001 | 273579 | 0.001 | 273629 | 0.002 | 273690 | 0.002
273714 | 0.001 | 273719 | 0.001 | 273695 | 0.001 | 273746 | 0.002 | 273808 | 0.002
273831 | 0.001 | 273837 | 0.001 | 273813 | 0.001 | 273864 | 0.002 | 273925 | 0.002
273948 | 0.001 | 273954 | 0.001 | 273931 | 0.001 | 273982 | 0.002 | 274042 | 0.002
274066 | 0.001 | 274072 | 0.001 | 274049 | 0.001 | 274099 | 0.002 | 274160 | 0.002
274184 | 0.001 | 274189 | 0.001 | 274165 | 0.001 | 274216 | 0.002 | 274278 | 0.002
274301 | 0.001 | 274307 | 0.001 | 274283 | 0.001 | 274334 | 0.002 | 274396 | 0.002
274418 | 0.001 | 274424 | 0.001 | 274401 | 0.001 | 274452 | 0.002 | 274512 | 0.002
274536 | 0.001 | 274542 | 0.001 | 274519 | 0.001 | 274570 | 0.002 | 274630 | 0.002
274654 | 0.001 | 274659 | 0.001 | 274636 | 0.001 | 274686 | 0.002 | 274769 | 0.002
274771 | 0.001 | 274777 | 0.001 | 274753 | 0.001 | 274804 | 0.002 | 274886 | 0.002
274889 | 0.001 | 274894 | 0.001 | 274871 | 0.001 | 274922 | 0.002 | 275004 | 0.002
275006 | 0.001 | 275012 | 0.001 | 274989 | 0.001 | 275040 | 0.002 | 275122 | 0.002
275124 | 0.001 | 275130 | 0.001 | 275106 | 0.001 | 275156 | 0.002 | 275239 | 0.002
275241 | 0.001 | 275247 | 0.001 | 275223 | 0.001 | 275274 | 0.002 | 275356 | 0.002
275359 | 0.001 | 275364 | 0.001 | 275341 | 0.001 | 275392 | 0.002 | 275474 | 0.002
275476 | 0.001 | 275482 | 0.001 | 275459 | 0.001 | 275510 | 0.002 | 275592 | 0.002
275594 | 0.001 | 275600 | 0.001 | 275576 | 0.001 | 275628 | 0.002 | 275709 | 0.002
275712 | 0.001 | 275717 | 0.001 | 275694 | 0.001 | 275744 | 0.002 | 275827 | 0.002
275829 | 0.001 | 275834 | 0.001 | 275811 | 0.001 | 275862 | 0.002 | 275944 | 0.002
275947 | 0.001 | 275952 | 0.001 | 275929 | 0.001 | 275980 | 0.002 | 276062 | 0.002
276064 | 0.001 | 276086 | 0.001 | 276046 | 0.001 | 276098 | 0.002 | 276179 | 0.002
276182 | 0.001 | 276204 | 0.001 | 276164 | 0.001 | 276214 | 0.002 | 276297 | 0.002
276299 | 0.001 | 276322 | 0.001 | 276281 | 0.001 | 276332 | 0.002 | 276414 | 0.002
276417 | 0.001 | 276440 | 0.001 | 276399 | 0.001 | 276450 | 0.002 | 276532 | 0.002
276534 | 0.001 | 276556 | 0.001 | 276516 | 0.001 | 276568 | 0.002 | 276649 | 0.002
276652 | 0.001 | 276674 | 0.001 | 276634 | 0.001 | 276684 | 0.002 | 276767 | 0.002
276769 | 0.001 | 276792 | 0.001 | 276752 | 0.001 | 276802 | 0.002 | 276885 | 0.002
276887 | 0.001 | 276910 | 0.001 | 276869 | 0.001 | 276920 | 0.002 | 277002 | 0.002
277005 | 0.001 | 277026 | 0.001 | 276986 | 0.001 | 277038 | 0.002 | 277120 | 0.002
277122 | 0.001 | 277144 | 0.001 | 277104 | 0.001 | 277155 | 0.002 | 277237 | 0.002
277239 | 0.001 | 277262 | 0.001 | 277225 | 0.001 | 277272 | 0.002 | 277355 | 0.002
277357 | 0.001 | 277380 | 0.001 | 277342 | 0.001 | 277390 | 0.002 | 277472 | 0.002
277475 | 0.001 | 277497 | 0.001 | 277459 | 0.001 | 277508 | 0.002 | 277590 | 0.002
277596 | 0.001 | 277614 | 0.001 | 277577 | 0.001 | 277625 | 0.002 | 277710 | 0.002
277712 | 0.001 | 277732 | 0.001 | 277694 | 0.001 | 277742 | 0.002 | 277828 | 0.002
277830 | 0.001 | 277850 | 0.001 | 277812 | 0.001 | 277864 | 0.002 | 277946 | 0.002
277948 | 0.001 | 277970 | 0.001 | 277928 | 0.001 | 277981 | 0.002 | 278063 | 0.002
278066 | 0.001 | 278088 | 0.001 | 278046 | 0.001 | 278098 | 0.002 | 278180 | 0.002
278182 | 0.001 | 278205 | 0.001 | 278164 | 0.001 | 278216 | 0.002 | 278298 | 0.002
278300 | 0.001 | 278323 | 0.001 | 278281 | 0.001 | 278333 | 0.002 | 278416 | 0.002
278418 | 0.001 | 278440 | 0.001 | 278398 | 0.001 | 278451 | 0.002 | 278533 | 0.002
278536 | 0.001 | 278558 | 0.001 | 278515 | 0.001 | 278568 | 0.002 | 278650 | 0.002
278652 | 0.001 | 278675 | 0.001 | 278633 | 0.001 | 278686 | 0.002 | 278768 | 0.002
278770 | 0.001 | 278793 | 0.001 | 278751 | 0.001 | 278803 | 0.002 | 278886 | 0.002
278888 | 0.001 | 278910 | 0.001 | 278867 | 0.001 | 278921 | 0.002 | 279003 | 0.002
279005 | 0.001 | 279027 | 0.001 | 278985 | 0.001 | 279038 | 0.002 | 279120 | 0.002
279138 | 0.001 | 279145 | 0.001 | 279102 | 0.001 | 279155 | 0.002 | 279238 | 0.002
279256 | 0.001 | 279263 | 0.001 | 279220 | 0.001 | 279273 | 0.002 | 279355 | 0.002
279373 | 0.001 | 279380 | 0.001 | 279337 | 0.001 | 279391 | 0.002 | 279473 | 0.002
279490 | 0.001 | 279497 | 0.001 | 279454 | 0.001 | 279508 | 0.002 | 279590 | 0.002
279608 | 0.001 | 279615 | 0.001 | 279572 | 0.001 | 279625 | 0.002 | 279708 | 0.002
279725 | 0.001 | 279733 | 0.001 | 279689 | 0.001 | 279743 | 0.002 | 279825 | 0.002
279842 | 0.001 | 279849 | 0.001 | 279806 | 0.001 | 279861 | 0.002 | 279943 | 0.002
279960 | 0.001 | 279967 | 0.001 | 279924 | 0.001 | 279979 | 0.002 | 280060 | 0.002
280077 | 0.001 | 280085 | 0.001 | 280041 | 0.001 | 280115 | 0.002 | 280177 | 0.002
280195 | 0.001 | 280203 | 0.001 | 280159 | 0.001 | 280232 | 0.002 | 280295 | 0.002
280311 | 0.001 | 280320 | 0.001 | 280275 | 0.001 | 280350 | 0.002 | 280413 | 0.002
280429 | 0.001 | 280437 | 0.001 | 280393 | 0.001 | 280467 | 0.002 | 280530 | 0.002
280547 | 0.001 | 280555 | 0.001 | 280511 | 0.001 | 280585 | 0.002 | 280647 | 0.002
280664 | 0.001 | 280673 | 0.001 | 280628 | 0.001 | 280701 | 0.002 | 280765 | 0.002
280781 | 0.001 | 280790 | 0.001 | 280745 | 0.001 | 280819 | 0.002 | 280883 | 0.002
280898 | 0.001 | 280907 | 0.002 | 280862 | 0.001 | 280937 | 0.002 | 281000 | 0.002
281016 | 0.001 | 281025 | 0.001 | 280980 | 0.001 | 281054 | 0.002 | 281117 | 0.002
281134 | 0.001 | 281142 | 0.001 | 281098 | 0.001 | 281171 | 0.002 | 281235 | 0.002
281250 | 0.001 | 281260 | 0.001 | 281214 | 0.001 | 281288 | 0.002 | 281353 | 0.002
281368 | 0.001 | 281377 | 0.001 | 281332 | 0.001 | 281406 | 0.002 | 281469 | 0.002
281485 | 0.001 | 281495 | 0.001 | 281449 | 0.001 | 281524 | 0.002 | 281587 | 0.002
281603 | 0.001 | 281612 | 0.001 | 281579 | 0.001 | 281640 | 0.002 | 281705 | 0.002
281720 | 0.001 | 281730 | 0.001 | 281699 | 0.001 | 281758 | 0.002 | 281823 | 0.002
281849 | 0.001 | 281847 | 0.001 | 281819 | 0.001 | 281875 | 0.002 | 281952 | 0.002
281969 | 0.001 | 281964 | 0.001 | 281938 | 0.001 | 281993 | 0.002 | 282072 | 0.002
282089 | 0.001 | 282082 | 0.001 | 282058 | 0.001 | 282110 | 0.002 | 282191 | 0.002
282209 | 0.001 | 282200 | 0.001 | 282178 | 0.001 | 282240 | 0.002 | 282311 | 0.002
282329 | 0.001 | 282330 | 0.001 | 282297 | 0.001 | 282359 | 0.002 | 282431 | 0.002
282448 | 0.001 | 282449 | 0.001 | 282417 | 0.001 | 282479 | 0.002 | 282551 | 0.002
282568 | 0.001 | 282569 | 0.001 | 282536 | 0.001 | 282599 | 0.002 | 282671 | 0.002
282688 | 0.001 | 282688 | 0.001 | 282656 | 0.001 | 282719 | 0.002 | 282790 | 0.002
282808 | 0.001 | 282808 | 0.001 | 282776 | 0.001 | 282839 | 0.002 | 282910 | 0.002
282928 | 0.001 | 282928 | 0.001 | 282896 | 0.001 | 282958 | 0.002 | 283030 | 0.002
283047 | 0.001 | 283047 | 0.001 | 283016 | 0.001 | 283078 | 0.002 | 283150 | 0.002
283166 | 0.001 | 283167 | 0.001 | 283136 | 0.001 | 283198 | 0.002 | 283270 | 0.002
283286 | 0.001 | 283287 | 0.001 | 283255 | 0.001 | 283318 | 0.002 | 283388 | 0.002
283406 | 0.001 | 283407 | 0.001 | 283375 | 0.001 | 283437 | 0.002 | 283508 | 0.002
283526 | 0.001 | 283527 | 0.001 | 283495 | 0.001 | 283556 | 0.002 | 283628 | 0.002
283645 | 0.001 | 283646 | 0.001 | 283615 | 0.001 | 283676 | 0.001 | 283748 | 0.002
283765 | 0.001 | 283766 | 0.001 | 283735 | 0.001 | 283796 | 0.001 | 283868 | 0.002
283885 | 0.001 | 283886 | 0.001 | 283853 | 0.001 | 283916 | 0.002 | 283987 | 0.002
284005 | 0.001 | 284006 | 0.001 | 283973 | 0.001 | 284036 | 0.002 | 284107 | 0.002
284125 | 0.001 | 284126 | 0.001 | 284093 | 0.001 | 284155 | 0.002 | 284227 | 0.002
284244 | 0.001 | 284244 | 0.001 | 284213 | 0.001 | 284275 | 0.002 | 284347 | 0.002
284364 | 0.001 | 284364 | 0.001 | 284333 | 0.001 | 284395 | 0.001 | 284467 | 0.002
284484 | 0.001 | 284484 | 0.001 | 284452 | 0.001 | 284515 | 0.002 | 284587 | 0.002
284604 | 0.001 | 284604 | 0.001 | 284572 | 0.001 | 284635 | 0.001 | 284706 | 0.002
284723 | 0.001 | 284724 | 0.001 | 284692 | 0.001 | 284754 | 0.001 | 284825 | 0.002
284843 | 0.001 | 284844 | 0.001 | 284812 | 0.001 | 284873 | 0.001 | 284945 | 0.002
284962 | 0.001 | 284963 | 0.001 | 284950 | 0.001 | 284993 | 0.001 | 285065 | 0.002
285082 | 0.001 | 285083 | 0.001 | 285070 | 0.001 | 285113 | 0.001 | 285185 | 0.002
285202 | 0.001 | 285203 | 0.001 | 285190 | 0.001 | 285233 | 0.001 | 285304 | 0.002
285322 | 0.001 | 285323 | 0.001 | 285309 | 0.001 | 285352 | 0.001 | 285424 | 0.002
285442 | 0.001 | 285443 | 0.001 | 285429 | 0.001 | 285472 | 0.001 | 285544 | 0.002
285561 | 0.001 | 285562 | 0.001 | 285549 | 0.001 | 285592 | 0.001 | 285664 | 0.002
285681 | 0.001 | 285682 | 0.001 | 285669 | 0.001 | 285712 | 0.002 | 285784 | 0.002
285801 | 0.001 | 285801 | 0.001 | 285788 | 0.001 | 285832 | 0.001 | 285903 | 0.002
285921 | 0.001 | 285921 | 0.001 | 285907 | 0.001 | 285952 | 0.001 | 286023 | 0.002
286041 | 0.001 | 286041 | 0.001 | 286025 | 0.001 | 286071 | 0.001 | 286143 | 0.002
286160 | 0.001 | 286160 | 0.001 | 286145 | 0.001 | 286191 | 0.001 | 286263 | 0.002
286283 | 0.001 | 286280 | 0.001 | 286265 | 0.001 | 286311 | 0.001 | 286386 | 0.002
286403 | 0.001 | 286400 | 0.001 | 286385 | 0.001 | 286430 | 0.001 | 286505 | 0.002
286523 | 0.001 | 286520 | 0.001 | 286504 | 0.001 | 286554 | 0.001 | 286625 | 0.002
286643 | 0.001 | 286643 | 0.001 | 286624 | 0.001 | 286673 | 0.001 | 286745 | 0.002
286763 | 0.001 | 286763 | 0.001 | 286744 | 0.001 | 286793 | 0.001 | 286865 | 0.002
286882 | 0.001 | 286882 | 0.001 | 286865 | 0.001 | 286913 | 0.001 | 286985 | 0.002
287002 | 0.001 | 287002 | 0.001 | 286985 | 0.001 | 287033 | 0.001 | 287105 | 0.002
287133 | 0.001 | 287122 | 0.001 | 287105 | 0.001 | 287153 | 0.001 | 287236 | 0.002
287253 | 0.001 | 287243 | 0.001 | 287235 | 0.001 | 287272 | 0.002 | 287377 | 0.002
287373 | 0.001 | 287363 | 0.001 | 287354 | 0.001 | 287392 | 0.001 | 287497 | 0.002
287493 | 0.001 | 287482 | 0.001 | 287474 | 0.001 | 287522 | 0.001 | 287617 | 0.002
287612 | 0.001 | 287602 | 0.001 | 287594 | 0.001 | 287642 | 0.001 | 287736 | 0.002
287732 | 0.001 | 287732 | 0.001 | 287714 | 0.001 | 287762 | 0.001 | 287856 | 0.002
287852 | 0.001 | 287852 | 0.001 | 287834 | 0.001 | 287883 | 0.001 | 287976 | 0.002
287972 | 0.001 | 287972 | 0.001 | 287953 | 0.001 | 288003 | 0.001 | 288096 | 0.002
288092 | 0.001 | 288092 | 0.001 | 288073 | 0.001 | 288122 | 0.001 | 288216 | 0.002
288211 | 0.001 | 288211 | 0.001 | 288193 | 0.001 | 288242 | 0.001 | 288335 | 0.002
288331 | 0.001 | 288331 | 0.001 | 288313 | 0.001 | 288362 | 0.001 | 288455 | 0.002
288451 | 0.001 | 288451 | 0.001 | 288433 | 0.001 | 288482 | 0.001 | 288575 | 0.002
288571 | 0.001 | 288571 | 0.001 | 288552 | 0.001 | 288602 | 0.001 | 288695 | 0.002
288691 | 0.001 | 288691 | 0.001 | 288672 | 0.001 | 288721 | 0.001 | 288815 | 0.002
288811 | 0.001 | 288811 | 0.001 | 288792 | 0.001 | 288841 | 0.001 | 288934 | 0.002
288930 | 0.001 | 288930 | 0.001 | 288912 | 0.001 | 288961 | 0.001 | 289054 | 0.002
289050 | 0.001 | 289050 | 0.001 | 289032 | 0.001 | 289081 | 0.001 | 289174 | 0.002
289170 | 0.001 | 289170 | 0.001 | 289152 | 0.001 | 289201 | 0.001 | 289294 | 0.002
289290 | 0.001 | 289307 | 0.001 | 289271 | 0.001 | 289320 | 0.001 | 289414 | 0.002
289410 | 0.001 | 289427 | 0.001 | 289391 | 0.001 | 289440 | 0.001 | 289533 | 0.002
289529 | 0.001 | 289547 | 0.001 | 289511 | 0.001 | 289560 | 0.001 | 289653 | 0.002
289649 | 0.001 | 289667 | 0.001 | 289631 | 0.001 | 289680 | 0.001 | 289773 | 0.002
289769 | 0.001 | 289786 | 0.001 | 289752 | 0.001 | 289800 | 0.001 | 289893 | 0.002
289889 | 0.001 | 289906 | 0.001 | 289871 | 0.001 | 289920 | 0.001 | 290013 | 0.002
290009 | 0.001 | 290026 | 0.001 | 289991 | 0.001 | 290039 | 0.001 | 290132 | 0.002
290128 | 0.001 | 290146 | 0.001 | 290111 | 0.001 | 290159 | 0.001 | 290252 | 0.002
290248 | 0.001 | 290267 | 0.001 | 290231 | 0.001 | 290279 | 0.001 | 290372 | 0.002
290368 | 0.001 | 290386 | 0.001 | 290354 | 0.001 | 290399 | 0.001 | 290492 | 0.002
290488 | 0.001 | 290506 | 0.001 | 290473 | 0.001 | 290519 | 0.001 | 290612 | 0.002
290608 | 0.001 | 290626 | 0.001 | 290593 | 0.001 | 290638 | 0.001 | 290732 | 0.002
290732 | 0.001 | 290746 | 0.001 | 290713 | 0.001 | 290758 | 0.001 | 290855 | 0.002
290851 | 0.001 | 290866 | 0.001 | 290833 | 0.001 | 290878 | 0.001 | 290975 | 0.002
290971 | 0.001 | 290985 | 0.001 | 290953 | 0.001 | 291001 | 0.001 | 291095 | 0.002
291091 | 0.001 | 291108 | 0.001 | 291073 | 0.001 | 291121 | 0.001 | 291215 | 0.002
291211 | 0.001 | 291228 | 0.001 | 291192 | 0.001 | 291240 | 0.001 | 291335 | 0.002
291331 | 0.001 | 291348 | 0.001 | 291312 | 0.001 | 291360 | 0.001 | 291454 | 0.002
291450 | 0.001 | 291468 | 0.001 | 291432 | 0.001 | 291480 | 0.001 | 291574 | 0.002
291570 | 0.001 | 291588 | 0.001 | 291552 | 0.001 | 291600 | 0.001 | 291694 | 0.002
291690 | 0.001 | 291707 | 0.001 | 291672 | 0.001 | 291720 | 0.001 | 291814 | 0.002
291810 | 0.001 | 291827 | 0.001 | 291791 | 0.001 | 291840 | 0.002 | 291934 | 0.002
291930 | 0.001 | 291947 | 0.001 | 291911 | 0.001 | 291959 | 0.001 | 292054 | 0.002
292049 | 0.001 | 292067 | 0.001 | 292031 | 0.001 | 292079 | 0.001 | 292173 | 0.002
292169 | 0.001 | 292187 | 0 | 292151 | 0.001 | 292199 | 0.002 | 292293 | 0.002
292289 | 0.001 | 292306 | 0.001 | 292271 | 0.001 | 292319 | 0.002 | 292413 | 0.002
292409 | 0.001 | 292426 | 0.001 | 292390 | 0.001 | 292439 | 0.001 | 292533 | 0.002
292529 | 0.001 | 292546 | 0.001 | 292510 | 0.001 | 292558 | 0.001 | 292653 | 0.002
292665 | 0.001 | 292666 | 0.001 | 292630 | 0.001 | 292679 | 0.001 | 292772 | 0.002
292785 | 0.001 | 292786 | 0.001 | 292750 | 0.001 | 292799 | 0.001 | 292892 | 0.002
292904 | 0.001 | 292905 | 0.001 | 292870 | 0.001 | 292919 | 0.002 | 293012 | 0.002
293024 | 0.001 | 293025 | 0.001 | 292989 | 0.001 | 293058 | 0.001 | 293132 | 0.002
293144 | 0.001 | 293145 | 0.001 | 293109 | 0.001 | 293178 | 0.001 | 293252 | 0.002
293264 | 0.001 | 293265 | 0.001 | 293229 | 0.001 | 293298 | 0.001 | 293371 | 0.002
293384 | 0.001 | 293385 | 0.001 | 293349 | 0.001 | 293417 | 0.002 | 293491 | 0.002
293504 | 0.001 | 293505 | 0.001 | 293469 | 0.001 | 293537 | 0.002 | 293611 | 0.002
293623 | 0.001 | 293624 | 0.001 | 293588 | 0.001 | 293657 | 0.002 | 293731 | 0.002
293743 | 0.001 | 293744 | 0.001 | 293708 | 0.001 | 293777 | 0.001 | 293851 | 0.002
293863 | 0.001 | 293865 | 0.001 | 293828 | 0.001 | 293897 | 0.001 | 293970 | 0.002
293983 | 0.001 | 293985 | 0.001 | 293948 | 0.001 | 294017 | 0.001 | 294090 | 0.002
294103 | 0.001 | 294105 | 0.001 | 294069 | 0.001 | 294136 | 0.002 | 294210 | 0.002
294222 | 0.001 | 294224 | 0.001 | 294189 | 0.002 | 294256 | 0.001 | 294330 | 0.002
294342 | 0.001 | 294344 | 0.001 | 294308 | 0.002 | 294376 | 0.001 | 294450 | 0.002
294462 | 0.001 | 294464 | 0.001 | 294428 | 0.001 | 294496 | 0.001 | 294569 | 0.002
294582 | 0.001 | 294584 | 0.001 | 294548 | 0.001 | 294616 | 0.001 | 294689 | 0.002
294702 | 0.001 | 294704 | 0.001 | 294668 | 0.001 | 294736 | 0.002 | 294809 | 0.002
294821 | 0.001 | 294823 | 0 | 294791 | 0.001 | 294856 | 0.002 | 294929 | 0.002
294941 | 0.001 | 294943 | 0 | 294910 | 0.001 | 294976 | 0.001 | 295050 | 0.002
295065 | 0.001 | 295063 | 0.001 | 295030 | 0.001 | 295096 | 0.001 | 295173 | 0.002
295185 | 0.001 | 295183 | 0.001 | 295150 | 0.001 | 295216 | 0.002 | 295292 | 0.002
295305 | 0.001 | 295303 | 0.001 | 295270 | 0.001 | 295335 | 0.002 | 295412 | 0.002
295425 | 0.001 | 295422 | 0.001 | 295390 | 0.001 | 295458 | 0.002 | 295532 | 0.002
295544 | 0.001 | 295545 | 0.001 | 295510 | 0.001 | 295578 | 0.002 | 295652 | 0.002
295664 | 0.001 | 295665 | 0.001 | 295629 | 0.001 | 295698 | 0.001 | 295772 | 0.002
295784 | 0.001 | 295785 | 0.001 | 295749 | 0.001 | 295818 | 0.001 | 295891 | 0.002
295904 | 0.001 | 295905 | 0.001 | 295869 | 0.001 | 295938 | 0.002 | 296011 | 0.002
296024 | 0.001 | 296025 | 0.001 | 295989 | 0.001 | 296057 | 0.002 | 296131 | 0.002
296143 | 0.001 | 296144 | 0.001 | 296109 | 0.001 | 296177 | 0.002 | 296251 | 0.002
296263 | 0.001 | 296264 | 0 | 296228 | 0.001 | 296297 | 0.002 | 296371 | 0.002
296383 | 0.001 | 296384 | 0.001 | 296348 | 0.001 | 296417 | 0.002 | 296491 | 0.002
296503 | 0.001 | 296504 | 0.001 | 296468 | 0.001 | 296537 | 0.002 | 296610 | 0.002
296623 | 0.001 | 296624 | 0.001 | 296588 | 0.002 | 296656 | 0.002 | 296730 | 0.002
296742 | 0.001 | 296743 | 0.001 | 296708 | 0.002 | 296776 | 0.002 | 296850 | 0.002
296862 | 0.001 | 296863 | 0 | 296827 | 0.001 | 296896 | 0.002 | 296970 | 0.002
296982 | 0.001 | 296983 | 0.001 | 296947 | 0.001 | 297016 | 0.002 | 297090 | 0.002
297102 | 0.001 | 297103 | 0 | 297067 | 0.001 | 297136 | 0.002 | 297209 | 0.002
297222 | 0.001 | 297223 | 0 | 297187 | 0.002 | 297255 | 0.002 | 297329 | 0.002
297341 | 0.001 | 297344 | 0.001 | 297307 | 0.002 | 297375 | 0.002 | 297449 | 0.002
297461 | 0.001 | 297463 | 0 | 297426 | 0.002 | 297495 | 0.002 | 297569 | 0.002
297581 | 0.001 | 297583 | 0.001 | 297546 | 0.002 | 297615 | 0.002 | 297689 | 0.002
297701 | 0.001 | 297703 | 0.001 | 297666 | 0.002 | 297735 | 0.002 | 297808 | 0.002
297821 | 0.001 | 297823 | 0.001 | 297786 | 0.002 | 297854 | 0.002 | 297928 | 0.002
297940 | 0.001 | 297943 | 0.001 | 297906 | 0.002 | 297974 | 0.002 | 298048 | 0.002
298060 | 0.001 | 298062 | 0.001 | 298025 | 0.002 | 298094 | 0.002 | 298168 | 0.002
298180 | 0.001 | 298182 | 0.001 | 298165 | 0.002 | 298214 | 0.002 | 298288 | 0.002
298300 | 0.001 | 298302 | 0.001 | 298285 | 0.002 | 298334 | 0.002 | 298407 | 0.002
298420 | 0.001 | 298422 | 0.001 | 298404 | 0.002 | 298454 | 0.002 | 298528 | 0.002
298540 | 0.001 | 298542 | 0.001 | 298524 | 0.002 | 298573 | 0.002 | 298648 | 0.002
298660 | 0.001 | 298661 | 0.001 | 298644 | 0.002 | 298693 | 0.002 | 298768 | 0.002
298780 | 0.001 | 298781 | 0.001 | 298764 | 0.002 | 298813 | 0.002 | 298888 | 0.002
298900 | 0.001 | 298901 | 0.001 | 298884 | 0.001 | 298934 | 0.002 | 299008 | 0.002
299020 | 0.001 | 299021 | 0.001 | 299003 | 0.001 | 299054 | 0.002 | 299127 | 0.002
299140 | 0.001 | 299141 | 0.001 | 299126 | 0.002 | 299173 | 0.002 | 299247 | 0.002
299259 | 0.001 | 299260 | 0.001 | 299246 | 0.002 | 299293 | 0.002 | 299367 | 0.002
299379 | 0.001 | 299380 | 0.001 | 299366 | 0.002 | 299413 | 0.002 | 299487 | 0.002
299502 | 0.001 | 299500 | 0.001 | 299486 | 0.002 | 299533 | 0.002 | 299610 | 0.002
299622 | 0.001 | 299620 | 0.001 | 299606 | 0.002 | 299653 | 0.002 | 299729 | 0.002
299742 | 0.001 | 299740 | 0.001 | 299725 | 0.002 | 299776 | 0.002 | 299849 | 0.002
299862 | 0.001 | 299863 | 0.001 | 299845 | 0.002 | 299895 | 0.002 | 299969 | 0.002
299981 | 0.001 | 299982 | 0.001 | 299965 | 0.002 | 300015 | 0.002 | 300110 | 0.002
300101 | 0.001 | 300102 | 0.001 | 300085 | 0.002 | 300135 | 0.002 | 300230 | 0.002
300221 | 0.001 | 300222 | 0.001 | 300205 | 0.002 | 300255 | 0.002 | 300350 | 0.002
300341 | 0.001 | 300342 | 0.001 | 300324 | 0.002 | 300375 | 0.002 | 300470 | 0.002
300461 | 0.001 | 300462 | 0.001 | 300444 | 0.002 | 300494 | 0.002 | 300590 | 0.002
300580 | 0.001 | 300581 | 0.001 | 300564 | 0.002 | 300614 | 0.002 | 300709 | 0.002
300700 | 0.001 | 300701 | 0.001 | 300684 | 0.002 | 300734 | 0.002 | 300829 | 0.002
300820 | 0.001 | 300822 | 0.001 | 300804 | 0.001 | 300854 | 0.002 | 300949 | 0.002
300940 | 0.001 | 300942 | 0.001 | 300923 | 0.002 | 300974 | 0.002 | 301069 | 0.002
301060 | 0.001 | 301062 | 0 | 301043 | 0.001 | 301093 | 0.002 | 301189 | 0.002
301179 | 0.001 | 301181 | 0 | 301163 | 0.001 | 301213 | 0.002 | 301308 | 0.002
301299 | 0.001 | 301301 | 0.011 | 301283 | 0.002 | 301333 | 0.002 | 301428 | 0.007
301419 | 0.034 | 301421 | 0.153 | 301403 | 0.002 | 301453 | 0.002 | 301549 | 0.404
301539 | 0.24 | 301541 | 0.905 | 301523 | 0.001 | 301573 | 0.002 | 301669 | 0.858
301659 | 0.638 | 301661 | 0.999 | 301643 | 0.012 | 301692 | 0.002 | 301789 | 0.999
301778 | 1 | 301780 | 1 | 301763 | 0.178 | 301812 | 0.008 | 301908 | 1.002
301898 | 1.001 | 301900 | 1 | 301883 | 0.847 | 301932 | 0.138 | 302028 | 1.002
302019 | 1.001 | 302020 | 1 | 302003 | 0.994 | 302052 | 0.883 | 302148 | 1.002
302139 | 1.001 | 302140 | 1 | 302123 | 1.001 | 302172 | 0.999 | 302268 | 1.002
302259 | 1.001 | 302260 | 1 | 302242 | 1.001 | 302291 | 1.002 | 302388 | 1.002
302379 | 1.001 | 302380 | 1 | 302362 | 1.001 | 302411 | 1.002 | 302508 | 1.002
302498 | 1.001 | 302499 | 1 | 302482 | 1.001 | 302531 | 1.002 | 302627 | 1.002
302618 | 1.001 | 302637 | 1 | 302602 | 1.001 | 302651 | 1.002 | 302747 | 1.002
302738 | 1.001 | 302756 | 1 | 302722 | 1.001 | 302771 | 1.002 | 302867 | 1.002
302858 | 1.001 | 302876 | 1 | 302841 | 1.001 | 302892 | 1.002 | 302987 | 1.002
302978 | 1.001 | 302996 | 1 | 302961 | 1.001 | 303011 | 1.002 | 303107 | 1.002
303097 | 1.001 | 303116 | 1 | 303081 | 1.001 | 303131 | 1.002 | 303226 | 1.002
303217 | 1.001 | 303236 | 1 | 303201 | 1.001 | 303251 | 1.002 | 303346 | 1.002
303337 | 1.001 | 303355 | 1 | 303321 | 1.001 | 303371 | 1.002 | 303466 | 1.002
303457 | 1.001 | 303475 | 1 | 303440 | 1.001 | 303491 | 1.002 | 303586 | 1.002
303577 | 1.001 | 303595 | 1 | 303563 | 1.001 | 303610 | 1.002 | 303706 | 1.002
303696 | 1.001 | 303715 | 1 | 303683 | 1.001 | 303730 | 1.002 | 303825 | 1.002
303819 | 1.001 | 303835 | 1 | 303803 | 1.001 | 303850 | 1.002 | 303948 | 1.002
303939 | 1.001 | 303954 | 1 | 303923 | 1.002 | 303970 | 1.002 | 304068 | 1.002
304059 | 1.001 | 304074 | 1 | 304043 | 1.001 | 304090 | 1.002 | 304188 | 1.002
304179 | 1.001 | 304194 | 1 | 304162 | 1.002 | 304212 | 1.002 | 304308 | 1.002
304299 | 1.001 | 304318 | 1 | 304282 | 1.002 | 304332 | 1.002 | 304429 | 1.002
304418 | 1.001 | 304438 | 1 | 304402 | 1.002 | 304452 | 1.002 | 304548 | 1.002
304538 | 1.001 | 304558 | 1 | 304522 | 1.002 | 304572 | 1.002 | 304668 | 1.002
304658 | 1.001 | 304677 | 1 | 304643 | 1.002 | 304692 | 1.002 | 304788 | 1.002
304778 | 1.001 | 304797 | 1 | 304762 | 1.002 | 304812 | 1.002 | 304908 | 1.002
304898 | 1.001 | 304917 | 1 | 304882 | 1.002 | 304931 | 1.002 | 305028 | 1.002
305017 | 1.001 | 305037 | 1 | 305002 | 1.002 | 305051 | 1.002 | 305147 | 1.002
305137 | 1.001 | 305157 | 1.001 | 305122 | 1.002 | 305171 | 1.002 | 305267 | 1.002
305258 | 1.001 | 305276 | 1 | 305242 | 1.002 | 305291 | 1.002 | 305387 | 1.002
305378 | 1.001 | 305396 | 1 | 305361 | 1.002 | 305411 | 1.002 | 305507 | 1.002
305498 | 1.001 | 305516 | 1 | 305481 | 1.002 | 305530 | 1.002 | 305627 | 1.002
305617 | 1.001 | 305636 | 1 | 305601 | 1.002 | 305650 | 1.002 | 305746 | 1.002
305737 | 1.001 | 305756 | 1.001 | 305721 | 1.002 | 305770 | 1.002 | 305866 | 1.002
305857 | 1.001 | 305875 | 1 | 305841 | 1.002 | 305890 | 1.002 | 305986 | 1.002
305977 | 1.001 | 305995 | 1 | 305960 | 1.002 | 306030 | 1.002 | 306106 | 1.002
306097 | 1.001 | 306115 | 1 | 306080 | 1.002 | 306150 | 1.002 | 306226 | 1.002
306233 | 1.001 | 306235 | 1 | 306200 | 1.002 | 306270 | 1.002 | 306345 | 1.002
306353 | 1.001 | 306355 | 1.001 | 306320 | 1.002 | 306390 | 1.002 | 306465 | 1.002
306472 | 1.001 | 306475 | 1.001 | 306440 | 1.002 | 306509 | 1.002 | 306585 | 1.002
306592 | 1.001 | 306594 | 1.001 | 306560 | 1.002 | 306629 | 1.002 | 306705 | 1.002
306712 | 1.001 | 306714 | 1.001 | 306679 | 1.002 | 306749 | 1.002 | 306825 | 1.002
306832 | 1.001 | 306834 | 1.001 | 306799 | 1.002 | 306869 | 1.002 | 306945 | 1.002
306952 | 1.001 | 306954 | 1.001 | 306919 | 1.002 | 306989 | 1.002 | 307064 | 1.002
307072 | 1.001 | 307074 | 1.001 | 307039 | 1.002 | 307108 | 1.002 | 307184 | 1.002
307191 | 1.001 | 307193 | 1.001 | 307159 | 1.002 | 307228 | 1.002 | 307304 | 1.002
307311 | 1.001 | 307313 | 1.001 | 307278 | 1.002 | 307349 | 1.002 | 307424 | 1.002
307431 | 1.001 | 307433 | 1.001 | 307398 | 1.002 | 307469 | 1.002 | 307544 | 1.002
307551 | 1.001 | 307553 | 1.001 | 307518 | 1.002 | 307589 | 1.002 | 307663 | 1.002
307671 | 1.001 | 307673 | 1 | 307638 | 1.002 | 307708 | 1.002 | 307783 | 1.002
307790 | 1.001 | 307792 | 1.001 | 307758 | 1.002 | 307828 | 1.002 | 307903 | 1.002
307910 | 1.001 | 307912 | 1.001 | 307881 | 1.002 | 307948 | 1.002 | 308023 | 1.002
308030 | 1.001 | 308032 | 1.001 | 308001 | 1.002 | 308068 | 1.002 | 308143 | 1.002
308150 | 1.001 | 308152 | 1.001 | 308121 | 1.002 | 308188 | 1.002 | 308262 | 1.002
308273 | 1.001 | 308272 | 1.001 | 308241 | 1.002 | 308307 | 1.002 | 308386 | 1.002
308392 | 1.001 | 308391 | 1.001 | 308361 | 1.002 | 308427 | 1.002 | 308506 | 1.002
308512 | 1.001 | 308512 | 1.001 | 308481 | 1.002 | 308550 | 1.002 | 308626 | 1.002
308632 | 1.001 | 308635 | 1.001 | 308600 | 1.002 | 308670 | 1.002 | 308746 | 1.002
308752 | 1.001 | 308755 | 1.001 | 308720 | 1.002 | 308790 | 1.002 | 308866 | 1.002
308872 | 1.001 | 308875 | 1.001 | 308840 | 1.002 | 308910 | 1.002 | 308985 | 1.002
308992 | 1.001 | 308995 | 1.001 | 308960 | 1.002 | 309029 | 1.002 | 309105 | 1.002
309111 | 1.001 | 309114 | 1.001 | 309080 | 1.002 | 309149 | 1.002 | 309225 | 1.002
309231 | 1.001 | 309234 | 1.001 | 309199 | 1.002 | 309269 | 1.002 | 309345 | 1.002
309352 | 1.001 | 309354 | 1.001 | 309319 | 1.002 | 309389 | 1.002 | 309465 | 1.002
309472 | 1.001 | 309474 | 1.001 | 309439 | 1.002 | 309509 | 1.002 | 309584 | 1.002
309592 | 1.001 | 309594 | 1.001 | 309559 | 1.002 | 309628 | 1.002 | 309704 | 1.002
309711 | 1.001 | 309713 | 1.001 | 309679 | 1.002 | 309748 | 1.002 | 309824 | 1.002
309831 | 1.001 | 309833 | 1.001 | 309798 | 1.002 | 309868 | 1.002 | 309944 | 1.002
309951 | 1.001 | 309953 | 1.001 | 309918 | 1.002 | 309988 | 1.002 | 310064 | 1.002
310071 | 1.001 | 310073 | 1.001 | 310038 | 1.002 | 310108 | 1.002 | 310183 | 1.002
310191 | 1.001 | 310193 | 1.001 | 310158 | 1.002 | 310227 | 1.002 | 310303 | 1.002
310310 | 1.001 | 310312 | 1.001 | 310278 | 1.002 | 310347 | 1.002 | 310423 | 1.002
310430 | 1.001 | 310432 | 1.001 | 310397 | 1.002 | 310468 | 1.002 | 310543 | 1.002
310550 | 1.001 | 310552 | 1.001 | 310517 | 1.002 | 310588 | 1.002 | 310663 | 1.002
310670 | 1.001 | 310672 | 1.001 | 310637 | 1.002 | 310708 | 1.002 | 310782 | 1.002
310790 | 1.001 | 310792 | 1.001 | 310757 | 1.002 | 310828 | 1.002 | 310902 | 1.002
310909 | 1.001 | 310912 | 1.001 | 310877 | 1.002 | 310947 | 1.002 | 311022 | 1.002
311029 | 1.001 | 311031 | 1.001 | 310996 | 1.002 | 311067 | 1.002 | 311143 | 1.002
311149 | 1.001 | 311151 | 1.001 | 311116 | 1.002 | 311187 | 1.002 | 311263 | 1.002
311269 | 1.001 | 311271 | 1.001 | 311237 | 1.002 | 311307 | 1.002 | 311383 | 1.002
311389 | 1.001 | 311391 | 1.001 | 311375 | 1.002 | 311427 | 1.002 | 311502 | 1.002
311508 | 1.001 | 311511 | 1.001 | 311495 | 1.003 | 311546 | 1.002 | 311622 | 1.002
311628 | 1.001 | 311631 | 1.001 | 311615 | 1.003 | 311666 | 1.002 | 311742 | 1.002
311748 | 1.001 | 311751 | 1.001 | 311735 | 1.002 | 311786 | 1.002 | 311862 | 1.002
311868 | 1.001 | 311871 | 1.001 | 311855 | 1.002 | 311906 | 1.002 | 311982 | 1.002
311988 | 1.001 | 311991 | 1.001 | 311974 | 1.003 | 312026 | 1.002 | 312101 | 1.002
312108 | 1.001 | 312111 | 1.001 | 312094 | 1.003 | 312145 | 1.002 | 312221 | 1.002
312227 | 1.001 | 312230 | 1.001 | 312214 | 1.003 | 312265 | 1.002 | 312341 | 1.002
312347 | 1.001 | 312350 | 1.001 | 312337 | 1.003 | 312385 | 1.002 | 312461 | 1.002
312467 | 1.001 | 312470 | 1.001 | 312457 | 1.002 | 312505 | 1.002 | 312581 | 1.002
312591 | 1.001 | 312590 | 1.001 | 312577 | 1.003 | 312625 | 1.002 | 312725 | 1.002
312711 | 1.001 | 312710 | 1.001 | 312696 | 1.002 | 312744 | 1.002 | 312846 | 1.002
312830 | 1.001 | 312829 | 1.001 | 312816 | 1.003 | 312864 | 1.002 | 312966 | 1.002
312950 | 1.001 | 312949 | 1.001 | 312936 | 1.003 | 312987 | 1.002 | 313085 | 1.002
313070 | 1.001 | 313072 | 1.001 | 313056 | 1.003 | 313107 | 1.002 | 313205 | 1.002
313190 | 1.001 | 313192 | 1.001 | 313176 | 1.003 | 313227 | 1.002 | 313325 | 1.002
313310 | 1.001 | 313312 | 1.001 | 313295 | 1.003 | 313347 | 1.002 | 313445 | 1.002
313430 | 1.001 | 313432 | 1.001 | 313415 | 1.003 | 313466 | 1.002 | 313565 | 1.002
313549 | 1.001 | 313551 | 1.001 | 313536 | 1.003 | 313587 | 1.002 | 313684 | 1.002
313669 | 1.001 | 313671 | 1.001 | 313656 | 1.003 | 313707 | 1.002 | 313804 | 1.002
313789 | 1.001 | 313791 | 1.001 | 313776 | 1.003 | 313827 | 1.002 | 313924 | 1.002
313909 | 1.001 | 313911 | 1.001 | 313895 | 1.003 | 313947 | 1.002 | 314044 | 1.002
314029 | 1.001 | 314031 | 1.001 | 314015 | 1.003 | 314066 | 1.002 | 314164 | 1.002
314148 | 1.001 | 314150 | 1.001 | 314135 | 1.003 | 314186 | 1.002 | 314284 | 1.002
314268 | 1.001 | 314270 | 1.001 | 314255 | 1.003 | 314306 | 1.002 | 314403 | 1.002
314388 | 1.001 | 314390 | 1.001 | 314375 | 1.003 | 314426 | 1.002 | 314523 | 1.002
314508 | 1.001 | 314510 | 1.001 | 314494 | 1.003 | 314546 | 1.002 | 314643 | 1.002
314628 | 1.001 | 314630 | 1.001 | 314614 | 1.003 | 314665 | 1.002 | 314763 | 1.002
314747 | 1.001 | 314749 | 1.001 | 314734 | 1.003 | 314785 | 1.002 | 314883 | 1.002
314867 | 1.001 | 314870 | 1.001 | 314854 | 1.003 | 314905 | 1.002 | 315002 | 1.002
314987 | 1.001 | 314990 | 1.001 | 314974 | 1.003 | 315025 | 1.002 | 315122 | 1.002
315107 | 1.001 | 315110 | 1.001 | 315094 | 1.003 | 315145 | 1.002 | 315242 | 1.002
315227 | 1.001 | 315230 | 1.001 | 315213 | 1.003 | 315265 | 1.002 | 315362 | 1.002
315346 | 1.001 | 315350 | 1.001 | 315333 | 1.003 | 315384 | 1.002 | 315482 | 1.002
315466 | 1.001 | 315469 | 1.001 | 315453 | 1.003 | 315504 | 1.002 | 315601 | 1.002
315586 | 1.001 | 315589 | 1.001 | 315573 | 1.003 | 315624 | 1.002 | 315721 | 1.002
315706 | 1.001 | 315709 | 1.001 | 315693 | 1.003 | 315744 | 1.002 | 315841 | 1.002
315826 | 1.001 | 315829 | 1.001 | 315812 | 1.003 | 315864 | 1.002 | 315962 | 1.002
315947 | 1.001 | 315966 | 1.001 | 315932 | 1.003 | 315983 | 1.002 | 316082 | 1.002
316066 | 1.001 | 316086 | 1.001 | 316052 | 1.003 | 316103 | 1.003 | 316201 | 1.002
316186 | 1.001 | 316206 | 1.001 | 316172 | 1.003 | 316223 | 1.003 | 316321 | 1.002
316306 | 1.001 | 316325 | 1.001 | 316292 | 1.003 | 316343 | 1.002 | 316441 | 1.002
316426 | 1.001 | 316445 | 1.001 | 316412 | 1.003 | 316463 | 1.002 | 316561 | 1.002
316546 | 1.001 | 316565 | 1.001 | 316532 | 1.003 | 316582 | 1.002 | 316691 | 1.002
316676 | 1.001 | 316685 | 1.001 | 316662 | 1.003 | 316702 | 1.002 | 316811 | 1.002
316795 | 1.001 | 316805 | 1.001 | 316780 | 1.003 | 316823 | 1.002 | 316931 | 1.002
316915 | 1.001 | 316924 | 1.001 | 316900 | 1.003 | 316953 | 1.002 | 317050 | 1.002
317039 | 1.001 | 317044 | 1.001 | 317020 | 1.003 | 317073 | 1.003 | 317174 | 1.002
317159 | 1.001 | 317175 | 1.001 | 317139 | 1.003 | 317193 | 1.003 | 317294 | 1.002
317279 | 1.001 | 317295 | 1.001 | 317259 | 1.003 | 317316 | 1.002 | 317414 | 1.002
317399 | 1.001 | 317418 | 1.001 | 317379 | 1.003 | 317435 | 1.002 | 317534 | 1.002
317518 | 1.001 | 317538 | 1.001 | 317499 | 1.003 | 317555 | 1.002 | 317654 | 1.002
317638 | 1.001 | 317658 | 1.001 | 317619 | 1.003 | 317675 | 1.002 | 317773 | 1.002
317758 | 1.001 | 317777 | 1.001 | 317739 | 1.003 | 317795 | 1.002 | 317893 | 1.002
317878 | 1.001 | 317897 | 1.001 | 317858 | 1.003 | 317915 | 1.002 | 318013 | 1.002
317998 | 1.001 | 318017 | 1.001 | 317978 | 1.003 | 318034 | 1.002 | 318133 | 1.002
318117 | 1.001 | 318137 | 1.001 | 318098 | 1.003 | 318154 | 1.002 | 318253 | 1.002
318237 | 1.001 | 318257 | 1.001 | 318218 | 1.003 | 318274 | 1.002 | 318372 | 1.002
318357 | 1.001 | 318376 | 1.001 | 318338 | 1.003 | 318395 | 1.002 | 318492 | 1.002
318477 | 1.001 | 318496 | 1.001 | 318457 | 1.003 | 318515 | 1.002 | 318612 | 1.002
318597 | 1.001 | 318616 | 1.001 | 318577 | 1.003 | 318635 | 1.002 | 318732 | 1.002
318716 | 1.001 | 318736 | 1.001 | 318697 | 1.003 | 318754 | 1.002 | 318852 | 1.002
318836 | 1.001 | 318856 | 1.001 | 318817 | 1.003 | 318894 | 1.002 | 318971 | 1.002
318956 | 1.001 | 318976 | 1.001 | 318937 | 1.003 | 319013 | 1.002 | 319091 | 1.002
319076 | 1.001 | 319095 | 1.001 | 319056 | 1.003 | 319133 | 1.002 | 319211 | 1.002
319196 | 1.001 | 319215 | 1.001 | 319176 | 1.003 | 319253 | 1.002 | 319331 | 1.002
319315 | 1.001 | 319335 | 1.001 | 319296 | 1.003 | 319373 | 1.002 | 319451 | 1.002
319435 | 1.001 | 319455 | 1.001 | 319416 | 1.003 | 319493 | 1.002 | 319570 | 1.002
319555 | 1.001 | 319575 | 1.001 | 319537 | 1.003 | 319612 | 1.002 | 319690 | 1.002
319691 | 1.001 | 319694 | 1.001 | 319656 | 1.003 | 319732 | 1.002 | 319810 | 1.002
319811 | 1.001 | 319814 | 1.001 | 319776 | 1.003 | 319853 | 1.002 | 319930 | 1.002
319931 | 1.001 | 319934 | 1.001 | 319896 | 1.003 | 319973 | 1.002 | 320050 | 1.002
320051 | 1.001 | 320055 | 1.001 | 320016 | 1.003 | 320093 | 1.002 | 320169 | 1.002
320171 | 1.001 | 320175 | 1.001 | 320136 | 1.003 | 320212 | 1.002 | 320290 | 1.002
320290 | 1.001 | 320294 | 1.001 | 320256 | 1.003 | 320332 | 1.002 | 320410 | 1.002
320410 | 1.001 | 320414 | 1.001 | 320375 | 1.003 | 320452 | 1.002 | 320530 | 1.002
320530 | 1.001 | 320534 | 1.001 | 320495 | 1.003 | 320572 | 1.002 | 320650 | 1.002
320650 | 1.001 | 320654 | 1.001 | 320615 | 1.003 | 320692 | 1.002 | 320770 | 1.002
320770 | 1.001 | 320774 | 1.001 | 320735 | 1.003 | 320812 | 1.002 | 320889 | 1.002
320889 | 1.001 | 320893 | 1.001 | 320855 | 1.003 | 320931 | 1.002 | 321009 | 1.002
321009 | 1.001 | 321013 | 1.001 | 320974 | 1.003 | 321051 | 1.002 | 321129 | 1.002
321130 | 1.001 | 321133 | 1.001 | 321097 | 1.003 | 321171 | 1.002 | 321249 | 1.002
321250 | 1.001 | 321253 | 1.001 | 321217 | 1.003 | 321291 | 1.002 | 321369 | 1.002
321370 | 1.001 | 321373 | 1.001 | 321337 | 1.003 | 321411 | 1.002 | 321488 | 1.002
321492 | 1.001 | 321492 | 1.001 | 321457 | 1.003 | 321530 | 1.002 | 321611 | 1.002
321612 | 1.001 | 321612 | 1.001 | 321576 | 1.003 | 321650 | 1.003 | 321731 | 1.002
321732 | 1.001 | 321732 | 1.001 | 321696 | 1.003 | 321773 | 1.003 | 321851 | 1.002
321852 | 1.001 | 321855 | 1.001 | 321817 | 1.003 | 321893 | 1.002 | 321971 | 1.002
321972 | 1.001 | 321975 | 1.001 | 321937 | 1.003 | 322013 | 1.002 | 322091 | 1.002
322092 | 1.001 | 322095 | 1.001 | 322057 | 1.003 | 322132 | 1.002 | 322210 | 1.002
322211 | 1.001 | 322214 | 1.001 | 322177 | 1.003 | 322252 | 1.002 | 322330 | 1.002
322331 | 1.001 | 322334 | 1.001 | 322296 | 1.003 | 322373 | 1.002 | 322451 | 1.002
322451 | 1.001 | 322454 | 1.001 | 322416 | 1.003 | 322493 | 1.002 | 322571 | 1.002
322571 | 1.001 | 322575 | 1.001 | 322536 | 1.003 | 322613 | 1.002 | 322691 | 1.002
322691 | 1.001 | 322695 | 1.001 | 322656 | 1.003 | 322733 | 1.002 | 322810 | 1.002
322810 | 1.001 | 322814 | 1.001 | 322776 | 1.003 | 322852 | 1.002 | 322930 | 1.002
322930 | 1.001 | 322934 | 1.001 | 322895 | 1.003 | 322972 | 1.002 | 323050 | 1.001
323051 | 1.001 | 323054 | 1.001 | 323015 | 1.003 | 323092 | 1.002 | 323170 | 1.002
323171 | 1.001 | 323174 | 1.001 | 323135 | 1.003 | 323212 | 1.002 | 323290 | 1.002
323291 | 1.001 | 323294 | 1.001 | 323255 | 1.003 | 323332 | 1.002 | 323409 | 1.002
323410 | 1.001 | 323414 | 1.001 | 323375 | 1.003 | 323451 | 1.002 | 323529 | 1.002
323530 | 1.001 | 323533 | 1.001 | 323494 | 1.003 | 323571 | 1.002 | 323649 | 1.002
323650 | 1.001 | 323653 | 1.001 | 323614 | 1.003 | 323691 | 1.002 | 323769 | 1.002
323770 | 1.001 | 323773 | 1.001 | 323734 | 1.003 | 323811 | 1.002 | 323889 | 1.002
323890 | 1.001 | 323893 | 1.001 | 323855 | 1.003 | 323931 | 1.003 | 324008 | 1.001
324009 | 1.001 | 324013 | 1.001 | 323975 | 1.003 | 324050 | 1.003 | 324128 | 1.001
324129 | 1.001 | 324132 | 1.001 | 324094 | 1.003 | 324170 | 1.002 | 324248 | 1.002
324249 | 1.001 | 324252 | 1.001 | 324214 | 1.003 | 324290 | 1.002 | 324368 | 1.002
324369 | 1.001 | 324372 | 1.001 | 324334 | 1.003 | 324410 | 1.002 | 324488 | 1.002
324489 | 1.001 | 324492 | 1.001 | 324454 | 1.003 | 324531 | 1.003 | 324609 | 1.002
324609 | 1.001 | 324612 | 1.001 | 324592 | 1.003 | 324651 | 1.002 | 324728 | 1.001
324728 | 1.001 | 324732 | 1.001 | 324712 | 1.003 | 324770 | 1.002 | 324848 | 1.001
324848 | 1.001 | 324852 | 1.001 | 324832 | 1.003 | 324890 | 1.002 | 324968 | 1.001
324968 | 1.001 | 324972 | 1.001 | 324952 | 1.003 | 325010 | 1.002 | 325088 | 1.001
325088 | 1.001 | 325092 | 1.001 | 325071 | 1.003 | 325130 | 1.003 | 325208 | 1.002
325208 | 1.001 | 325212 | 1.001 | 325191 | 1.003 | 325250 | 1.002 | 325327 | 1.002
325328 | 1.001 | 325331 | 1.001 | 325311 | 1.003 | 325369 | 1.002 | 325469 | 1.002
325448 | 1.001 | 325451 | 1.001 | 325431 | 1.003 | 325489 | 1.002 | 325588 | 1.001
325568 | 1.001 | 325571 | 1.001 | 325555 | 1.003 | 325609 | 1.002 | 325708 | 1.001
325688 | 1.001 | 325691 | 1.001 | 325675 | 1.003 | 325729 | 1.002 | 325828 | 1.001
325811 | 1.001 | 325811 | 1.001 | 325794 | 1.003 | 325849 | 1.002 | 325952 | 1.001
325931 | 1.001 | 325931 | 1.001 | 325914 | 1.003 | 325968 | 1.002 | 326072 | 1.001
326050 | 1.001 | 326050 | 1.001 | 326034 | 1.003 | 326088 | 1.002 | 326192 | 1.001
326170 | 1.001 | 326170 | 1.001 | 326154 | 1.003 | 326212 | 1.002 | 326311 | 1.002
326290 | 1.001 | 326293 | 1.001 | 326274 | 1.003 | 326332 | 1.002 | 326431 | 1.002
326410 | 1.001 | 326414 | 1.001 | 326393 | 1.003 | 326452 | 1.002 | 326551 | 1.002
326530 | 1.001 | 326534 | 1.001 | 326513 | 1.003 | 326572 | 1.002 | 326671 | 1.002
326649 | 1.001 | 326653 | 1.001 | 326633 | 1.003 | 326691 | 1.002 | 326791 | 1.002
326769 | 1.001 | 326773 | 1.001 | 326753 | 1.003 | 326811 | 1.002 | 326910 | 1.002
326889 | 1.001 | 326893 | 1.001 | 326873 | 1.003 | 326931 | 1.002 | 327030 | 1.002
327010 | 1.001 | 327013 | 1.001 | 326993 | 1.003 | 327051 | 1.002 | 327150 | 1.002
327130 | 1.001 | 327133 | 1.001 | 327113 | 1.003 | 327171 | 1.002 | 327270 | 1.002
327249 | 1.001 | 327252 | 1.001 | 327233 | 1.003 | 327290 | 1.003 | 327390 | 1.002
327369 | 1.002 | 327372 | 1.001 | 327353 | 1.003 | 327410 | 1.002 | 327510 | 1.002
327489 | 1.001 | 327492 | 1.001 | 327473 | 1.003 | 327530 | 1.002 | 327630 | 1.002
327609 | 1.001 | 327612 | 1.001 | 327592 | 1.003 | 327650 | 1.002 | 327750 | 1.002
327729 | 1.001 | 327732 | 1.001 | 327712 | 1.003 | 327770 | 1.002 | 327870 | 1.002
327848 | 1.001 | 327852 | 1.001 | 327832 | 1.003 | 327889 | 1.002 | 327990 | 1.001
327968 | 1.002 | 327971 | 1.001 | 327952 | 1.003 | 328009 | 1.002 | 328110 | 1.001
328088 | 1.002 | 328091 | 1.001 | 328072 | 1.003 | 328129 | 1.002 | 328229 | 1.002
328208 | 1.001 | 328211 | 1.001 | 328192 | 1.003 | 328250 | 1.002 | 328349 | 1.002
328328 | 1.001 | 328331 | 1.001 | 328311 | 1.003 | 328370 | 1.002 | 328469 | 1.002
328448 | 1.001 | 328451 | 1.001 | 328431 | 1.003 | 328489 | 1.002 | 328589 | 1.002
328567 | 1.002 | 328570 | 1.001 | 328551 | 1.003 | 328609 | 1.002 | 328709 | 1.002
328687 | 1.002 | 328691 | 1.001 | 328671 | 1.003 | 328729 | 1.002 | 328828 | 1.002
328807 | 1.002 | 328811 | 1.001 | 328791 | 1.003 | 328849 | 1.002 | 328948 | 1.002
328927 | 1.002 | 328931 | 1.001 | 328910 | 1.003 | 328969 | 1.002 | 329068 | 1.002
329047 | 1.002 | 329051 | 1.001 | 329030 | 1.003 | 329089 | 1.002 | 329188 | 1.002
329167 | 1.002 | 329170 | 1.001 | 329151 | 1.003 | 329208 | 1.002 | 329308 | 1.002
329287 | 1.002 | 329308 | 1.001 | 329271 | 1.003 | 329328 | 1.002 | 329427 | 1.002
329407 | 1.002 | 329427 | 1.001 | 329391 | 1.003 | 329448 | 1.002 | 329547 | 1.002
329527 | 1.002 | 329547 | 1.001 | 329510 | 1.003 | 329568 | 1.002 | 329667 | 1.002
329647 | 1.002 | 329667 | 1.001 | 329630 | 1.003 | 329688 | 1.002 | 329788 | 1.002
329766 | 1.002 | 329787 | 1.001 | 329750 | 1.003 | 329807 | 1.002 | 329908 | 1.002
329886 | 1.002 | 329907 | 1.001 | 329873 | 1.003 | 329927 | 1.002 | 330028 | 1.001
330006 | 1.002 | 330027 | 1.001 | 329993 | 1.003 | 330047 | 1.002 | 330147 | 1.002
330126 | 1.002 | 330146 | 1.001 | 330113 | 1.003 | 330167 | 1.002 | 330267 | 1.002
330249 | 1.002 | 330266 | 1.001 | 330232 | 1.003 | 330287 | 1.002 | 330390 | 1.002
330369 | 1.002 | 330386 | 1.001 | 330352 | 1.003 | 330406 | 1.002 | 330510 | 1.002
330488 | 1.002 | 330506 | 1.001 | 330472 | 1.003 | 330530 | 1.002 | 330630 | 1.002
330608 | 1.002 | 330630 | 1.001 | 330592 | 1.003 | 330650 | 1.002 | 330749 | 1.002
330729 | 1.002 | 330749 | 1.001 | 330713 | 1.003 | 330770 | 1.002 | 330869 | 1.002
330849 | 1.002 | 330869 | 1.001 | 330832 | 1.003 | 330890 | 1.002 | 330989 | 1.002
330969 | 1.002 | 330989 | 1.001 | 330952 | 1.003 | 331010 | 1.002 | 331109 | 1.002
331088 | 1.002 | 331109 | 1.001 | 331072 | 1.003 | 331129 | 1.002 | 331229 | 1.002
331208 | 1.002 | 331229 | 1.001 | 331192 | 1.003 | 331249 | 1.002 | 331350 | 1.002
331328 | 1.002 | 331348 | 1.001 | 331312 | 1.003 | 331369 | 1.002 | 331469 | 1.002
331448 | 1.002 | 331468 | 1.001 | 331431 | 1.003 | 331489 | 1.002 | 331589 | 1.002
331568 | 1.002 | 331588 | 1.001 | 331551 | 1.003 | 331609 | 1.002 | 331709 | 1.002
331687 | 1.002 | 331708 | 1.001 | 331671 | 1.003 | 331728 | 1.002 | 331829 | 1.002
331807 | 1.002 | 331828 | 1.001 | 331791 | 1.003 | 331869 | 1.002 | 331949 | 1.002
331927 | 1.002 | 331948 | 1.001 | 331911 | 1.003 | 331988 | 1.002 | 332068 | 1.002
332047 | 1.002 | 332067 | 1 | 332030 | 1.003 | 332108 | 1.002 | 332188 | 1.002
332167 | 1.002 | 332187 | 1 | 332150 | 1.003 | 332228 | 1.002 | 332308 | 1.002
332286 | 1.002 | 332307 | 1.001 | 332270 | 1.003 | 332348 | 1.002 | 332428 | 1.002
332406 | 1.002 | 332427 | 1.001 | 332390 | 1.003 | 332468 | 1.002 | 332548 | 1.002
332526 | 1.002 | 332547 | 1.001 | 332510 | 1.003 | 332588 | 1.002 | 332667 | 1.002
332646 | 1.002 | 332666 | 1.001 | 332630 | 1.003 | 332708 | 1.002 | 332787 | 1.002
332766 | 1.002 | 332786 | 1.001 | 332750 | 1.003 | 332828 | 1.003 | 332907 | 1.002
332886 | 1.002 | 332907 | 1.001 | 332870 | 1.003 | 332948 | 1.002 | 333027 | 1.002
333006 | 1.002 | 333027 | 1.001 | 332990 | 1.003 | 333068 | 1.002 | 333147 | 1.002
333126 | 1.002 | 333147 | 1.001 | 333110 | 1.003 | 333188 | 1.002 | 333266 | 1.002
333261 | 1.002 | 333266 | 1 | 333230 | 1.003 | 333307 | 1.002 | 333386 | 1.001
333381 | 1.002 | 333386 | 1 | 333349 | 1.003 | 333427 | 1.002 | 333507 | 1.002
333501 | 1.002 | 333506 | 1.001 | 333469 | 1.003 | 333547 | 1.002 | 333627 | 1.002
333622 | 1.002 | 333626 | 1 | 333589 | 1.003 | 333667 | 1.002 | 333747 | 1.002
333742 | 1.002 | 333746 | 1 | 333709 | 1.003 | 333787 | 1.002 | 333867 | 1.001
333861 | 1.002 | 333865 | 1 | 333829 | 1.003 | 333906 | 1.002 | 333986 | 1.001
333981 | 1.002 | 333985 | 1 | 333948 | 1.003 | 334026 | 1.002 | 334106 | 1.002
334101 | 1.002 | 334105 | 1 | 334068 | 1.003 | 334146 | 1.002 | 334226 | 1.001
334221 | 1.002 | 334225 | 1 | 334188 | 1.003 | 334266 | 1.002 | 334346 | 1.002
334341 | 1.002 | 334345 | 1 | 334312 | 1.003 | 334386 | 1.002 | 334466 | 1.002
334460 | 1.002 | 334465 | 1 | 334432 | 1.003 | 334505 | 1.002 | 334585 | 1.001
334583 | 1.002 | 334584 | 1 | 334552 | 1.003 | 334625 | 1.002 | 334708 | 1.002
334703 | 1.002 | 334704 | 1 | 334671 | 1.003 | 334745 | 1.002 | 334828 | 1.001
334823 | 1.002 | 334824 | 1 | 334791 | 1.003 | 334865 | 1.002 | 334948 | 1.001
334943 | 1.002 | 334944 | 1 | 334911 | 1.003 | 334989 | 1.002 | 335068 | 1.002
335063 | 1.002 | 335068 | 1 | 335031 | 1.003 | 335109 | 1.002 | 335188 | 1.002
335183 | 1.002 | 335187 | 1 | 335151 | 1.003 | 335228 | 1.002 | 335308 | 1.002
335303 | 1.002 | 335307 | 1 | 335270 | 1.003 | 335348 | 1.002 | 335428 | 1.002
335423 | 1.002 | 335427 | 1.001 | 335390 | 1.003 | 335468 | 1.002 | 335548 | 1.001
335543 | 1.002 | 335547 | 1 | 335510 | 1.003 | 335588 | 1.002 | 335668 | 1.001
335663 | 1.002 | 335667 | 1 | 335630 | 1.003 | 335708 | 1.002 | 335788 | 1.001
335782 | 1.002 | 335787 | 1 | 335750 | 1.003 | 335827 | 1.002 | 335907 | 1.002
335902 | 1.002 | 335906 | 1 | 335869 | 1.003 | 335947 | 1.002 | 336027 | 1.001
336022 | 1.002 | 336026 | 1 | 335989 | 1.003 | 336067 | 1.002 | 336147 | 1.001
336142 | 1.002 | 336146 | 1 | 336109 | 1.002 | 336187 | 1.002 | 336267 | 1.001
336262 | 1.002 | 336266 | 1 | 336229 | 1.003 | 336307 | 1.002 | 336387 | 1.001
336381 | 1.002 | 336386 | 1 | 336350 | 1.003 | 336427 | 1.002 | 336506 | 1.001
336501 | 1.002 | 336505 | 1 | 336470 | 1.003 | 336547 | 1.002 | 336626 | 1.001
336621 | 1.002 | 336625 | 1 | 336589 | 1.003 | 336667 | 1.002 | 336746 | 1.001
336741 | 1.002 | 336745 | 1 | 336709 | 1.003 | 336787 | 1.002 | 336866 | 1.002
336861 | 1.002 | 336866 | 1 | 336829 | 1.003 | 336907 | 1.002 | 336986 | 1.001
336980 | 1.002 | 336986 | 1 | 336949 | 1.002 | 337027 | 1.002 | 337105 | 1.001
337100 | 1.002 | 337105 | 1 | 337069 | 1.003 | 337146 | 1.002 | 337226 | 1.001
337220 | 1.002 | 337225 | 1 | 337188 | 1.003 | 337266 | 1.002 | 337346 | 1.001
337341 | 1.002 | 337345 | 1 | 337308 | 1.003 | 337386 | 1.002 | 337466 | 1.001
337461 | 1.002 | 337465 | 1 | 337428 | 1.002 | 337506 | 1.002 | 337586 | 1.001
337581 | 1.002 | 337585 | 1 | 337548 | 1.003 | 337626 | 1.002 | 337705 | 1.001
337700 | 1.003 | 337704 | 1 | 337686 | 1.002 | 337745 | 1.002 | 337825 | 1.001
337820 | 1.003 | 337824 | 1 | 337806 | 1.003 | 337865 | 1.002 | 337945 | 1.001
337940 | 1.003 | 337944 | 1 | 337926 | 1.003 | 337985 | 1.002 | 338065 | 1.001
338060 | 1.002 | 338064 | 1 | 338045 | 1.003 | 338105 | 1.002 | 338206 | 1.001
338180 | 1.003 | 338184 | 1 | 338166 | 1.002 | 338225 | 1.002 | 338326 | 1.001
338299 | 1.003 | 338304 | 1 | 338286 | 1.002 | 338344 | 1.002 | 338446 | 1.001
338419 | 1.003 | 338423 | 1 | 338406 | 1.003 | 338464 | 1.002 | 338566 | 1.001
338539 | 1.003 | 338543 | 1 | 338526 | 1.003 | 338584 | 1.002 | 338685 | 1.001
338659 | 1.003 | 338663 | 1 | 338649 | 1.002 | 338704 | 1.002 | 338805 | 1.001
338779 | 1.003 | 338783 | 1 | 338768 | 1.003 | 338825 | 1.002 | 338926 | 1.001
338898 | 1.003 | 338903 | 1 | 338888 | 1.003 | 338945 | 1.002 | 339046 | 1.001
339022 | 1.003 | 339022 | 1 | 339008 | 1.003 | 339064 | 1.002 | 339169 | 1.001
339142 | 1.003 | 339143 | 1 | 339128 | 1.003 | 339184 | 1.002 | 339289 | 1.001
339262 | 1.003 | 339263 | 1 | 339248 | 1.003 | 339307 | 1.002 | 339408 | 1.001
339382 | 1.003 | 339386 | 1 | 339367 | 1.003 | 339427 | 1.002 | 339528 | 1.001
339502 | 1.003 | 339506 | 1 | 339487 | 1.003 | 339547 | 1.002 | 339648 | 1.001
339621 | 1.003 | 339625 | 1 | 339607 | 1.003 | 339666 | 1.002 | 339768 | 1.001
339741 | 1.003 | 339745 | 1 | 339728 | 1.003 | 339786 | 1.002 | 339888 | 1.001
339861 | 1.003 | 339865 | 1 | 339848 | 1.003 | 339906 | 1.002 | 340007 | 1.001
339981 | 1.003 | 339985 | 1 | 339968 | 1.003 | 340026 | 1.002 | 340127 | 1.001
340101 | 1.003 | 340105 | 1 | 340087 | 1.003 | 340146 | 1.002 | 340247 | 1.001
340220 | 1.003 | 340225 | 1 | 340207 | 1.003 | 340265 | 1.002 | 340367 | 1.001
340340 | 1.003 | 340344 | 1 | 340327 | 1.003 | 340386 | 1.002 | 340488 | 1.001
340460 | 1.003 | 340464 | 1 | 340447 | 1.003 | 340506 | 1.002 | 340608 | 1.001
340580 | 1.003 | 340584 | 1 | 340567 | 1.003 | 340626 | 1.002 | 340727 | 1.001
340700 | 1.003 | 340704 | 1 | 340686 | 1.003 | 340746 | 1.002 | 340847 | 1.001
340819 | 1.003 | 340824 | 1 | 340806 | 1.003 | 340866 | 1.002 | 340967 | 1.001
340940 | 1.003 | 340944 | 1 | 340926 | 1.003 | 340985 | 1.002 | 341087 | 1.002
341060 | 1.003 | 341064 | 1 | 341046 | 1.003 | 341105 | 1.002 | 341207 | 1.001
341180 | 1.002 | 341184 | 1 | 341166 | 1.003 | 341225 | 1.002 | 341326 | 1.001
341300 | 1.003 | 341304 | 1 | 341285 | 1.003 | 341345 | 1.002 | 341446 | 1.001
341420 | 1.003 | 341424 | 1 | 341405 | 1.003 | 341465 | 1.002 | 341566 | 1.001
341539 | 1.003 | 341543 | 1 | 341525 | 1.003 | 341584 | 1.002 | 341686 | 1.001
341659 | 1.003 | 341663 | 1 | 341645 | 1.003 | 341704 | 1.002 | 341806 | 1.001
341779 | 1.003 | 341783 | 1 | 341766 | 1.003 | 341824 | 1.002 | 341925 | 1.001
341899 | 1.002 | 341903 | 1 | 341885 | 1.003 | 341944 | 1.002 | 342045 | 1.001
342019 | 1.002 | 342023 | 1 | 342005 | 1.003 | 342064 | 1.002 | 342165 | 1.001
342138 | 1.003 | 342142 | 1 | 342125 | 1.003 | 342183 | 1.002 | 342285 | 1.001
342258 | 1.003 | 342262 | 1 | 342245 | 1.003 | 342303 | 1.002 | 342405 | 1.001
342378 | 1.003 | 342382 | 1.001 | 342365 | 1.003 | 342423 | 1.002 | 342524 | 1.001
342498 | 1.003 | 342502 | 1 | 342484 | 1.003 | 342544 | 1.002 | 342645 | 1.001
342618 | 1.003 | 342639 | 1.001 | 342604 | 1.003 | 342664 | 1.002 | 342765 | 1.001
342737 | 1.003 | 342759 | 1 | 342724 | 1.003 | 342784 | 1.002 | 342885 | 1.001
342857 | 1.002 | 342879 | 1 | 342844 | 1.003 | 342903 | 1.002 | 343005 | 1.001
342977 | 1.002 | 342999 | 1 | 342964 | 1.003 | 343023 | 1.002 | 343124 | 1.001
343098 | 1.003 | 343118 | 1 | 343087 | 1.003 | 343143 | 1.002 | 343244 | 1.001
343218 | 1.002 | 343238 | 1 | 343206 | 1.003 | 343263 | 1.002 | 343364 | 1.001
343341 | 1.003 | 343358 | 1 | 343327 | 1.003 | 343383 | 1.002 | 343487 | 1.001
343460 | 1.002 | 343478 | 1 | 343447 | 1.003 | 343502 | 1.002 | 343607 | 1.001
343580 | 1.002 | 343599 | 1 | 343567 | 1.003 | 343622 | 1.002 | 343727 | 1.001
343700 | 1.002 | 343718 | 1 | 343687 | 1.003 | 343745 | 1.002 | 343846 | 1.002
343820 | 1.002 | 343841 | 1 | 343806 | 1.003 | 343865 | 1.002 | 343966 | 1.001
343940 | 1.002 | 343961 | 1 | 343926 | 1.003 | 343985 | 1.002 | 344086 | 1.001
344059 | 1.002 | 344081 | 1 | 344046 | 1.003 | 344104 | 1.002 | 344206 | 1.001
344179 | 1.003 | 344201 | 1 | 344166 | 1.003 | 344225 | 1.002 | 344327 | 1.001
344299 | 1.002 | 344321 | 1 | 344286 | 1.003 | 344345 | 1.002 | 344446 | 1.001
344419 | 1.002 | 344440 | 1 | 344406 | 1.003 | 344465 | 1.002 | 344566 | 1.001
344539 | 1.002 | 344560 | 1 | 344525 | 1.003 | 344585 | 1.002 | 344686 | 1.001
344658 | 1.002 | 344680 | 1 | 344645 | 1.003 | 344724 | 1.002 | 344806 | 1.001
344779 | 1.002 | 344800 | 1 | 344765 | 1.003 | 344845 | 1.002 | 344926 | 1.001
344899 | 1.002 | 344920 | 1 | 344885 | 1.003 | 344965 | 1.002 | 345046 | 1.001
345019 | 1.002 | 345039 | 1.001 | 345005 | 1.003 | 345084 | 1.002 | 345165 | 1.001
345139 | 1.002 | 345159 | 1.001 | 345124 | 1.003 | 345204 | 1.002 | 345285 | 1.001
345259 | 1.002 | 345280 | 1.001 | 345244 | 1.003 | 345324 | 1.002 | 345405 | 1.001
345378 | 1.002 | 345400 | 1 | 345365 | 1.003 | 345444 | 1.002 | 345525 | 1.001
345498 | 1.002 | 345520 | 1 | 345485 | 1.003 | 345564 | 1.002 | 345645 | 1.002
345618 | 1.002 | 345639 | 1.001 | 345605 | 1.003 | 345683 | 1.002 | 345764 | 1.001
345738 | 1.002 | 345759 | 1.001 | 345724 | 1.003 | 345803 | 1.002 | 345884 | 1.001
345858 | 1.002 | 345879 | 1.001 | 345844 | 1.003 | 345923 | 1.002 | 346005 | 1.001
345977 | 1.002 | 345999 | 1 | 345964 | 1.003 | 346043 | 1.002 | 346135 | 1.001
346108 | 1.002 | 346119 | 1.001 | 346084 | 1.003 | 346163 | 1.002 | 346255 | 1.001
346228 | 1.002 | 346238 | 1.001 | 346215 | 1.003 | 346283 | 1.002 | 346375 | 1.001
346348 | 1.002 | 346358 | 1.001 | 346335 | 1.003 | 346414 | 1.002 | 346494 | 1.001
346468 | 1.002 | 346478 | 1.001 | 346455 | 1.003 | 346533 | 1.002 | 346614 | 1.001
346588 | 1.002 | 346609 | 1.001 | 346574 | 1.003 | 346653 | 1.002 | 346734 | 1.001
346724 | 1.002 | 346729 | 1.001 | 346694 | 1.003 | 346773 | 1.002 | 346854 | 1.001
346844 | 1.002 | 346849 | 1.001 | 346814 | 1.003 | 346893 | 1.002 | 346974 | 1.001
346963 | 1.002 | 346969 | 1.001 | 346934 | 1.003 | 347013 | 1.002 | 347094 | 1.001
347083 | 1.002 | 347088 | 1.001 | 347054 | 1.003 | 347132 | 1.002 | 347213 | 1.001
347203 | 1.002 | 347208 | 1.001 | 347173 | 1.003 | 347252 | 1.002 | 347333 | 1.001
347323 | 1.002 | 347328 | 1.001 | 347293 | 1.003 | 347372 | 1.002 | 347454 | 1.001
347443 | 1.002 | 347448 | 1.001 | 347413 | 1.003 | 347492 | 1.002 | 347574 | 1.002
347563 | 1.002 | 347568 | 1.001 | 347531 | 1.003 | 347612 | 1.002 | 347694 | 1.002
347682 | 1.002 | 347687 | 1.001 | 347651 | 1.003 | 347732 | 1.002 | 347813 | 1.002
347806 | 1.002 | 347807 | 1.001 | 347770 | 1.003 | 347852 | 1.002 | 347936 | 1.001
347926 | 1.002 | 347927 | 1.001 | 347890 | 1.003 | 347972 | 1.002 | 348056 | 1.001
348046 | 1.002 | 348047 | 1 | 348010 | 1.003 | 348095 | 1.002 | 348176 | 1.001
348166 | 1.002 | 348171 | 1.001 | 348130 | 1.003 | 348215 | 1.002 | 348296 | 1.002
348285 | 1.002 | 348291 | 1.001 | 348250 | 1.003 | 348335 | 1.002 | 348416 | 1.001
348405 | 1.002 | 348410 | 1.001 | 348369 | 1.003 | 348454 | 1.002 | 348535 | 1.001
348525 | 1.002 | 348530 | 1.001 | 348489 | 1.003 | 348574 | 1.002 | 348655 | 1.002
348645 | 1.002 | 348650 | 1.001 | 348609 | 1.003 | 348694 | 1.002 | 348775 | 1.001
348765 | 1.002 | 348770 | 1.001 | 348729 | 1.003 | 348814 | 1.002 | 348895 | 1.001
348884 | 1.002 | 348890 | 1.001 | 348850 | 1.003 | 348934 | 1.002 | 349016 | 1.002
349004 | 1.003 | 349009 | 1.001 | 348969 | 1.003 | 349053 | 1.002 | 349135 | 1.002
349124 | 1.002 | 349129 | 1.001 | 349089 | 1.003 | 349173 | 1.002 | 349255 | 1.002
349244 | 1.002 | 349249 | 1.001 | 349209 | 1.003 | 349294 | 1.002 | 349375 | 1.001
349364 | 1.002 | 349369 | 1.001 | 349329 | 1.003 | 349414 | 1.002 | 349495 | 1.002
349484 | 1.002 | 349489 | 1.001 | 349449 | 1.003 | 349534 | 1.002 | 349615 | 1.002
349604 | 1.002 | 349608 | 1.001 | 349569 | 1.003 | 349654 | 1.002 | 349734 | 1.002
349724 | 1.002 | 349728 | 1.001 | 349688 | 1.003 | 349773 | 1.002 | 349854 | 1.002
349844 | 1.002 | 349849 | 1.001 | 349808 | 1.003 | 349893 | 1.002 | 349974 | 1.001
349964 | 1.002 | 349969 | 1.001 | 349928 | 1.003 | 350013 | 1.002 | 350094 | 1.002
350084 | 1.002 | 350089 | 1.001 | 350048 | 1.003 | 350133 | 1.002 | 350214 | 1.002
350203 | 1.002 | 350209 | 1.001 | 350168 | 1.003 | 350253 | 1.002 | 350333 | 1.002
350323 | 1.002 | 350328 | 1.001 | 350287 | 1.003 | 350372 | 1.002 | 350453 | 1.002
350443 | 1.002 | 350448 | 1.001 | 350407 | 1.003 | 350492 | 1.002 | 350573 | 1.002
350563 | 1.002 | 350568 | 1.001 | 350527 | 1.003 | 350612 | 1.002 | 350693 | 1.002
350683 | 1.002 | 350688 | 1.001 | 350647 | 1.003 | 350732 | 1.002 | 350835 | 1.002
350802 | 1.003 | 350808 | 1.001 | 350767 | 1.003 | 350852 | 1.002 | 350955 | 1.002
350922 | 1.002 | 350927 | 1.001 | 350906 | 1.003 | 350971 | 1.002 | 351075 | 1.002
351042 | 1.003 | 351047 | 1.001 | 351026 | 1.003 | 351091 | 1.002 | 351195 | 1.002
351162 | 1.002 | 351167 | 1.001 | 351145 | 1.003 | 351211 | 1.002 | 351314 | 1.002
351282 | 1.003 | 351287 | 1.001 | 351265 | 1.003 | 351331 | 1.002 | 351434 | 1.002
351401 | 1.003 | 351407 | 1.001 | 351385 | 1.003 | 351452 | 1.002 | 351554 | 1.002
351521 | 1.003 | 351526 | 1.001 | 351505 | 1.003 | 351571 | 1.002 | 351674 | 1.002
351641 | 1.003 | 351646 | 1.001 | 351625 | 1.003 | 351691 | 1.002 | 351794 | 1.002
351762 | 1.003 | 351766 | 1.001 | 351745 | 1.003 | 351811 | 1.002 | 351913 | 1.002
351882 | 1.003 | 351886 | 1.001 | 351867 | 1.003 | 351931 | 1.002 | 352033 | 1.002
352002 | 1.003 | 352006 | 1.001 | 351987 | 1.003 | 352051 | 1.002 | 352153 | 1.002
352121 | 1.002 | 352126 | 1.001 | 352107 | 1.002 | 352171 | 1.002 | 352273 | 1.001
352244 | 1.002 | 352246 | 1.001 | 352228 | 1.002 | 352290 | 1.002 | 352397 | 1.002
352364 | 1.003 | 352366 | 1.001 | 352348 | 1.003 | 352410 | 1.002 | 352517 | 1.002
352484 | 1.003 | 352486 | 1.001 | 352467 | 1.002 | 352533 | 1.002 | 352636 | 1.002
352604 | 1.003 | 352609 | 1.001 | 352587 | 1.003 | 352653 | 1.002 | 352756 | 1.002
352723 | 1.003 | 352729 | 1.001 | 352707 | 1.003 | 352773 | 1.002 | 352876 | 1.002
352843 | 1.003 | 352848 | 1.001 | 352827 | 1.003 | 352892 | 1.002 | 352996 | 1.002
352963 | 1.002 | 352968 | 1.001 | 352947 | 1.003 | 353012 | 1.002 | 353116 | 1.002
353083 | 1.002 | 353088 | 1.001 | 353067 | 1.003 | 353133 | 1.002 | 353235 | 1.001
353203 | 1.002 | 353208 | 1.001 | 353186 | 1.003 | 353253 | 1.002 | 353355 | 1.001
353324 | 1.003 | 353328 | 1.001 | 353306 | 1.002 | 353373 | 1.002 | 353475 | 1.002
353443 | 1.002 | 353447 | 1.001 | 353426 | 1.002 | 353492 | 1.002 | 353595 | 1.002
353563 | 1.003 | 353567 | 1.001 | 353546 | 1.002 | 353612 | 1.002 | 353715 | 1.002
353683 | 1.003 | 353687 | 1.001 | 353666 | 1.002 | 353732 | 1.002 | 353835 | 1.002
353803 | 1.003 | 353808 | 1.001 | 353785 | 1.002 | 353852 | 1.002 | 353954 | 1.002
353923 | 1.003 | 353928 | 1.001 | 353905 | 1.002 | 353972 | 1.002 | 354074 | 1.002
354042 | 1.002 | 354048 | 1.001 | 354025 | 1.003 | 354092 | 1.002 | 354195 | 1.002
354162 | 1.002 | 354167 | 1.001 | 354145 | 1.002 | 354211 | 1.002 | 354315 | 1.002
354282 | 1.003 | 354287 | 1.001 | 354265 | 1.002 | 354331 | 1.002 | 354435 | 1.002
354402 | 1.003 | 354407 | 1.001 | 354385 | 1.002 | 354451 | 1.002 | 354554 | 1.002
354522 | 1.003 | 354527 | 1.001 | 354505 | 1.002 | 354571 | 1.002 | 354674 | 1.002
354641 | 1.002 | 354647 | 1.001 | 354625 | 1.002 | 354691 | 1.002 | 354794 | 1.002
354761 | 1.002 | 354766 | 1.001 | 354745 | 1.003 | 354810 | 1.002 | 354914 | 1.002
354881 | 1.002 | 354886 | 1.001 | 354865 | 1.003 | 354930 | 1.002 | 355034 | 1.002
355001 | 1.002 | 355006 | 1.001 | 354984 | 1.002 | 355050 | 1.002 | 355153 | 1.002
355121 | 1.002 | 355126 | 1.001 | 355104 | 1.002 | 355170 | 1.002 | 355273 | 1.002
355240 | 1.002 | 355246 | 1.001 | 355224 | 1.002 | 355291 | 1.002 | 355393 | 1.002
355360 | 1.002 | 355365 | 1.001 | 355344 | 1.003 | 355410 | 1.002 | 355513 | 1.002
355481 | 1.002 | 355485 | 1.001 | 355464 | 1.002 | 355530 | 1.002 | 355633 | 1.002
355601 | 1.002 | 355605 | 1.001 | 355584 | 1.002 | 355650 | 1.002 | 355752 | 1.002
355721 | 1.002 | 355725 | 1.001 | 355703 | 1.002 | 355770 | 1.002 | 355872 | 1.002
355841 | 1.002 | 355845 | 1.001 | 355823 | 1.002 | 355890 | 1.002 | 355992 | 1.002
355960 | 1.002 | 355982 | 1.001 | 355943 | 1.002 | 356009 | 1.002 | 356112 | 1.002
356080 | 1.002 | 356102 | 1.001 | 356063 | 1.002 | 356129 | 1.002 | 356232 | 1.002
356200 | 1.002 | 356221 | 1.001 | 356183 | 1.002 | 356249 | 1.002 | 356353 | 1.002
356320 | 1.002 | 356341 | 1.001 | 356306 | 1.002 | 356369 | 1.002 | 356472 | 1.002
356440 | 1.002 | 356462 | 1.001 | 356426 | 1.002 | 356489 | 1.002 | 356592 | 1.002
356562 | 1.002 | 356582 | 1.001 | 356546 | 1.002 | 356609 | 1.002 | 356715 | 1.002
356682 | 1.002 | 356702 | 1.001 | 356666 | 1.002 | 356728 | 1.002 | 356835 | 1.002
356802 | 1.003 | 356822 | 1.001 | 356786 | 1.002 | 356848 | 1.002 | 356955 | 1.002
356923 | 1.002 | 356941 | 1.001 | 356905 | 1.002 | 356972 | 1.002 | 357074 | 1.002
357043 | 1.002 | 357064 | 1.001 | 357025 | 1.002 | 357092 | 1.002 | 357194 | 1.001
357163 | 1.002 | 357184 | 1.001 | 357145 | 1.002 | 357212 | 1.002 | 357314 | 1.002
357282 | 1.002 | 357304 | 1.001 | 357265 | 1.002 | 357331 | 1.002 | 357434 | 1.002
357402 | 1.002 | 357424 | 1.001 | 357385 | 1.002 | 357451 | 1.002 | 357555 | 1.002
357522 | 1.002 | 357543 | 1.001 | 357506 | 1.002 | 357571 | 1.002 | 357675 | 1.002
357642 | 1.002 | 357664 | 1.001 | 357625 | 1.002 | 357711 | 1.002 | 357794 | 1.002
357762 | 1.002 | 357784 | 1.001 | 357745 | 1.002 | 357831 | 1.002 | 357914 | 1.002
357881 | 1.002 | 357904 | 1.001 | 357865 | 1.002 | 357951 | 1.002 | 358034 | 1.002
358001 | 1.002 | 358024 | 1.001 | 357985 | 1.002 | 358071 | 1.002 | 358154 | 1.002
358121 | 1.002 | 358144 | 1.001 | 358105 | 1.002 | 358191 | 1.002 | 358274 | 1.002
358241 | 1.002 | 358263 | 1.001 | 358224 | 1.002 | 358310 | 1.002 | 358393 | 1.002
358362 | 1.002 | 358383 | 1.001 | 358344 | 1.002 | 358430 | 1.002 | 358513 | 1.001
358481 | 1.002 | 358503 | 1.001 | 358464 | 1.002 | 358550 | 1.002 | 358633 | 1.001
358601 | 1.002 | 358623 | 1.001 | 358584 | 1.002 | 358670 | 1.002 | 358753 | 1.002
358721 | 1.002 | 358743 | 1.001 | 358704 | 1.002 | 358791 | 1.002 | 358873 | 1.002
358841 | 1.002 | 358862 | 1.001 | 358823 | 1.002 | 358910 | 1.002 | 358992 | 1.002
358961 | 1.002 | 358982 | 1.001 | 358944 | 1.002 | 359030 | 1.002 | 359113 | 1.002
359080 | 1.002 | 359102 | 1.001 | 359064 | 1.002 | 359150 | 1.002 | 359233 | 1.002
359200 | 1.002 | 359223 | 1.001 | 359184 | 1.002 | 359270 | 1.001 | 359353 | 1.002
359320 | 1.002 | 359343 | 1.001 | 359304 | 1.002 | 359390 | 1.002 | 359473 | 1.002
359440 | 1.002 | 359462 | 1.001 | 359424 | 1.002 | 359510 | 1.002 | 359592 | 1.002
359560 | 1.002 | 359582 | 1.001 | 359543 | 1.002 | 359629 | 1.002 | 359712 | 1.002
359680 | 1.002 | 359702 | 1.001 | 359663 | 1.002 | 359749 | 1.002 | 359832 | 1.002
359799 | 1.002 | 359822 | 1.001 | 359783 | 1.002 | 359869 | 1.002 | 359952 | 1.002
359920 | 1.002 | 359942 | 1.001 | 359903 | 1.002 | 359989 | 1.002 | 360072 | 1.002
360040 | 1.002 | 360061 | 1.001 | 360023 | 1.002 | 360109 | 1.002 | 360192 | 1.002
360160 | 1.002 | 360181 | 1.001 | 360142 | 1.002 | 360229 | 1.001 | 360311 | 1.002
360296 | 1.002 | 360301 | 1.001 | 360262 | 1.002 | 360349 | 1.002 | 360431 | 1.002
360416 | 1.002 | 360421 | 1.001 | 360383 | 1.002 | 360469 | 1.001 | 360552 | 1.002
360536 | 1.002 | 360541 | 1.001 | 360503 | 1.002 | 360589 | 1.001 | 360672 | 1.002
360655 | 1.002 | 360660 | 1.001 | 360626 | 1.002 | 360709 | 1.002 | 360792 | 1.002
360775 | 1.002 | 360781 | 1.002 | 360745 | 1.003 | 360828 | 1.001 | 360911 | 1.002
360895 | 1.002 | 360901 | 1.012 | 360865 | 0.991 | 360948 | 1.001 | 361031 | 1.002
361018 | 1.002 | 361021 | 0.86 | 360985 | 0.802 | 361068 | 1.002 | 361154 | 1.002
361138 | 1.002 | 361141 | 0.014 | 361105 | 0.285 | 361188 | 1.002 | 361274 | 1.002
361259 | 1.002 | 361261 | 0.002 | 361225 | 0.002 | 361312 | 1.001 | 361394 | 1.003
361378 | 1.004 | 361383 | 0.002 | 361345 | 0.002 | 361432 | 1.001 | 361513 | 0.976
361498 | 0.583 | 361503 | 0.001 | 361465 | 0.002 | 361551 | 1.001 | 361634 | 0.677
361618 | 0.087 | 361623 | 0.001 | 361585 | 0.002 | 361671 | 1.001 | 361754 | 0.153
361738 | 0.003 | 361743 | 0.001 | 361705 | 0.002 | 361791 | 1.011 | 361874 | 0.006
361858 | 0.003 | 361864 | 0.001 | 361825 | 0.002 | 361911 | 1.01 | 361994 | 0.002
361977 | 0.003 | 361984 | 0.002 | 361945 | 0.002 | 362031 | 0.772 | 362114 | 0.002
362097 | 0.003 | 362103 | 0.001 | 362064 | 0.002 | 362150 | 0.01 | 362233 | 0.002
362217 | 0.003 | 362223 | 0.001 | 362184 | 0.002 | 362270 | 0.002 | 362353 | 0.002
362337 | 0.003 | 362343 | 0.001 | 362304 | 0.002 | 362390 | 0.002 | 362473 | 0.002
362457 | 0.003 | 362463 | 0.001 | 362424 | 0.002 | 362510 | 0.002 | 362593 | 0.002
362576 | 0.003 | 362583 | 0.001 | 362544 | 0.002 | 362630 | 0.002 | 362713 | 0.002
362696 | 0.003 | 362702 | 0.001 | 362663 | 0.002 | 362750 | 0.002 | 362832 | 0.002
362817 | 0.003 | 362822 | 0.001 | 362783 | 0.002 | 362870 | 0.002 | 362952 | 0.002
362937 | 0.003 | 362942 | 0.001 | 362903 | 0.002 | 362990 | 0.002 | 363072 | 0.002
363057 | 0.003 | 363062 | 0.001 | 363023 | 0.002 | 363110 | 0.002 | 363193 | 0.002
363176 | 0.003 | 363182 | 0.001 | 363144 | 0.002 | 363230 | 0.002 | 363313 | 0.002
363296 | 0.003 | 363301 | 0.001 | 363264 | 0.002 | 363350 | 0.002 | 363432 | 0.002
363416 | 0.003 | 363422 | 0.001 | 363383 | 0.002 | 363469 | 0.002 | 363574 | 0.002
363536 | 0.002 | 363542 | 0.001 | 363503 | 0.002 | 363589 | 0.002 | 363694 | 0.002
363656 | 0.003 | 363662 | 0.001 | 363623 | 0.002 | 363709 | 0.002 | 363813 | 0.002
363776 | 0.003 | 363782 | 0.001 | 363743 | 0.002 | 363829 | 0.002 | 363933 | 0.002
363895 | 0.003 | 363901 | 0.001 | 363863 | 0.002 | 363949 | 0.002 | 364053 | 0.002
364015 | 0.003 | 364021 | 0.001 | 363982 | 0.002 | 364068 | 0.002 | 364173 | 0.002
364135 | 0.003 | 364141 | 0.002 | 364121 | 0.002 | 364188 | 0.002 | 364293 | 0.002
364255 | 0.003 | 364261 | 0.002 | 364240 | 0.002 | 364308 | 0.002 | 364413 | 0.002
364375 | 0.003 | 364381 | 0.002 | 364360 | 0.002 | 364429 | 0.002 | 364533 | 0.002
364495 | 0.003 | 364500 | 0.002 | 364480 | 0.002 | 364549 | 0.002 | 364653 | 0.002
364615 | 0.003 | 364620 | 0.002 | 364601 | 0.002 | 364668 | 0.002 | 364773 | 0.002
364735 | 0.003 | 364740 | 0.002 | 364721 | 0.002 | 364788 | 0.002 | 364893 | 0.002
364855 | 0.003 | 364860 | 0.002 | 364840 | 0.002 | 364908 | 0.002 | 365012 | 0.002
364975 | 0.002 | 364980 | 0.002 | 364960 | 0.003 | 365028 | 0.002 | 365132 | 0.002
365094 | 0.002 | 365101 | 0.002 | 365083 | 0.003 | 365148 | 0.002 | 365252 | 0.002
365214 | 0.002 | 365220 | 0.002 | 365203 | 0.003 | 365267 | 0.002 | 365372 | 0.002
365337 | 0.003 | 365340 | 0.001 | 365323 | 0.003 | 365387 | 0.002 | 365495 | 0.002
365457 | 0.003 | 365460 | 0.002 | 365443 | 0.003 | 365507 | 0.002 | 365615 | 0.002
365577 | 0.003 | 365580 | 0.002 | 365562 | 0.003 | 365627 | 0.002 | 365735 | 0.002
365698 | 0.002 | 365700 | 0.002 | 365683 | 0.003 | 365751 | 0.002 | 365855 | 0.002
365817 | 0.002 | 365822 | 0.002 | 365803 | 0.003 | 365871 | 0.002 | 365975 | 0.002
365937 | 0.002 | 365942 | 0.002 | 365923 | 0.003 | 365990 | 0.002 | 366095 | 0.002
366057 | 0.003 | 366062 | 0.002 | 366043 | 0.003 | 366110 | 0.002 | 366215 | 0.002
366177 | 0.003 | 366182 | 0.002 | 366162 | 0.003 | 366230 | 0.002 | 366334 | 0.002
366297 | 0.003 | 366303 | 0.002 | 366282 | 0.002 | 366350 | 0.002 | 366454 | 0.002
366416 | 0.003 | 366423 | 0.002 | 366402 | 0.003 | 366470 | 0.002 | 366574 | 0.002
366536 | 0.003 | 366542 | 0.002 | 366522 | 0.003 | 366589 | 0.002 | 366694 | 0.002
366656 | 0.003 | 366662 | 0.002 | 366642 | 0.003 | 366709 | 0.002 | 366814 | 0.002
366776 | 0.003 | 366782 | 0.002 | 366761 | 0.003 | 366829 | 0.002 | 366934 | 0.002
366896 | 0.003 | 366902 | 0.002 | 366881 | 0.003 | 366949 | 0.002 | 367053 | 0.002
367015 | 0.003 | 367022 | 0.002 | 367001 | 0.003 | 367069 | 0.002 | 367173 | 0.002
367135 | 0.003 | 367141 | 0.002 | 367121 | 0.003 | 367190 | 0.002 | 367293 | 0.002
367255 | 0.002 | 367261 | 0.002 | 367242 | 0.003 | 367309 | 0.002 | 367414 | 0.002
367376 | 0.003 | 367381 | 0.002 | 367362 | 0.003 | 367429 | 0.002 | 367534 | 0.002
367496 | 0.003 | 367501 | 0.002 | 367481 | 0.003 | 367549 | 0.002 | 367653 | 0.002
367616 | 0.003 | 367621 | 0.002 | 367601 | 0.003 | 367669 | 0.002 | 367773 | 0.002
367735 | 0.003 | 367740 | 0.002 | 367721 | 0.003 | 367789 | 0.002 | 367893 | 0.002
367855 | 0.003 | 367860 | 0.002 | 367841 | 0.003 | 367908 | 0.002 | 368013 | 0.002
367975 | 0.003 | 367981 | 0.002 | 367961 | 0.003 | 368028 | 0.002 | 368133 | 0.002
368095 | 0.003 | 368101 | 0.002 | 368080 | 0.003 | 368148 | 0.002 | 368252 | 0.002
368215 | 0.002 | 368221 | 0.002 | 368200 | 0.003 | 368268 | 0.002 | 368372 | 0.002
368334 | 0.002 | 368340 | 0.002 | 368320 | 0.003 | 368388 | 0.002 | 368492 | 0.002
368454 | 0.002 | 368460 | 0.002 | 368440 | 0.003 | 368507 | 0.002 | 368612 | 0.002
368574 | 0.002 | 368580 | 0.002 | 368560 | 0.003 | 368627 | 0.002 | 368732 | 0.002
368694 | 0.003 | 368700 | 0.002 | 368679 | 0.003 | 368748 | 0.002 | 368851 | 0.002
368814 | 0.003 | 368820 | 0.002 | 368799 | 0.003 | 368868 | 0.002 | 368971 | 0.002
368933 | 0.003 | 368940 | 0.001 | 368920 | 0.003 | 368988 | 0.002 | 369092 | 0.002
369054 | 0.002 | 369059 | 0.001 | 369040 | 0.003 | 369107 | 0.002 | 369212 | 0.002
369174 | 0.002 | 369179 | 0.002 | 369160 | 0.002 | 369227 | 0.002 | 369332 | 0.002
369294 | 0.002 | 369316 | 0.002 | 369280 | 0.002 | 369347 | 0.002 | 369452 | 0.002
369414 | 0.002 | 369436 | 0.001 | 369402 | 0.003 | 369467 | 0.002 | 369571 | 0.002
369533 | 0.003 | 369556 | 0.001 | 369522 | 0.002 | 369587 | 0.002 | 369691 | 0.002
369653 | 0.002 | 369676 | 0.001 | 369642 | 0.003 | 369707 | 0.002 | 369811 | 0.002
369776 | 0.002 | 369796 | 0.002 | 369762 | 0.003 | 369826 | 0.002 | 369934 | 0.002
369896 | 0.002 | 369916 | 0.001 | 369882 | 0.003 | 369946 | 0.002 | 370054 | 0.002
370016 | 0.002 | 370036 | 0.001 | 370001 | 0.002 | 370070 | 0.002 | 370173 | 0.002
370137 | 0.002 | 370159 | 0.001 | 370122 | 0.003 | 370190 | 0.002 | 370293 | 0.002
370256 | 0.002 | 370279 | 0.001 | 370242 | 0.002 | 370310 | 0.002 | 370414 | 0.002
370376 | 0.002 | 370399 | 0.001 | 370362 | 0.002 | 370429 | 0.002 | 370534 | 0.002
370496 | 0.002 | 370519 | 0.001 | 370482 | 0.003 | 370570 | 0.002 | 370654 | 0.002
370616 | 0.002 | 370638 | 0.002 | 370601 | 0.002 | 370690 | 0.002 | 370774 | 0.002
370736 | 0.002 | 370758 | 0.001 | 370721 | 0.002 | 370809 | 0.002 | 370893 | 0.002
370855 | 0.002 | 370878 | 0.001 | 370841 | 0.003 | 370929 | 0.002 | 371013 | 0.002
370975 | 0.002 | 370998 | 0.001 | 370961 | 0.002 | 371049 | 0.002 | 371133 | 0.002
371095 | 0.002 | 371119 | 0.001 | 371081 | 0.002 | 371169 | 0.002 | 371253 | 0.002
371215 | 0.002 | 371238 | 0.001 | 371201 | 0.002 | 371289 | 0.002 | 371373 | 0.002
371335 | 0.002 | 371358 | 0.001 | 371320 | 0.002 | 371408 | 0.002 | 371492 | 0.002
371454 | 0.002 | 371478 | 0.002 | 371440 | 0.002 | 371528 | 0.002 | 371612 | 0.002
371574 | 0.002 | 371598 | 0.001 | 371560 | 0.002 | 371648 | 0.002 | 371732 | 0.002
371694 | 0.002 | 371718 | 0.001 | 371680 | 0.002 | 371768 | 0.002 | 371852 | 0.002
371815 | 0.002 | 371837 | 0.001 | 371801 | 0.002 | 371889 | 0.002 | 371972 | 0.002
371935 | 0.002 | 371957 | 0.001 | 371920 | 0.002 | 372008 | 0.002 | 372092 | 0.002
372055 | 0.002 | 372077 | 0.001 | 372040 | 0.002 | 372128 | 0.002 | 372212 | 0.002
372174 | 0.002 | 372197 | 0.001 | 372160 | 0.002 | 372248 | 0.002 | 372332 | 0.002
372294 | 0.002 | 372317 | 0.001 | 372280 | 0.002 | 372368 | 0.002 | 372452 | 0.002
372414 | 0.002 | 372436 | 0.001 | 372400 | 0.002 | 372488 | 0.002 | 372572 | 0.002
372534 | 0.002 | 372556 | 0.001 | 372519 | 0.002 | 372608 | 0.002 | 372691 | 0.002
372654 | 0.002 | 372677 | 0.001 | 372639 | 0.002 | 372727 | 0.002 | 372811 | 0.002
372773 | 0.002 | 372797 | 0.001 | 372759 | 0.002 | 372847 | 0.002 | 372931 | 0.002
372893 | 0.002 | 372917 | 0.001 | 372879 | 0.002 | 372967 | 0.002 | 373051 | 0.002
373013 | 0.002 | 373037 | 0.001 | 372999 | 0.002 | 373087 | 0.002 | 373171 | 0.002
373133 | 0.002 | 373156 | 0.001 | 373118 | 0.002 | 373207 | 0.002 | 373291 | 0.003
373253 | 0.002 | 373276 | 0.001 | 373238 | 0.002 | 373326 | 0.002 | 373410 | 0.002
373372 | 0.002 | 373396 | 0.002 | 373359 | 0.002 | 373447 | 0.002 | 373530 | 0.002
373493 | 0.002 | 373516 | 0.002 | 373479 | 0.002 | 373567 | 0.002 | 373650 | 0.002
373613 | 0.002 | 373636 | 0.002 | 373599 | 0.002 | 373687 | 0.002 | 373771 | 0.002
373748 | 0.002 | 373755 | 0.002 | 373719 | 0.002 | 373807 | 0.002 | 373891 | 0.003
373868 | 0.002 | 373875 | 0.002 | 373841 | 0.002 | 373926 | 0.002 | 374010 | 0.003
373989 | 0.002 | 373995 | 0.002 | 373961 | 0.002 | 374046 | 0.002 | 374130 | 0.003
374112 | 0.002 | 374115 | 0.002 | 374081 | 0.002 | 374166 | 0.002 | 374253 | 0.002
374232 | 0.003 | 374235 | 0.002 | 374201 | 0.002 | 374286 | 0.002 | 374373 | 0.002
374351 | 0.002 | 374355 | 0.002 | 374321 | 0.002 | 374406 | 0.002 | 374493 | 0.002
374471 | 0.003 | 374475 | 0.002 | 374440 | 0.002 | 374529 | 0.002 | 374612 | 0.002
374591 | 0.003 | 374598 | 0.002 | 374561 | 0.002 | 374649 | 0.002 | 374732 | 0.002
374711 | 0.003 | 374718 | 0.002 | 374681 | 0.002 | 374769 | 0.002 | 374853 | 0.002
374831 | 0.003 | 374838 | 0.002 | 374801 | 0.002 | 374889 | 0.002 | 374973 | 0.003
374951 | 0.003 | 374958 | 0.002 | 374921 | 0.002 | 375009 | 0.002 | 375093 | 0.003
375071 | 0.003 | 375077 | 0.002 | 375041 | 0.002 | 375129 | 0.002 | 375213 | 0.003
375191 | 0.003 | 375197 | 0.002 | 375160 | 0.002 | 375248 | 0.002 | 375332 | 0.002
375311 | 0.003 | 375317 | 0.002 | 375280 | 0.002 | 375368 | 0.002 | 375452 | 0.002
375441 | 0.003 | 375438 | 0.002 | 375400 | 0.002 | 375488 | 0.002 | 375583 | 0.003
375561 | 0.003 | 375558 | 0.002 | 375520 | 0.002 | 375608 | 0.002 | 375703 | 0.002
375681 | 0.003 | 375677 | 0.002 | 375651 | 0.002 | 375728 | 0.002 | 375823 | 0.003
375801 | 0.003 | 375797 | 0.002 | 375771 | 0.002 | 375859 | 0.002 | 375943 | 0.002
375921 | 0.003 | 375917 | 0.002 | 375890 | 0.002 | 375979 | 0.002 | 376062 | 0.003
376041 | 0.003 | 376047 | 0.002 | 376010 | 0.002 | 376098 | 0.002 | 376204 | 0.002
376161 | 0.003 | 376168 | 0.002 | 376130 | 0.002 | 376218 | 0.002 | 376324 | 0.002
376281 | 0.003 | 376288 | 0.002 | 376250 | 0.002 | 376338 | 0.002 | 376443 | 0.002
376400 | 0.003 | 376408 | 0.002 | 376370 | 0.002 | 376458 | 0.002 | 376563 | 0.002
376520 | 0.003 | 376527 | 0.002 | 376489 | 0.002 | 376578 | 0.002 | 376683 | 0.002
376640 | 0.003 | 376647 | 0.002 | 376609 | 0.002 | 376697 | 0.002 | 376804 | 0.002
376760 | 0.003 | 376767 | 0.001 | 376730 | 0.002 | 376817 | 0.002 | 376924 | 0.003
376880 | 0.003 | 376887 | 0.001 | 376850 | 0.002 | 376938 | 0.002 | 377043 | 0.002
376999 | 0.003 | 377007 | 0.002 | 376970 | 0.002 | 377058 | 0.002 | 377163 | 0.002
377119 | 0.003 | 377126 | 0.002 | 377090 | 0.002 | 377178 | 0.002 | 377283 | 0.003
377239 | 0.003 | 377246 | 0.002 | 377228 | 0.002 | 377297 | 0.002 | 377403 | 0.003
377359 | 0.003 | 377366 | 0.002 | 377348 | 0.002 | 377417 | 0.002 | 377523 | 0.002
377480 | 0.003 | 377486 | 0.002 | 377467 | 0.002 | 377537 | 0.002 | 377643 | 0.002
377600 | 0.003 | 377606 | 0.002 | 377587 | 0.002 | 377657 | 0.002 | 377762 | 0.002
377719 | 0.003 | 377725 | 0.002 | 377707 | 0.002 | 377777 | 0.002 | 377882 | 0.002
377839 | 0.003 | 377846 | 0.002 | 377827 | 0.002 | 377896 | 0.002 | 378002 | 0.002
377959 | 0.003 | 377966 | 0.002 | 377947 | 0.002 | 378016 | 0.002 | 378122 | 0.002
378079 | 0.003 | 378086 | 0.002 | 378066 | 0.002 | 378136 | 0.002 | 378242 | 0.002
378199 | 0.003 | 378206 | 0.002 | 378187 | 0.002 | 378256 | 0.002 | 378361 | 0.002
378318 | 0.003 | 378326 | 0.001 | 378304 | 0.002 | 378376 | 0.002 | 378482 | 0.002
378438 | 0.003 | 378445 | 0.001 | 378424 | 0.002 | 378497 | 0.002 | 378602 | 0.002
378561 | 0.003 | 378565 | 0.001 | 378544 | 0.002 | 378616 | 0.002 | 378725 | 0.002
378682 | 0.003 | 378685 | 0.001 | 378664 | 0.002 | 378736 | 0.002 | 378845 | 0.002
378802 | 0.003 | 378805 | 0.001 | 378784 | 0.002 | 378859 | 0.002 | 378964 | 0.002
378921 | 0.003 | 378929 | 0.001 | 378904 | 0.002 | 378979 | 0.002 | 379084 | 0.002
379041 | 0.003 | 379048 | 0.001 | 379024 | 0.002 | 379099 | 0.002 | 379204 | 0.003
379161 | 0.003 | 379168 | 0.001 | 379144 | 0.002 | 379218 | 0.002 | 379324 | 0.002
379281 | 0.003 | 379288 | 0.001 | 379264 | 0.002 | 379338 | 0.002 | 379444 | 0.002
379401 | 0.003 | 379408 | 0.001 | 379383 | 0.002 | 379458 | 0.002 | 379564 | 0.002
379521 | 0.003 | 379528 | 0.001 | 379503 | 0.002 | 379578 | 0.002 | 379684 | 0.002
379640 | 0.003 | 379648 | 0.001 | 379623 | 0.002 | 379699 | 0.002 | 379804 | 0.002
379760 | 0.003 | 379767 | 0.001 | 379743 | 0.002 | 379819 | 0.002 | 379924 | 0.002
379880 | 0.003 | 379887 | 0.001 | 379863 | 0.002 | 379938 | 0.002 | 380044 | 0.002
380000 | 0.003 | 380007 | 0.001 | 379982 | 0.002 | 380058 | 0.002 | 380164 | 0.002
380120 | 0.003 | 380127 | 0.001 | 380102 | 0.002 | 380178 | 0.002 | 380283 | 0.002
380239 | 0.003 | 380247 | 0.001 | 380222 | 0.002 | 380298 | 0.002 | 380403 | 0.002
380360 | 0.003 | 380366 | 0.001 | 380343 | 0.002 | 380418 | 0.002 | 380523 | 0.002
380480 | 0.003 | 380486 | 0.001 | 380463 | 0.002 | 380537 | 0.002 | 380643 | 0.002
380600 | 0.003 | 380607 | 0.001 | 380582 | 0.002 | 380657 | 0.002 | 380763 | 0.002
380720 | 0.003 | 380727 | 0.001 | 380702 | 0.002 | 380777 | 0.002 | 380882 | 0.002
380839 | 0.003 | 380847 | 0.001 | 380822 | 0.002 | 380897 | 0.002 | 381002 | 0.002
380959 | 0.003 | 380966 | 0.001 | 380942 | 0.002 | 381017 | 0.002 | 381122 | 0.002
381079 | 0.003 | 381086 | 0.001 | 381062 | 0.002 | 381136 | 0.002 | 381243 | 0.002
381199 | 0.003 | 381206 | 0.001 | 381181 | 0.002 | 381256 | 0.002 | 381363 | 0.002
381319 | 0.003 | 381326 | 0.001 | 381301 | 0.002 | 381377 | 0.002 | 381483 | 0.002
381438 | 0.003 | 381446 | 0.001 | 381421 | 0.002 | 381497 | 0.002 | 381602 | 0.002
381558 | 0.003 | 381565 | 0.001 | 381541 | 0.002 | 381617 | 0.002 | 381722 | 0.002
381678 | 0.003 | 381685 | 0.001 | 381661 | 0.002 | 381736 | 0.002 | 381842 | 0.002
381798 | 0.003 | 381805 | 0.001 | 381780 | 0.002 | 381856 | 0.002 | 381962 | 0.002
381918 | 0.002 | 381925 | 0.001 | 381901 | 0.002 | 381976 | 0.002 | 382082 | 0.002
382038 | 0.003 | 382045 | 0.001 | 382021 | 0.002 | 382096 | 0.002 | 382201 | 0.002
382158 | 0.003 | 382164 | 0.001 | 382141 | 0.002 | 382216 | 0.002 | 382321 | 0.002
382278 | 0.003 | 382285 | 0.001 | 382261 | 0.002 | 382336 | 0.002 | 382441 | 0.002
382398 | 0.003 | 382405 | 0.001 | 382381 | 0.002 | 382455 | 0.002 | 382561 | 0.002
382518 | 0.003 | 382525 | 0.001 | 382500 | 0.002 | 382575 | 0.002 | 382681 | 0.002
382638 | 0.003 | 382662 | 0.001 | 382623 | 0.002 | 382695 | 0.002 | 382800 | 0.002
382757 | 0.003 | 382782 | 0.001 | 382743 | 0.002 | 382815 | 0.002 | 382921 | 0.002
382877 | 0.003 | 382902 | 0.001 | 382863 | 0.002 | 382935 | 0.002 | 383041 | 0.002
383000 | 0.003 | 383022 | 0.001 | 382983 | 0.002 | 383055 | 0.002 | 383164 | 0.002
383120 | 0.003 | 383141 | 0.001 | 383104 | 0.002 | 383175 | 0.002 | 383284 | 0.002
383240 | 0.003 | 383261 | 0.001 | 383223 | 0.002 | 383298 | 0.002 | 383404 | 0.002
383361 | 0.003 | 383384 | 0.001 | 383343 | 0.002 | 383418 | 0.002 | 383523 | 0.002
383480 | 0.003 | 383504 | 0.001 | 383463 | 0.002 | 383558 | 0.002 | 383643 | 0.002
383600 | 0.002 | 383624 | 0.001 | 383583 | 0.002 | 383678 | 0.002 | 383763 | 0.002
383720 | 0.002 | 383745 | 0.001 | 383703 | 0.002 | 383798 | 0.002 | 383883 | 0.002
383840 | 0.003 | 383864 | 0.001 | 383822 | 0.002 | 383918 | 0.002 | 384003 | 0.002
383960 | 0.003 | 383984 | 0.001 | 383942 | 0.002 | 384037 | 0.002 | 384123 | 0.002
384079 | 0.003 | 384104 | 0.001 | 384062 | 0.002 | 384157 | 0.002 | 384243 | 0.002
384199 | 0.003 | 384224 | 0.001 | 384182 | 0.002 | 384277 | 0.002 | 384363 | 0.002
384319 | 0.003 | 384344 | 0.001 | 384302 | 0.002 | 384397 | 0.002 | 384483 | 0.002
384439 | 0.003 | 384463 | 0.001 | 384421 | 0.002 | 384517 | 0.002 | 384603 | 0.002
384559 | 0.002 | 384583 | 0.001 | 384541 | 0.002 | 384636 | 0.002 | 384722 | 0.002
384678 | 0.002 | 384703 | 0.001 | 384662 | 0.002 | 384757 | 0.002 | 384842 | 0.002
384798 | 0.003 | 384823 | 0.001 | 384782 | 0.002 | 384877 | 0.002 | 384962 | 0.002
384918 | 0.002 | 384943 | 0.001 | 384902 | 0.002 | 384997 | 0.002 | 385082 | 0.002
385039 | 0.002 | 385062 | 0.001 | 385021 | 0.002 | 385117 | 0.002 | 385202 | 0.002
385159 | 0.003 | 385183 | 0.001 | 385141 | 0.002 | 385236 | 0.002 | 385321 | 0.002
385278 | 0.002 | 385303 | 0.001 | 385261 | 0.002 | 385356 | 0.002 | 385441 | 0.002
385398 | 0.002 | 385423 | 0.001 | 385381 | 0.002 | 385476 | 0.002 | 385561 | 0.002
385518 | 0.002 | 385543 | 0.001 | 385501 | 0.002 | 385596 | 0.002 | 385681 | 0.002
385638 | 0.002 | 385662 | 0.001 | 385620 | 0.002 | 385716 | 0.002 | 385802 | 0.002
385758 | 0.002 | 385782 | 0.001 | 385740 | 0.002 | 385836 | 0.002 | 385922 | 0.002
385878 | 0.002 | 385902 | 0.001 | 385860 | 0.002 | 385955 | 0.002 | 386041 | 0.002
385997 | 0.002 | 386022 | 0.001 | 385980 | 0.002 | 386075 | 0.002 | 386161 | 0.002
386117 | 0.002 | 386142 | 0.001 | 386100 | 0.002 | 386195 | 0.002 | 386281 | 0.002
386237 | 0.002 | 386262 | 0.001 | 386221 | 0.002 | 386315 | 0.002 | 386401 | 0.002
386357 | 0.002 | 386381 | 0.001 | 386340 | 0.002 | 386436 | 0.002 | 386521 | 0.002
386477 | 0.002 | 386501 | 0.001 | 386460 | 0.002 | 386555 | 0.002 | 386640 | 0.002
386596 | 0.002 | 386621 | 0.001 | 386580 | 0.002 | 386675 | 0.002 | 386760 | 0.002
386717 | 0.002 | 386742 | 0.001 | 386700 | 0.002 | 386795 | 0.002 | 386880 | 0.002
386837 | 0.002 | 386862 | 0.001 | 386820 | 0.002 | 386915 | 0.002 | 387000 | 0.002
386957 | 0.002 | 386981 | 0.001 | 386939 | 0.002 | 387035 | 0.002 | 387120 | 0.002
387077 | 0.002 | 387101 | 0.001 | 387062 | 0.002 | 387154 | 0.002 | 387239 | 0.002
387196 | 0.002 | 387221 | 0.001 | 387182 | 0.002 | 387274 | 0.002 | 387359 | 0.002
387336 | 0.002 | 387341 | 0.001 | 387302 | 0.002 | 387394 | 0.002 | 387483 | 0.002
387456 | 0.002 | 387461 | 0.001 | 387423 | 0.002 | 387514 | 0.002 | 387603 | 0.002
387575 | 0.002 | 387580 | 0.001 | 387543 | 0.002 | 387634 | 0.002 | 387723 | 0.002
387695 | 0.002 | 387700 | 0.001 | 387662 | 0.002 | 387758 | 0.002 | 387843 | 0.002
387815 | 0.002 | 387824 | 0.001 | 387782 | 0.002 | 387877 | 0.002 | 387962 | 0.002
387935 | 0.002 | 387944 | 0.001 | 387902 | 0.002 | 387997 | 0.002 | 388082 | 0.002
388055 | 0.002 | 388064 | 0.001 | 388022 | 0.002 | 388117 | 0.002 | 388202 | 0.002
388175 | 0.002 | 388184 | 0.001 | 388142 | 0.002 | 388237 | 0.002 | 388322 | 0.002
388295 | 0.002 | 388303 | 0.001 | 388261 | 0.002 | 388357 | 0.002 | 388442 | 0.002
388415 | 0.002 | 388423 | 0.001 | 388381 | 0.002 | 388476 | 0.002 | 388561 | 0.002
388535 | 0.002 | 388543 | 0.001 | 388501 | 0.002 | 388596 | 0.002 | 388682 | 0.002
388655 | 0.002 | 388663 | 0.001 | 388621 | 0.002 | 388716 | 0.002 | 388802 | 0.002
388774 | 0.002 | 388783 | 0.001 | 388741 | 0.002 | 388836 | 0.002 | 388943 | 0.002
388894 | 0.002 | 388902 | 0.001 | 388860 | 0.002 | 388956 | 0.002 | 389063 | 0.002
389014 | 0.002 | 389022 | 0.001 | 388981 | 0.002 | 389075 | 0.002 | 389183 | 0.002
389134 | 0.002 | 389142 | 0.001 | 389101 | 0.002 | 389196 | 0.002 | 389303 | 0.002
389254 | 0.002 | 389263 | 0.001 | 389221 | 0.002 | 389316 | 0.002 | 389423 | 0.002
389373 | 0.002 | 389383 | 0.001 | 389341 | 0.002 | 389436 | 0.002 | 389542 | 0.002
389493 | 0.002 | 389502 | 0.001 | 389460 | 0.002 | 389556 | 0.002 | 389662 | 0.002
389613 | 0.002 | 389622 | 0.001 | 389580 | 0.002 | 389676 | 0.002 | 389782 | 0.002
389733 | 0.002 | 389742 | 0.001 | 389700 | 0.002 | 389795 | 0.002 | 389903 | 0.002
389854 | 0.002 | 389862 | 0.001 | 389820 | 0.002 | 389915 | 0.002 | 390023 | 0.002
389974 | 0.002 | 389982 | 0.001 | 389940 | 0.002 | 390035 | 0.002 | 390142 | 0.002
390093 | 0.002 | 390102 | 0.001 | 390060 | 0.002 | 390155 | 0.002 | 390262 | 0.002
390213 | 0.002 | 390221 | 0.001 | 390179 | 0.002 | 390275 | 0.002 | 390382 | 0.002
390333 | 0.002 | 390341 | 0.001 | 390299 | 0.002 | 390394 | 0.002 | 390502 | 0.002
390453 | 0.002 | 390461 | 0.001 | 390438 | 0.002 | 390514 | 0.002 | 390622 | 0.002
390573 | 0.002 | 390581 | 0.001 | 390558 | 0.002 | 390634 | 0.002 | 390742 | 0.002
390692 | 0.002 | 390702 | 0.001 | 390678 | 0.002 | 390754 | 0.002 | 390861 | 0.002
390812 | 0.002 | 390821 | 0.001 | 390798 | 0.002 | 390875 | 0.002 | 390981 | 0.002
390932 | 0.002 | 390941 | 0.001 | 390918 | 0.002 | 390994 | 0.002 | 391101 | 0.002
391052 | 0.002 | 391061 | 0.001 | 391037 | 0.002 | 391114 | 0.002 | 391221 | 0.002
391172 | 0.002 | 391181 | 0.001 | 391157 | 0.002 | 391234 | 0.002 | 391341 | 0.002
391291 | 0.002 | 391301 | 0.001 | 391277 | 0.002 | 391354 | 0.002 | 391460 | 0.002
391411 | 0.002 | 391420 | 0.001 | 391400 | 0.002 | 391474 | 0.002 | 391580 | 0.002
391532 | 0.002 | 391540 | 0.001 | 391521 | 0.002 | 391593 | 0.002 | 391701 | 0.002
391652 | 0.002 | 391660 | 0.001 | 391641 | 0.002 | 391713 | 0.002 | 391821 | 0.002
391775 | 0.002 | 391780 | 0.001 | 391760 | 0.002 | 391833 | 0.002 | 391944 | 0.002
391895 | 0.002 | 391900 | 0.001 | 391880 | 0.002 | 391953 | 0.002 | 392064 | 0.002
392014 | 0.002 | 392019 | 0.001 | 392000 | 0.002 | 392077 | 0.002 | 392183 | 0.002
392134 | 0.002 | 392143 | 0.001 | 392120 | 0.002 | 392197 | 0.002 | 392303 | 0.002
392254 | 0.002 | 392263 | 0.001 | 392240 | 0.002 | 392316 | 0.002 | 392423 | 0.002
392374 | 0.002 | 392383 | 0 | 392359 | 0.002 | 392436 | 0.002 | 392543 | 0.002
392495 | 0.002 | 392503 | 0.001 | 392479 | 0.002 | 392556 | 0.002 | 392664 | 0.002
392614 | 0.002 | 392623 | 0.001 | 392599 | 0.002 | 392676 | 0.002 | 392783 | 0.002
392734 | 0.002 | 392742 | 0.001 | 392720 | 0.002 | 392796 | 0.002 | 392903 | 0.002
392854 | 0.002 | 392862 | 0.001 | 392840 | 0.002 | 392915 | 0.002 | 393023 | 0.002
392974 | 0.002 | 392983 | 0.001 | 392960 | 0.002 | 393035 | 0.002 | 393143 | 0.002
393094 | 0.002 | 393103 | 0.001 | 393079 | 0.002 | 393155 | 0.002 | 393263 | 0.002
393213 | 0.002 | 393223 | 0.001 | 393199 | 0.002 | 393276 | 0.002 | 393382 | 0.002
393333 | 0.002 | 393342 | 0.001 | 393319 | 0.002 | 393396 | 0.002 | 393502 | 0.002
393453 | 0.002 | 393462 | 0.001 | 393439 | 0.002 | 393516 | 0.002 | 393622 | 0.002
393573 | 0.002 | 393582 | 0.001 | 393559 | 0.002 | 393635 | 0.002 | 393742 | 0.002
393694 | 0.002 | 393702 | 0.001 | 393678 | 0.002 | 393755 | 0.002 | 393863 | 0.002
393814 | 0.002 | 393822 | 0.001 | 393798 | 0.002 | 393875 | 0.002 | 393982 | 0.002
393933 | 0.002 | 393942 | 0.001 | 393919 | 0.002 | 393995 | 0.002 | 394102 | 0.002
394053 | 0.002 | 394061 | 0.001 | 394039 | 0.002 | 394115 | 0.002 | 394222 | 0.002
394173 | 0.002 | 394182 | 0.001 | 394159 | 0.002 | 394234 | 0.002 | 394342 | 0.002
394293 | 0.002 | 394302 | 0.001 | 394278 | 0.002 | 394354 | 0.002 | 394462 | 0.002
394413 | 0.002 | 394422 | 0.001 | 394398 | 0.001 | 394474 | 0.002 | 394582 | 0.002
394532 | 0.002 | 394542 | 0.001 | 394518 | 0.001 | 394595 | 0.003 | 394701 | 0.002
394652 | 0.002 | 394661 | 0.001 | 394638 | 0.001 | 394715 | 0.002 | 394821 | 0.002
394772 | 0.002 | 394781 | 0.001 | 394758 | 0.002 | 394834 | 0.002 | 394941 | 0.002
394893 | 0.002 | 394901 | 0.001 | 394877 | 0.001 | 394954 | 0.002 | 395062 | 0.002
395013 | 0.002 | 395021 | 0.001 | 394997 | 0.001 | 395074 | 0.002 | 395182 | 0.002
395132 | 0.002 | 395141 | 0.001 | 395118 | 0.002 | 395194 | 0.002 | 395301 | 0.002
395252 | 0.002 | 395261 | 0.001 | 395238 | 0.002 | 395314 | 0.002 | 395421 | 0.002
395372 | 0.002 | 395381 | 0.001 | 395358 | 0.002 | 395433 | 0.002 | 395541 | 0.002
395492 | 0.002 | 395501 | 0 | 395478 | 0.002 | 395553 | 0.002 | 395661 | 0.002
395612 | 0.002 | 395621 | 0.001 | 395597 | 0.002 | 395673 | 0.002 | 395781 | 0.002
395731 | 0.002 | 395741 | 0.001 | 395717 | 0.002 | 395793 | 0.002 | 395900 | 0.002
395851 | 0.002 | 395860 | 0.001 | 395840 | 0.002 | 395914 | 0.002 | 396020 | 0.002
395971 | 0.002 | 395998 | 0.001 | 395961 | 0.002 | 396034 | 0.003 | 396140 | 0.002
396095 | 0.002 | 396118 | 0.001 | 396081 | 0.001 | 396153 | 0.002 | 396264 | 0.002
396215 | 0.002 | 396237 | 0 | 396200 | 0.002 | 396273 | 0.002 | 396384 | 0.002
396335 | 0.002 | 396357 | 0 | 396320 | 0.002 | 396393 | 0.002 | 396504 | 0.002
396454 | 0.002 | 396477 | 0 | 396440 | 0.002 | 396536 | 0.002 | 396623 | 0.002
396574 | 0.002 | 396601 | 0 | 396560 | 0.002 | 396656 | 0.002 | 396743 | 0.002
396694 | 0.002 | 396721 | 0.001 | 396680 | 0.002 | 396776 | 0.002 | 396863 | 0.002
396814 | 0.002 | 396840 | 0 | 396800 | 0.002 | 396896 | 0.002 | 396983 | 0.002
396935 | 0.002 | 396960 | 0.001 | 396919 | 0.002 | 397016 | 0.002 | 397104 | 0.002
397054 | 0.002 | 397080 | 0 | 397039 | 0.002 | 397136 | 0.002 | 397223 | 0.002
397174 | 0.002 | 397200 | 0 | 397160 | 0.001 | 397256 | 0.002 | 397343 | 0.002
397294 | 0.002 | 397320 | 0.001 | 397280 | 0.001 | 397376 | 0.002 | 397463 | 0.002
397414 | 0.002 | 397441 | 0.001 | 397400 | 0.002 | 397496 | 0.002 | 397583 | 0.002
397534 | 0.002 | 397560 | 0.001 | 397519 | 0.002 | 397616 | 0.002 | 397703 | 0.002
397654 | 0.002 | 397680 | 0.001 | 397639 | 0.002 | 397735 | 0.002 | 397822 | 0.002
397773 | 0.002 | 397800 | 0 | 397759 | 0.002 | 397855 | 0.002 | 397942 | 0.002
397893 | 0.002 | 397920 | 0.001 | 397879 | 0.002 | 397975 | 0.003 | 398062 | 0.002
398013 | 0.002 | 398040 | 0 | 397999 | 0.002 | 398095 | 0.002 | 398182 | 0.002
398134 | 0.002 | 398159 | 0 | 398118 | 0.002 | 398215 | 0.002 | 398303 | 0.002
398254 | 0.002 | 398279 | 0 | 398238 | 0.001 | 398334 | 0.002 | 398423 | 0.002
398373 | 0.002 | 398399 | 0.001 | 398359 | 0.001 | 398455 | 0.002 | 398542 | 0.002
398493 | 0.002 | 398519 | 0.001 | 398479 | 0.002 | 398575 | 0.002 | 398662 | 0.002
398613 | 0.002 | 398640 | 0.001 | 398599 | 0.002 | 398695 | 0.002 | 398782 | 0.002
398733 | 0.002 | 398759 | 0.001 | 398718 | 0.002 | 398815 | 0.002 | 398902 | 0.002
398853 | 0.002 | 398879 | 0 | 398838 | 0.002 | 398935 | 0.002 | 399022 | 0.002
398972 | 0.002 | 398999 | 0 | 398958 | 0.002 | 399054 | 0.002 | 399141 | 0.002
399092 | 0.002 | 399119 | 0 | 399078 | 0.002 | 399174 | 0.002 | 399261 | 0.002
399212 | 0.002 | 399239 | 0 | 399198 | 0.002 | 399294 | 0.002 | 399381 | 0.002
399333 | 0.002 | 399358 | 0.001 | 399318 | 0.002 | 399414 | 0.002 | 399501 | 0.002
399453 | 0.002 | 399478 | 0.001 | 399437 | 0.002 | 399534 | 0.002 | 399622 | 0.002
399572 | 0.002 | 399598 | 0 | 399558 | 0.002 | 399653 | 0.002 | 399741 | 0.002
399692 | 0.002 | 399718 | 0 | 399678 | 0.001 | 399774 | 0.002 | 399861 | 0.002
399812 | 0.002 | 399839 | 0 | 399798 | 0.002 | 399894 | 0.002 | 399981 | 0.002
399932 | 0.002 | 399959 | 0 | 399918 | 0.002 | 400014 | 0.002 | 400101 | 0.002
400052 | 0.002 | 400078 | 0 | 400037 | 0.002 | 400134 | 0.002 | 400221 | 0.002
400172 | 0.002 | 400198 | 0 | 400160 | 0.002 | 400253 | 0.002 | 400340 | 0.002
400291 | 0.002 | 400318 | 0 | 400280 | 0.002 | 400373 | 0.002 | 400460 | 0.002
400411 | 0.002 | 400438 | 0 | 400401 | 0.002 | 400493 | 0.002 | 400580 | 0.002
400535 | 0.002 | 400558 | 0 | 400521 | 0.002 | 400613 | 0.002 | 400704 | 0.002
400655 | 0.002 | 400677 | 0 | 400641 | 0.002 | 400733 | 0.002 | 400824 | 0.002
400791 | 0.002 | 400797 | 0.001 | 400760 | 0.002 | 400857 | 0.002 | 400944 | 0.002
400911 | 0.002 | 400921 | 0.001 | 400880 | 0.002 | 400976 | 0.002 | 401063 | 0.002
401031 | 0.002 | 401041 | 0 | 401000 | 0.002 | 401096 | 0.002 | 401183 | 0.002
401150 | 0.002 | 401161 | 0 | 401120 | 0.002 | 401216 | 0.002 | 401303 | 0.002
401270 | 0.002 | 401281 | 0 | 401240 | 0.001 | 401336 | 0.002 | 401423 | 0.002
401390 | 0.002 | 401400 | 0 | 401359 | 0.001 | 401456 | 0.002 | 401565 | 0.002
401510 | 0.002 | 401520 | 0 | 401479 | 0.001 | 401575 | 0.002 | 401685 | 0.002
401631 | 0.002 | 401640 | 0 | 401600 | 0.002 | 401695 | 0.002 | 401805 | 0.002
401751 | 0.002 | 401760 | 0 | 401720 | 0.002 | 401816 | 0.002 | 401925 | 0.002
401870 | 0.002 | 401881 | 0 | 401840 | 0.002 | 401936 | 0.002 | 402044 | 0.002
401990 | 0.002 | 402000 | 0 | 401959 | 0.002 | 402056 | 0.002 | 402164 | 0.002
402110 | 0.002 | 402120 | 0 | 402079 | 0.002 | 402176 | 0.002 | 402284 | 0.002
402230 | 0.002 | 402240 | 0 | 402199 | 0.002 | 402295 | 0.002 | 402404 | 0.002
402350 | 0.002 | 402360 | 0 | 402319 | 0.002 | 402415 | 0.002 | 402524 | 0.002
402469 | 0.002 | 402480 | 0 | 402439 | 0.001 | 402535 | 0.002 | 402644 | 0.002
402589 | 0.002 | 402599 | 0 | 402558 | 0.002 | 402655 | 0.002 | 402764 | 0.002
402709 | 0.002 | 402719 | 0 | 402678 | 0.002 | 402775 | 0.002 | 402884 | 0.002
402830 | 0.002 | 402839 | 0 | 402798 | 0.002 | 402894 | 0.002 | 403004 | 0.002
402950 | 0.002 | 402959 | 0 | 402919 | 0.002 | 403015 | 0.002 | 403124 | 0.002
403069 | 0.002 | 403080 | 0 | 403039 | 0.002 | 403135 | 0.002 | 403244 | 0.002
403189 | 0.002 | 403200 | 0 | 403159 | 0.002 | 403255 | 0.002 | 403363 | 0.002
403309 | 0.002 | 403319 | 0 | 403278 | 0.002 | 403375 | 0.002 | 403483 | 0.002
403429 | 0.002 | 403439 | 0 | 403398 | 0.002 | 403494 | 0.002 | 403603 | 0.002
403549 | 0.002 | 403559 | 0 | 403518 | 0.002 | 403614 | 0.002 | 403723 | 0.002
403668 | 0.002 | 403679 | 0 | 403656 | 0.002 | 403734 | 0.002 | 403844 | 0.002
403788 | 0.002 | 403799 | 0 | 403776 | 0.001 | 403854 | 0.002 | 403963 | 0.002
403908 | 0.002 | 403918 | 0 | 403897 | 0.001 | 403974 | 0.002 | 404083 | 0.002
404029 | 0.002 | 404038 | 0 | 404017 | 0.002 | 404093 | 0.002 | 404203 | 0.002
404149 | 0.002 | 404158 | 0 | 404136 | 0.002 | 404213 | 0.002 | 404323 | 0.002
404269 | 0.002 | 404279 | 0 | 404256 | 0.002 | 404334 | 0.002 | 404443 | 0.002
404388 | 0.002 | 404399 | 0 | 404376 | 0.001 | 404454 | 0.002 | 404562 | 0.002
404508 | 0.002 | 404518 | 0 | 404496 | 0.001 | 404574 | 0.002 | 404682 | 0.002
404628 | 0.002 | 404638 | 0.001 | 404620 | 0.002 | 404694 | 0.002 | 404802 | 0.002
404748 | 0.002 | 404758 | 0 | 404740 | 0.002 | 404813 | 0.002 | 404923 | 0.002
404882 | 0.002 | 404878 | 0.001 | 404859 | 0.002 | 404933 | 0.002 | 405056 | 0.002
405002 | 0.002 | 404998 | 0.001 | 404979 | 0.002 | 405053 | 0.002 | 405177 | 0.002
405122 | 0.002 | 405117 | 0.001 | 405109 | 0.002 | 405173 | 0.002 | 405297 | 0.002
405241 | 0.002 | 405238 | 0 | 405230 | 0.002 | 405307 | 0.003 | 405416 | 0.002
405362 | 0.002 | 405361 | 0.001 | 405350 | 0.002 | 405427 | 0.003 | 405536 | 0.002
405482 | 0.002 | 405492 | 0.001 | 405470 | 0.002 | 405547 | 0.003 | 405656 | 0.002
405602 | 0.002 | 405612 | 0.001 | 405590 | 0.002 | 405667 | 0.002 | 405776 | 0.002
405722 | 0.002 | 405732 | 0.001 | 405709 | 0.002 | 405787 | 0.002 | 405896 | 0.002
405841 | 0.002 | 405852 | 0.001 | 405829 | 0.002 | 405907 | 0.002 | 406016 | 0.002
405961 | 0.002 | 405971 | 0.001 | 405949 | 0.002 | 406027 | 0.003 | 406135 | 0.002
406081 | 0.002 | 406091 | 0.001 | 406069 | 0.002 | 406147 | 0.003 | 406255 | 0.002
406201 | 0.002 | 406211 | 0.001 | 406189 | 0.002 | 406266 | 0.002 | 406375 | 0.002
406321 | 0.002 | 406331 | 0.001 | 406308 | 0.002 | 406386 | 0.002 | 406496 | 0.002
406440 | 0.002 | 406451 | 0.001 | 406429 | 0.002 | 406506 | 0.002 | 406616 | 0.002
406561 | 0.002 | 406571 | 0.001 | 406549 | 0.002 | 406626 | 0.003 | 406735 | 0.002
406681 | 0.002 | 406691 | 0.001 | 406669 | 0.002 | 406746 | 0.002 | 406855 | 0.002
406801 | 0.002 | 406811 | 0.001 | 406789 | 0.002 | 406865 | 0.002 | 406975 | 0.002
406921 | 0.002 | 406931 | 0.001 | 406908 | 0.002 | 406986 | 0.003 | 407095 | 0.002
407041 | 0.001 | 407051 | 0.001 | 407028 | 0.002 | 407106 | 0.003 | 407215 | 0.002
407160 | 0.001 | 407171 | 0.001 | 407148 | 0.002 | 407226 | 0.003 | 407334 | 0.002
407280 | 0.002 | 407290 | 0.001 | 407268 | 0.002 | 407346 | 0.003 | 407454 | 0.002
407400 | 0.002 | 407410 | 0.001 | 407388 | 0.002 | 407465 | 0.003 | 407574 | 0.002
407520 | 0.001 | 407530 | 0.001 | 407507 | 0.002 | 407585 | 0.003 | 407695 | 0.002
407640 | 0.002 | 407650 | 0.001 | 407628 | 0.002 | 407705 | 0.003 | 407815 | 0.002
407759 | 0.002 | 407770 | 0.001 | 407748 | 0.002 | 407825 | 0.003 | 407934 | 0.002
407880 | 0.002 | 407890 | 0.001 | 407868 | 0.002 | 407945 | 0.003 | 408054 | 0.002
408000 | 0.002 | 408010 | 0.001 | 407988 | 0.002 | 408065 | 0.003 | 408174 | 0.002
408120 | 0.001 | 408130 | 0.001 | 408108 | 0.002 | 408185 | 0.003 | 408294 | 0.002
408240 | 0.002 | 408250 | 0.001 | 408227 | 0.002 | 408305 | 0.003 | 408414 | 0.002
408359 | 0.001 | 408370 | 0.001 | 408347 | 0.002 | 408425 | 0.003 | 408534 | 0.002
408479 | 0.002 | 408489 | 0.001 | 408467 | 0.002 | 408545 | 0.003 | 408653 | 0.002
408599 | 0.001 | 408609 | 0 | 408587 | 0.002 | 408665 | 0.003 | 408773 | 0.002
408719 | 0.001 | 408729 | 0.001 | 408707 | 0.002 | 408784 | 0.003 | 408894 | 0.002
408839 | 0.002 | 408849 | 0.001 | 408827 | 0.002 | 408904 | 0.003 | 409014 | 0.002
408958 | 0.001 | 408969 | 0.001 | 408947 | 0.002 | 409024 | 0.003 | 409134 | 0.002
409079 | 0.002 | 409090 | 0.001 | 409064 | 0.002 | 409144 | 0.003 | 409253 | 0.002
409199 | 0.002 | 409209 | 0 | 409185 | 0.002 | 409264 | 0.003 | 409373 | 0.002
409322 | 0.002 | 409347 | 0.001 | 409305 | 0.002 | 409404 | 0.003 | 409496 | 0.002
409442 | 0.002 | 409466 | 0 | 409424 | 0.002 | 409524 | 0.003 | 409616 | 0.002
409562 | 0.002 | 409586 | 0 | 409544 | 0.002 | 409648 | 0.003 | 409737 | 0.002
409681 | 0.002 | 409709 | 0 | 409664 | 0.002 | 409767 | 0.003 | 409857 | 0.002
409802 | 0.002 | 409829 | 0 | 409784 | 0.002 | 409887 | 0.003 | 409976 | 0.002
409922 | 0.002 | 409949 | 0.001 | 409904 | 0.002 | 410007 | 0.003 | 410096 | 0.002
410042 | 0.002 | 410070 | 0.001 | 410023 | 0.002 | 410127 | 0.003 | 410216 | 0.002
410162 | 0.002 | 410189 | 0.001 | 410143 | 0.002 | 410247 | 0.003 | 410336 | 0.002
410281 | 0.002 | 410309 | 0.001 | 410263 | 0.002 | 410366 | 0.003 | 410456 | 0.002
410401 | 0.002 | 410429 | 0.001 | 410384 | 0.002 | 410486 | 0.003 | 410575 | 0.002
410521 | 0.001 | 410549 | 0.001 | 410504 | 0.002 | 410606 | 0.003 | 410695 | 0.002
410641 | 0.001 | 410669 | 0.001 | 410624 | 0.002 | 410727 | 0.003 | 410815 | 0.002
410761 | 0.002 | 410788 | 0.001 | 410743 | 0.002 | 410847 | 0.003 | 410935 | 0.002
410881 | 0.002 | 410908 | 0.001 | 410863 | 0.002 | 410967 | 0.003 | 411056 | 0.002
411000 | 0.002 | 411028 | 0.001 | 410983 | 0.002 | 411086 | 0.003 | 411175 | 0.002
411121 | 0.002 | 411148 | 0.001 | 411103 | 0.002 | 411206 | 0.003 | 411295 | 0.002
411241 | 0.002 | 411269 | 0.001 | 411223 | 0.002 | 411326 | 0.003 | 411415 | 0.002
411361 | 0.002 | 411388 | 0.001 | 411342 | 0.002 | 411446 | 0.003 | 411535 | 0.002
411481 | 0.002 | 411508 | 0.001 | 411462 | 0.002 | 411566 | 0.003 | 411655 | 0.002
411600 | 0.002 | 411628 | 0.001 | 411583 | 0.002 | 411685 | 0.003 | 411774 | 0.002
411720 | 0.002 | 411748 | 0.001 | 411703 | 0.002 | 411805 | 0.003 | 411894 | 0.002
411840 | 0.002 | 411868 | 0.001 | 411823 | 0.002 | 411925 | 0.003 | 412014 | 0.002
411960 | 0.002 | 411987 | 0.001 | 411942 | 0.002 | 412046 | 0.003 | 412134 | 0.002
412080 | 0.002 | 412107 | 0.001 | 412062 | 0.002 | 412166 | 0.003 | 412255 | 0.002
412199 | 0.002 | 412227 | 0.001 | 412182 | 0.002 | 412285 | 0.003 | 412375 | 0.002
412320 | 0.002 | 412347 | 0.001 | 412302 | 0.002 | 412405 | 0.003 | 412494 | 0.002
412440 | 0.002 | 412468 | 0.001 | 412422 | 0.002 | 412525 | 0.003 | 412614 | 0.002
412560 | 0.002 | 412588 | 0.001 | 412541 | 0.002 | 412645 | 0.003 | 412734 | 0.002
412680 | 0.002 | 412707 | 0.001 | 412661 | 0.002 | 412765 | 0.003 | 412854 | 0.002
412800 | 0.002 | 412827 | 0.001 | 412781 | 0.002 | 412884 | 0.003 | 412974 | 0.002
412919 | 0.002 | 412947 | 0.001 | 412902 | 0.002 | 413004 | 0.003 | 413093 | 0.002
413039 | 0.002 | 413067 | 0.001 | 413022 | 0.002 | 413124 | 0.003 | 413213 | 0.002
413159 | 0.002 | 413187 | 0.001 | 413142 | 0.002 | 413245 | 0.003 | 413333 | 0.002
413279 | 0.002 | 413306 | 0.001 | 413261 | 0.002 | 413365 | 0.003 | 413454 | 0.002
413399 | 0.002 | 413426 | 0.001 | 413384 | 0.002 | 413485 | 0.003 | 413574 | 0.002
413519 | 0.002 | 413546 | 0.001 | 413504 | 0.002 | 413604 | 0.003 | 413693 | 0.002
413639 | 0.002 | 413667 | 0.001 | 413625 | 0.002 | 413724 | 0.002 | 413813 | 0.002
413762 | 0.002 | 413787 | 0.001 | 413745 | 0.002 | 413844 | 0.003 | 413936 | 0.002
413882 | 0.002 | 413906 | 0.001 | 413864 | 0.002 | 413964 | 0.003 | 414056 | 0.002
414002 | 0.002 | 414026 | 0.001 | 413984 | 0.002 | 414088 | 0.003 | 414176 | 0.002
414121 | 0.002 | 414149 | 0.001 | 414104 | 0.002 | 414208 | 0.003 | 414318 | 0.002
414258 | 0.002 | 414269 | 0.001 | 414224 | 0.002 | 414327 | 0.003 | 414438 | 0.002
414377 | 0.002 | 414389 | 0.001 | 414344 | 0.002 | 414447 | 0.002 | 414558 | 0.002
414497 | 0.002 | 414510 | 0.001 | 414464 | 0.002 | 414567 | 0.002 | 414678 | 0.002
414618 | 0.002 | 414629 | 0.001 | 414583 | 0.002 | 414687 | 0.003 | 414797 | 0.002
414738 | 0.001 | 414749 | 0.001 | 414703 | 0.002 | 414807 | 0.003 | 414917 | 0.002
414858 | 0.002 | 414869 | 0.001 | 414823 | 0.002 | 414926 | 0.003 | 415037 | 0.002
414978 | 0.002 | 414989 | 0.001 | 414944 | 0.002 | 415046 | 0.002 | 415158 | 0.002
415097 | 0.002 | 415109 | 0.001 | 415064 | 0.002 | 415166 | 0.002 | 415278 | 0.002
415217 | 0.002 | 415228 | 0.001 | 415183 | 0.002 | 415286 | 0.003 | 415397 | 0.002
415337 | 0.002 | 415348 | 0.001 | 415303 | 0.002 | 415407 | 0.003 | 415517 | 0.002
415457 | 0.002 | 415468 | 0.001 | 415423 | 0.002 | 415526 | 0.002 | 415637 | 0.002
415577 | 0.002 | 415589 | 0.001 | 415543 | 0.002 | 415646 | 0.002 | 415757 | 0.002
415696 | 0.002 | 415709 | 0.001 | 415663 | 0.002 | 415766 | 0.003 | 415877 | 0.002
415817 | 0.002 | 415828 | 0.001 | 415782 | 0.002 | 415886 | 0.002 | 415996 | 0.002
415937 | 0.001 | 415948 | 0.001 | 415902 | 0.002 | 416006 | 0.002 | 416116 | 0.002
416057 | 0.002 | 416068 | 0.001 | 416022 | 0.002 | 416125 | 0.002 | 416236 | 0.002
416177 | 0.002 | 416188 | 0.001 | 416143 | 0.002 | 416245 | 0.002 | 416357 | 0.002
416296 | 0.002 | 416308 | 0.001 | 416263 | 0.002 | 416365 | 0.002 | 416477 | 0.002
416416 | 0.002 | 416428 | 0.001 | 416382 | 0.002 | 416485 | 0.002 | 416596 | 0.002
416536 | 0.002 | 416547 | 0.001 | 416502 | 0.002 | 416605 | 0.002 | 416716 | 0.002
416656 | 0.001 | 416667 | 0.001 | 416622 | 0.002 | 416726 | 0.002 | 416836 | 0.002
416776 | 0.002 | 416788 | 0.001 | 416742 | 0.002 | 416845 | 0.002 | 416956 | 0.002
416896 | 0.002 | 416908 | 0.001 | 416880 | 0.002 | 416965 | 0.002 | 417076 | 0.002
417015 | 0.002 | 417028 | 0.001 | 417000 | 0.002 | 417085 | 0.002 | 417196 | 0.002
417136 | 0.001 | 417147 | 0.001 | 417121 | 0.002 | 417205 | 0.002 | 417315 | 0.002
417256 | 0.001 | 417267 | 0.001 | 417241 | 0.002 | 417325 | 0.003 | 417435 | 0.002
417376 | 0.001 | 417387 | 0.001 | 417360 | 0.002 | 417444 | 0.002 | 417555 | 0.002
417496 | 0.002 | 417507 | 0.001 | 417480 | 0.002 | 417564 | 0.003 | 417676 | 0.002
417615 | 0.002 | 417627 | 0.001 | 417600 | 0.002 | 417684 | 0.002 | 417796 | 0.002
417735 | 0.001 | 417746 | 0.001 | 417720 | 0.002 | 417804 | 0.002 | 417915 | 0.002
417855 | 0.001 | 417866 | 0.001 | 417843 | 0.002 | 417925 | 0.002 | 418035 | 0.002
417975 | 0.001 | 417987 | 0.001 | 417963 | 0.002 | 418044 | 0.003 | 418155 | 0.002
418099 | 0.002 | 418107 | 0.001 | 418082 | 0.002 | 418164 | 0.003 | 418278 | 0.002
418217 | 0.002 | 418227 | 0.001 | 418202 | 0.002 | 418284 | 0.003 | 418398 | 0.002
418337 | 0.002 | 418347 | 0.001 | 418323 | 0.002 | 418404 | 0.003 | 418518 | 0.002
418457 | 0.002 | 418466 | 0.001 | 418443 | 0.002 | 418527 | 0.003 | 418637 | 0.002
418577 | 0.002 | 418589 | 0.001 | 418563 | 0.002 | 418647 | 0.003 | 418758 | 0.002
418697 | 0.002 | 418709 | 0.001 | 418682 | 0.002 | 418766 | 0.003 | 418878 | 0.002
418817 | 0.002 | 418829 | 0.001 | 418802 | 0.002 | 418886 | 0.003 | 418998 | 0.002
418936 | 0.002 | 418949 | 0.001 | 418922 | 0.002 | 419006 | 0.003 | 419118 | 0.002
419056 | 0.001 | 419069 | 0.001 | 419042 | 0.002 | 419126 | 0.003 | 419237 | 0.002
419176 | 0.002 | 419189 | 0.001 | 419162 | 0.002 | 419246 | 0.002 | 419357 | 0.002
419296 | 0.002 | 419309 | 0.001 | 419281 | 0.002 | 419365 | 0.002 | 419477 | 0.002
419416 | 0.002 | 419429 | 0.001 | 419401 | 0.002 | 419484 | 0.003 | 419597 | 0.002
419535 | 0.002 | 419549 | 0.001 | 419521 | 0.002 | 419604 | 0.002 | 419717 | 0.002
419655 | 0.002 | 419668 | 0.001 | 419641 | 0.002 | 419724 | 0.002 | 419836 | 0.002
419774 | 0.002 | 419788 | 0.001 | 419761 | 0.002 | 419844 | 0.002 | 419956 | 0.002
419894 | 0.002 | 419908 | 0.001 | 419880 | 0.002 | 419963 | 0.002 | 420076 | 0.002
420014 | 0.002 | 420028 | 0.001 | 420000 | 0.002 | 420083 | 0.002 | 420196 | 0.002
420133 | 0.002 | 420148 | 0.001 | 420120 | 0.002 | 420203 | 0.002 | 420316 | 0.002
420253 | 0.002 | 420268 | 0.001 | 420240 | 0.002 | 420323 | 0.002 | 420435 | 0.002
420373 | 0.002 | 420387 | 0.001 | 420360 | 0.002 | 420443 | 0.003 | 420555 | 0.002
420493 | 0.002 | 420507 | 0.001 | 420480 | 0.002 | 420562 | 0.002 | 420675 | 0.002
420613 | 0.002 | 420627 | 0.001 | 420599 | 0.002 | 420682 | 0.003 | 420795 | 0.002
420732 | 0.002 | 420747 | 0.001 | 420719 | 0.002 | 420802 | 0.002 | 420915 | 0.002
420852 | 0.002 | 420867 | 0.001 | 420840 | 0.002 | 420921 | 0.002 | 421035 | 0.002
420972 | 0.002 | 420986 | 0.001 | 420960 | 0.002 | 421041 | 0.002 | 421154 | 0.002
421092 | 0.002 | 421106 | 0.001 | 421080 | 0.002 | 421160 | 0.002 | 421274 | 0.002
421212 | 0.002 | 421226 | 0.001 | 421199 | 0.002 | 421280 | 0.002 | 421395 | 0.002
421330 | 0.002 | 421346 | 0.001 | 421319 | 0.002 | 421400 | 0.002 | 421515 | 0.002
421450 | 0.002 | 421466 | 0.001 | 421439 | 0.002 | 421520 | 0.002 | 421635 | 0.002
421570 | 0.002 | 421585 | 0.001 | 421559 | 0.002 | 421640 | 0.002 | 421754 | 0.002
|  | 421706 | 0.001 | 421679 | 0.002 | 421760 | 0.002 | 421874 | 0.002
|  |  |  | 421798 | 0.002 | 421879 | 0.002 | 421994 | 0.002
|  |  |  | 421918 | 0.002 | 421999 | 0.002 | 422114 | 0.002
|  |  |  | 422038 | 0.002 | 422119 | 0.002 | 422234 | 0.002
|  |  |  | 422161 | 0.002 | 422239 | 0.002 | 422353 | 0.002
|  |  |  | 422281 | 0.002 | 422378 | 0.002 | 422473 | 0.002
|  |  |  | 422401 | 0.002 | 422498 | 0.002 | 422593 | 0.002
|  |  |  | 422520 | 0.002 | 422618 | 0.006 | 422716 | 0.002
|  |  |  | 422640 | 0.002 | 422737 | 0.01 | 422836 | 0.002
|  |  |  | 422760 | 0.002 | 422860 | 0.014
|  |  |  | 422880 | 0.002 | 422980 | 0.017
|  |  |  | 423000 | 0.002
|  |  |  | 423119 | 0.002
|  |  |  | 423239 | 0.002
|  |  |  | 423358 | 0.002
|  |  |  | 423478 | 0.002
|  |  |  | 423598 | 0.002
|  |  |  | 423717 | 0.002
|  |  |  | 423837 | 0.002
|  |  |  | 423957 | 0.002
|  |  |  | 424077 | 0.002
|  |  |  | 424197 | 0.002
|  |  |  | 424316 | 0.002
|  |  |  | 424436 | 0.002
|  |  |  | 424556 | 0.002
|  |  |  | 424676 | 0.002
|  |  |  | 424796 | 0.002
|  |  |  | 424914 | 0.002
|  |  |  | 425034 | 0.002
|  |  |  | 425154 | 0.002
|  |  |  | 425274 | 0.002
|  |  |  | 425394 | 0.002
|  |  |  | 425513 | 0.002

## Zero Offset
Test 1 |  | Test 2 |  | Test 3 |  | Test 4 |  | Test 5
TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg) | TimeStamp (ms) | ScaleWeight (kg)
760 | 0.002 | 817 | 0.002 | 843 | 0.002 | 811 | 0.002 | 829 | 0.002
944 | 0.002 | 999 | 0.002 | 1026 | 0.002 | 994 | 0.002 | 1012 | 0.002
1061 | 0.002 | 1117 | 0.002 | 1143 | 0.002 | 1111 | 0.002 | 1129 | 0.002
1178 | 0.002 | 1233 | 0.002 | 1260 | 0.002 | 1228 | 0.002 | 1247 | 0.002
1296 | 0.002 | 1351 | 0.002 | 1378 | 0.002 | 1345 | 0.002 | 1364 | 0.002
1413 | 0.002 | 1468 | 0.002 | 1495 | 0.002 | 1463 | 0.002 | 1481 | 0.002
1530 | 0.002 | 1586 | 0.002 | 1612 | 0.002 | 1580 | 0.002 | 1599 | 0.002
1647 | 0.002 | 1702 | 0.002 | 1729 | 0.002 | 1697 | 0.002 | 1716 | 0.002
1765 | 0.002 | 1820 | 0.002 | 1847 | 0.002 | 1814 | 0.002 | 1833 | 0.002
1882 | 0.002 | 1937 | 0.002 | 1964 | 0.002 | 1932 | 0.002 | 1950 | 0.002
1999 | 0.002 | 2055 | 0.002 | 2081 | 0.002 | 2049 | 0.002 | 2068 | 0.002
2116 | 0.002 | 2171 | 0.002 | 2199 | 0.001 | 2166 | 0.002 | 2185 | 0.002
2234 | 0.002 | 2289 | 0.002 | 2316 | 0.002 | 2283 | 0.002 | 2302 | 0.002
2351 | 0.002 | 2407 | 0.002 | 2434 | 0.002 | 2401 | 0.002 | 2419 | 0.002
2468 | 0.002 | 2524 | 0.002 | 2550 | 0.002 | 2518 | 0.002 | 2537 | 0.002
2589 | 0.002 | 2641 | 0.002 | 2668 | 0.002 | 2635 | 0.002 | 2654 | 0.002
2706 | 0.002 | 2758 | 0.002 | 2785 | 0.002 | 2753 | 0.002 | 2771 | 0.002
2824 | 0.002 | 2876 | 0.002 | 2906 | 0.002 | 2870 | 0.002 | 2888 | 0.002
2941 | 0.002 | 2993 | 0.002 | 3023 | 0.002 | 2988 | 0.002 | 3009 | 0.002
3058 | 0.002 | 3110 | 0.002 | 3140 | 0.001 | 3104 | 0.002 | 3127 | 0.002
3176 | 0.002 | 3227 | 0.002 | 3258 | 0.001 | 3222 | 0.002 | 3244 | 0.002
3293 | 0.002 | 3345 | 0.002 | 3375 | 0.001 | 3342 | 0.002 | 3361 | 0.002
3410 | 0.002 | 3462 | 0.002 | 3492 | 0.001 | 3460 | 0.002 | 3479 | 0.002
3528 | 0.002 | 3579 | 0.002 | 3610 | 0.002 | 3577 | 0.002 | 3596 | 0.002
3645 | 0.002 | 3696 | 0.002 | 3727 | 0.002 | 3694 | 0.002 | 3714 | 0.002
3763 | 0.002 | 3814 | 0.002 | 3845 | 0.001 | 3812 | 0.002 | 3830 | 0.002
3880 | 0.002 | 3931 | 0.002 | 3962 | 0.001 | 3930 | 0.002 | 3948 | 0.002
3997 | 0.002 | 4051 | 0.002 | 4079 | 0.001 | 4046 | 0.002 | 4066 | 0.002
4115 | 0.002 | 4169 | 0.002 | 4197 | 0.002 | 4164 | 0.002 | 4183 | 0.002
4232 | 0.002 | 4286 | 0.002 | 4314 | 0.001 | 4281 | 0.002 | 4300 | 0.002
4349 | 0.002 | 4404 | 0.002 | 4431 | 0.001 | 4399 | 0.002 | 4418 | 0.002
4467 | 0.002 | 4521 | 0.002 | 4549 | 0.002 | 4516 | 0.002 | 4535 | 0.002
4584 | 0.002 | 4638 | 0.002 | 4666 | 0.002 | 4633 | 0.002 | 4653 | 0.002
4702 | 0.002 | 4756 | 0.002 | 4784 | 0.002 | 4751 | 0.002 | 4770 | 0.002
4819 | 0.002 | 4874 | 0.002 | 4901 | 0.001 | 4869 | 0.002 | 4887 | 0.002
4936 | 0.002 | 4990 | 0.002 | 5018 | 0.001 | 4985 | 0.002 | 5005 | 0.002
5054 | 0.002 | 5108 | 0.002 | 5136 | 0.001 | 5103 | 0.002 | 5122 | 0.002
5172 | 0.002 | 5225 | 0.002 | 5253 | 0.001 | 5220 | 0.002 | 5239 | 0.002
5288 | 0.002 | 5343 | 0.002 | 5370 | 0.002 | 5338 | 0.002 | 5357 | 0.002
5406 | 0.002 | 5460 | 0.002 | 5488 | 0.001 | 5455 | 0.002 | 5474 | 0.002
5523 | 0.002 | 5577 | 0.002 | 5605 | 0.001 | 5572 | 0.002 | 5592 | 0.002
5641 | 0.002 | 5695 | 0.002 | 5723 | 0.001 | 5690 | 0.002 | 5709 | 0.002
5758 | 0.002 | 5813 | 0.002 | 5840 | 0.001 | 5808 | 0.002 | 5826 | 0.002
5875 | 0.002 | 5929 | 0.002 | 5957 | 0.001 | 5924 | 0.002 | 5944 | 0.002
5993 | 0.002 | 6047 | 0.002 | 6075 | 0.001 | 6042 | 0.002 | 6061 | 0.002
6111 | 0.002 | 6164 | 0.002 | 6192 | 0.001 | 6159 | 0.002 | 6178 | 0.002
6227 | 0.002 | 6282 | 0.002 | 6309 | 0.001 | 6277 | 0.002 | 6296 | 0.002
6345 | 0.002 | 6400 | 0.002 | 6427 | 0.001 | 6394 | 0.002 | 6413 | 0.002
6462 | 0.002 | 6516 | 0.002 | 6544 | 0.001 | 6511 | 0.002 | 6531 | 0.002
6580 | 0.002 | 6634 | 0.002 | 6662 | 0.001 | 6629 | 0.002 | 6648 | 0.002
6697 | 0.002 | 6752 | 0.002 | 6779 | 0.001 | 6747 | 0.002 | 6765 | 0.002
6814 | 0.002 | 6868 | 0.002 | 6896 | 0.002 | 6863 | 0.002 | 6883 | 0.002
6932 | 0.002 | 6986 | 0.002 | 7014 | 0.001 | 6981 | 0.002 | 7000 | 0.002
7050 | 0.002 | 7103 | 0.002 | 7131 | 0.001 | 7098 | 0.002 | 7117 | 0.002
7166 | 0.002 | 7221 | 0.002 | 7248 | 0.001 | 7216 | 0.002 | 7235 | 0.002
7284 | 0.002 | 7339 | 0.002 | 7366 | 0.002 | 7333 | 0.002 | 7352 | 0.002
7401 | 0.002 | 7455 | 0.002 | 7483 | 0.001 | 7450 | 0.002 | 7470 | 0.002
7519 | 0.002 | 7573 | 0.002 | 7601 | 0.002 | 7568 | 0.002 | 7587 | 0.002
7640 | 0.002 | 7691 | 0.002 | 7718 | 0.001 | 7686 | 0.002 | 7704 | 0.002
7757 | 0.002 | 7808 | 0.002 | 7835 | 0.001 | 7802 | 0.002 | 7822 | 0.002
7874 | 0.002 | 7925 | 0.002 | 7956 | 0.002 | 7920 | 0.002 | 7939 | 0.002
7992 | 0.002 | 8042 | 0.002 | 8074 | 0.001 | 8038 | 0.002 | 8059 | 0.002
8109 | 0.002 | 8160 | 0.002 | 8190 | 0.001 | 8155 | 0.002 | 8177 | 0.002
8226 | 0.002 | 8278 | 0.002 | 8308 | 0.002 | 8272 | 0.002 | 8295 | 0.002
8344 | 0.002 | 8394 | 0.002 | 8426 | 0.001 | 8393 | 0.002 | 8412 | 0.002
8461 | 0.002 | 8512 | 0.002 | 8543 | 0.001 | 8510 | 0.002 | 8529 | 0.002
8579 | 0.002 | 8630 | 0.002 | 8660 | 0.001 | 8628 | 0.002 | 8646 | 0.002
8696 | 0.002 | 8747 | 0.002 | 8777 | 0.001 | 8745 | 0.001 | 8764 | 0.002
8813 | 0.002 | 8864 | 0.002 | 8895 | 0.002 | 8862 | 0.002 | 8882 | 0.002
8931 | 0.002 | 8981 | 0.002 | 9013 | 0.002 | 8980 | 0.002 | 8998 | 0.002
9048 | 0.002 | 9102 | 0.002 | 9129 | 0.002 | 9097 | 0.002 | 9116 | 0.002
9165 | 0.002 | 9220 | 0.002 | 9247 | 0.001 | 9214 | 0.002 | 9234 | 0.002
9316 | 0.002 | 9337 | 0.002 | 9365 | 0.001 | 9332 | 0.002 | 9351 | 0.002
9434 | 0.002 | 9454 | 0.002 | 9516 | 0.001 | 9449 | 0.002 | 9503 | 0.002
9550 | 0.002 | 9572 | 0.002 | 9633 | 0.002 | 9567 | 0.002 | 9620 | 0.002
9668 | 0.002 | 9689 | 0.002 | 9751 | 0.001 | 9684 | 0.002 | 9738 | 0.002
9785 | 0.002 | 9806 | 0.002 | 9868 | 0.001 | 9801 | 0.002 | 9856 | 0.002
9903 | 0.002 | 9924 | 0.002 | 9986 | 0.002 | 9953 | 0.002 | 9972 | 0.002
10020 | 0.002 | 10041 | 0.002 | 10102 | 0.002 | 10071 | 0.002 | 10090 | 0.002
10137 | 0.002 | 10159 | 0.002 | 10220 | 0.001 | 10187 | 0.002 | 10207 | 0.002
10255 | 0.002 | 10276 | 0.002 | 10338 | 0.001 | 10305 | 0.002 | 10324 | 0.002
10373 | 0.002 | 10393 | 0.002 | 10455 | 0.001 | 10423 | 0.002 | 10442 | 0.002
10489 | 0.002 | 10511 | 0.002 | 10572 | 0.002 | 10541 | 0.002 | 10560 | 0.002
10607 | 0.002 | 10629 | 0.002 | 10690 | 0.001 | 10657 | 0.002 | 10677 | 0.002
10725 | 0.002 | 10779 | 0.002 | 10808 | 0.002 | 10775 | 0.002 | 10795 | 0.002
10842 | 0.002 | 10897 | 0.002 | 10925 | 0.002 | 10893 | 0.002 | 10912 | 0.002
10959 | 0.002 | 11014 | 0.002 | 11042 | 0.002 | 11010 | 0.002 | 11029 | 0.002
11077 | 0.002 | 11131 | 0.002 | 11160 | 0.001 | 11127 | 0.002 | 11147 | 0.002
11194 | 0.002 | 11249 | 0.002 | 11277 | 0.001 | 11245 | 0.002 | 11265 | 0.002
11312 | 0.002 | 11366 | 0.002 | 11395 | 0.001 | 11362 | 0.002 | 11381 | 0.002
11429 | 0.002 | 11484 | 0.002 | 11512 | 0.002 | 11480 | 0.002 | 11499 | 0.002
11547 | 0.002 | 11601 | 0.002 | 11629 | 0.002 | 11597 | 0.002 | 11617 | 0.002
11664 | 0.002 | 11718 | 0.002 | 11747 | 0.002 | 11714 | 0.002 | 11734 | 0.002
11782 | 0.002 | 11836 | 0.003 | 11865 | 0.002 | 11832 | 0.002 | 11851 | 0.002
11899 | 0.002 | 11954 | 0.002 | 11981 | 0.001 | 11950 | 0.002 | 11969 | 0.002
12016 | 0.002 | 12070 | 0.002 | 12099 | 0.002 | 12066 | 0.002 | 12086 | 0.002
12134 | 0.002 | 12188 | 0.002 | 12217 | 0.001 | 12184 | 0.002 | 12204 | 0.002
12252 | 0.002 | 12306 | 0.002 | 12334 | 0.002 | 12302 | 0.002 | 12321 | 0.002
12368 | 0.002 | 12423 | 0.002 | 12451 | 0.002 | 12419 | 0.002 | 12438 | 0.002
12489 | 0.002 | 12540 | 0.002 | 12569 | 0.001 | 12536 | 0.002 | 12556 | 0.002
12607 | 0.002 | 12658 | 0.002 | 12686 | 0.001 | 12654 | 0.002 | 12674 | 0.002
12725 | 0.002 | 12775 | 0.002 | 12807 | 0.002 | 12771 | 0.002 | 12790 | 0.002
12841 | 0.002 | 12893 | 0.002 | 12924 | 0.002 | 12889 | 0.002 | 12911 | 0.002
12959 | 0.002 | 13010 | 0.002 | 13042 | 0.001 | 13006 | 0.002 | 13029 | 0.002
13077 | 0.002 | 13128 | 0.002 | 13159 | 0.001 | 13123 | 0.002 | 13147 | 0.002
13194 | 0.002 | 13245 | 0.002 | 13277 | 0.001 | 13244 | 0.002 | 13263 | 0.002
13312 | 0.002 | 13363 | 0.002 | 13394 | 0.001 | 13362 | 0.002 | 13381 | 0.002
13429 | 0.002 | 13480 | 0.002 | 13511 | 0.001 | 13479 | 0.002 | 13499 | 0.002
13546 | 0.002 | 13597 | 0.002 | 13629 | 0.001 | 13596 | 0.002 | 13616 | 0.002
13664 | 0.002 | 13715 | 0.002 | 13747 | 0.001 | 13714 | 0.002 | 13733 | 0.002
13782 | 0.002 | 13836 | 0.002 | 13863 | 0.001 | 13832 | 0.002 | 13851 | 0.002
13898 | 0.002 | 13953 | 0.002 | 13981 | 0.001 | 13948 | 0.002 | 13968 | 0.002
14016 | 0.002 | 14070 | 0.002 | 14099 | 0.001 | 14066 | 0.002 | 14086 | 0.002
14134 | 0.002 | 14188 | 0.002 | 14216 | 0.001 | 14184 | 0.002 | 14203 | 0.002
14251 | 0.002 | 14306 | 0.002 | 14333 | 0.001 | 14302 | 0.002 | 14320 | 0.002
14368 | 0.002 | 14423 | 0.002 | 14451 | 0.001 | 14418 | 0.002 | 14438 | 0.002
14486 | 0.002 | 14540 | 0.002 | 14568 | 0.001 | 14536 | 0.002 | 14556 | 0.002
14603 | 0.002 | 14658 | 0.002 | 14686 | 0.001 | 14654 | 0.002 | 14673 | 0.002
14721 | 0.002 | 14775 | 0.002 | 14803 | 0.001 | 14771 | 0.002 | 14790 | 0.002
14838 | 0.002 | 14893 | 0.002 | 14920 | 0.001 | 14888 | 0.002 | 14908 | 0.002
14955 | 0.002 | 15010 | 0.002 | 15038 | 0.001 | 15006 | 0.002 | 15026 | 0.002
15073 | 0.002 | 15127 | 0.002 | 15156 | 0.001 | 15123 | 0.002 | 15142 | 0.002
15191 | 0.002 | 15245 | 0.002 | 15272 | 0.001 | 15241 | 0.002 | 15260 | 0.002
15308 | 0.002 | 15363 | 0.002 | 15390 | 0.001 | 15358 | 0.002 | 15378 | 0.002
15425 | 0.002 | 15479 | 0.002 | 15508 | 0.001 | 15475 | 0.002 | 15495 | 0.002
15543 | 0.002 | 15597 | 0.002 | 15626 | 0.001 | 15593 | 0.002 | 15612 | 0.002
15661 | 0.002 | 15715 | 0.002 | 15742 | 0.001 | 15711 | 0.002 | 15730 | 0.002
15777 | 0.002 | 15832 | 0.002 | 15860 | 0.001 | 15827 | 0.002 | 15847 | 0.003
15895 | 0.002 | 15949 | 0.002 | 15978 | 0.001 | 15945 | 0.002 | 15965 | 0.002
16013 | 0.002 | 16067 | 0.002 | 16095 | 0.001 | 16063 | 0.002 | 16082 | 0.002
16130 | 0.002 | 16184 | 0.002 | 16212 | 0.001 | 16180 | 0.002 | 16199 | 0.003
16247 | 0.002 | 16302 | 0.002 | 16330 | 0.001 | 16297 | 0.002 | 16317 | 0.002
16365 | 0.002 | 16419 | 0.002 | 16447 | 0.001 | 16415 | 0.002 | 16435 | 0.002
16482 | 0.002 | 16536 | 0.002 | 16565 | 0.001 | 16532 | 0.002 | 16551 | 0.002
16600 | 0.002 | 16654 | 0.002 | 16683 | 0.001 | 16650 | 0.002 | 16669 | 0.002
16717 | 0.002 | 16772 | 0.002 | 16799 | 0.001 | 16768 | 0.002 | 16787 | 0.002
16834 | 0.002 | 16888 | 0.002 | 16917 | 0.001 | 16884 | 0.002 | 16904 | 0.003
16952 | 0.002 | 17006 | 0.002 | 17035 | 0.001 | 17002 | 0.002 | 17021 | 0.002
17073 | 0.002 | 17124 | 0.002 | 17152 | 0.001 | 17120 | 0.002 | 17139 | 0.002
17190 | 0.002 | 17241 | 0.002 | 17269 | 0.001 | 17238 | 0.002 | 17256 | 0.002
17307 | 0.002 | 17358 | 0.002 | 17390 | 0.001 | 17354 | 0.002 | 17374 | 0.002
17425 | 0.002 | 17476 | 0.002 | 17508 | 0.001 | 17472 | 0.002 | 17495 | 0.002
17543 | 0.002 | 17593 | 0.002 | 17625 | 0.001 | 17590 | 0.002 | 17612 | 0.002
17659 | 0.002 | 17711 | 0.002 | 17742 | 0.001 | 17707 | 0.002 | 17729 | 0.002
17777 | 0.002 | 17828 | 0.002 | 17860 | 0.001 | 17827 | 0.002 | 17847 | 0.002
17895 | 0.002 | 17946 | 0.002 | 17977 | 0.001 | 17945 | 0.002 | 17965 | 0.002
18012 | 0.002 | 18063 | 0.002 | 18095 | 0.001 | 18063 | 0.002 | 18081 | 0.002
18129 | 0.003 | 18181 | 0.002 | 18212 | 0.001 | 18180 | 0.002 | 18199 | 0.002
18247 | 0.003 | 18298 | 0.002 | 18329 | 0.001 | 18297 | 0.002 | 18317 | 0.002
18364 | 0.002 | 18415 | 0.002 | 18447 | 0.001 | 18415 | 0.002 | 18434 | 0.002
18482 | 0.003 | 18531 | 0.002 | 18565 | 0.001 | 18532 | 0.002 | 18551 | 0.003
18599 | 0.003 | 18648 | 0.002 | 18681 | 0.001 | 18650 | 0.002 | 18669 | 0.002
18716 | 0.003 | 18765 | 0.002 | 18799 | 0.001 | 18767 | 0.002 | 18787 | 0.002
18834 | 0.003 | 18883 | 0.002 | 18917 | 0.001 | 18884 | 0.002 | 18904 | 0.002
18952 | 0.003 | 19000 | 0.002 | 19034 | 0.001 | 19002 | 0.002 | 19021 | 0.002
19068 | 0.003 | 19118 | 0.002 | 19151 | 0.001 | 19120 | 0.002 | 19139 | 0.002
19186 | 0.003 | 19235 | 0.002 | 19269 | 0.001 | 19236 | 0.002 | 19256 | 0.002
19304 | 0.003 | 19352 | 0.002 | 19386 | 0.001 | 19354 | 0.002 | 19374 | 0.002
19422 | 0.003 | 19470 | 0.002 | 19504 | 0.001 | 19472 | 0.002 | 19491 | 0.002
19538 | 0.003 | 19588 | 0.002 | 19621 | 0.001 | 19589 | 0.002 | 19608 | 0.002
19656 | 0.003 | 19704 | 0.002 | 19738 | 0.001 | 19706 | 0.002 | 19726 | 0.002
19774 | 0.003 | 19822 | 0.002 | 19856 | 0.001 | 19824 | 0.002 | 19844 | 0.002
19891 | 0.003 | 19940 | 0.002 | 19974 | 0.001 | 19941 | 0.002 | 19960 | 0.002
20008 | 0.003 | 20057 | 0.002 | 20090 | 0.001 | 20059 | 0.002 | 20113 | 0.002
20159 | 0.003 | 20174 | 0.002 | 20242 | 0.002 | 20176 | 0.002 | 20230 | 0.002
20277 | 0.003 | 20292 | 0.002 | 20360 | 0.001 | 20293 | 0.002 | 20347 | 0.002
20393 | 0.003 | 20409 | 0.002 | 20477 | 0.001 | 20411 | 0.002 | 20465 | 0.002
20511 | 0.003 | 20527 | 0.002 | 20594 | 0.001 | 20562 | 0.002 | 20582 | 0.002
20629 | 0.003 | 20644 | 0.002 | 20712 | 0.001 | 20680 | 0.002 | 20700 | 0.002
20746 | 0.003 | 20762 | 0.002 | 20830 | 0.001 | 20797 | 0.002 | 20817 | 0.002
20864 | 0.003 | 20879 | 0.002 | 20946 | 0.001 | 20915 | 0.002 | 20935 | 0.002
20981 | 0.003 | 20997 | 0.002 | 21064 | 0.001 | 21032 | 0.002 | 21052 | 0.002
21098 | 0.003 | 21114 | 0.002 | 21182 | 0.001 | 21149 | 0.002 | 21170 | 0.002
21216 | 0.003 | 21231 | 0.002 | 21299 | 0.001 | 21267 | 0.002 | 21287 | 0.002
21334 | 0.003 | 21349 | 0.002 | 21416 | 0.001 | 21385 | 0.002 | 21404 | 0.002
21450 | 0.003 | 21467 | 0.002 | 21534 | 0.001 | 21502 | 0.002 | 21522 | 0.002
21568 | 0.003 | 21617 | 0.002 | 21651 | 0.001 | 21619 | 0.002 | 21640 | 0.002
21686 | 0.003 | 21734 | 0.002 | 21769 | 0.001 | 21737 | 0.002 | 21756 | 0.002
21807 | 0.003 | 21852 | 0.003 | 21886 | 0.001 | 21855 | 0.002 | 21874 | 0.002
21923 | 0.003 | 21969 | 0.003 | 22003 | 0.001 | 21971 | 0.002 | 21992 | 0.002
22041 | 0.003 | 22086 | 0.003 | 22124 | 0.002 | 22089 | 0.002 | 22109 | 0.002
22159 | 0.003 | 22204 | 0.003 | 22242 | 0.002 | 22207 | 0.002 | 22224 | 0.002
22276 | 0.003 | 22322 | 0.002 | 22360 | 0.002 | 22324 | 0.002 | 22341 | 0.002
22393 | 0.003 | 22439 | 0.002 | 22476 | 0.002 | 22445 | 0.002 | 22459 | 0.002
22511 | 0.003 | 22556 | 0.002 | 22594 | 0.002 | 22562 | 0.002 | 22577 | 0.002
22628 | 0.003 | 22674 | 0.002 | 22712 | 0.002 | 22680 | 0.002 | 22693 | 0.002
22746 | 0.003 | 22792 | 0.002 | 22829 | 0.002 | 22797 | 0.002 | 22811 | 0.002
22863 | 0.003 | 22908 | 0.003 | 22946 | 0.002 | 22915 | 0.002 | 22929 | 0.002
22980 | 0.003 | 23026 | 0.003 | 23064 | 0.002 | 23032 | 0.002 | 23046 | 0.002
23098 | 0.003 | 23147 | 0.002 | 23181 | 0.002 | 23149 | 0.002 | 23163 | 0.002
23216 | 0.003 | 23264 | 0.003 | 23299 | 0.002 | 23267 | 0.002 | 23281 | 0.002
23332 | 0.003 | 23382 | 0.003 | 23416 | 0.002 | 23385 | 0.002 | 23398 | 0.002
23450 | 0.003 | 23499 | 0.003 | 23533 | 0.002 | 23501 | 0.002 | 23516 | 0.002
23568 | 0.003 | 23617 | 0.003 | 23651 | 0.002 | 23619 | 0.002 | 23633 | 0.002
23686 | 0.003 | 23734 | 0.003 | 23769 | 0.002 | 23737 | 0.002 | 23751 | 0.002
23802 | 0.003 | 23852 | 0.003 | 23885 | 0.002 | 23854 | 0.002 | 23868 | 0.002
23920 | 0.003 | 23969 | 0.003 | 24003 | 0.002 | 23971 | 0.002 | 23986 | 0.002
24038 | 0.003 | 24086 | 0.002 | 24121 | 0.002 | 24089 | 0.002 | 24103 | 0.002
24155 | 0.003 | 24204 | 0.003 | 24238 | 0.002 | 24206 | 0.002 | 24220 | 0.002
24272 | 0.003 | 24322 | 0.003 | 24355 | 0.002 | 24324 | 0.002 | 24338 | 0.002
24390 | 0.003 | 24438 | 0.003 | 24473 | 0.002 | 24441 | 0.002 | 24456 | 0.002
24507 | 0.003 | 24556 | 0.003 | 24590 | 0.002 | 24558 | 0.002 | 24572 | 0.002
24625 | 0.003 | 24674 | 0.003 | 24708 | 0.002 | 24676 | 0.002 | 24690 | 0.002
24742 | 0.003 | 24791 | 0.003 | 24825 | 0.002 | 24794 | 0.002 | 24808 | 0.002
24859 | 0.003 | 24908 | 0.003 | 24943 | 0.002 | 24911 | 0.002 | 24925 | 0.002
24977 | 0.003 | 25026 | 0.003 | 25060 | 0.002 | 25028 | 0.001 | 25042 | 0.002
25095 | 0.003 | 25143 | 0.003 | 25178 | 0.002 | 25146 | 0.001 | 25160 | 0.002
25211 | 0.003 | 25261 | 0.003 | 25295 | 0.002 | 25264 | 0.002 | 25277 | 0.002
25329 | 0.003 | 25378 | 0.003 | 25412 | 0.002 | 25380 | 0.001 | 25395 | 0.002
25447 | 0.003 | 25495 | 0.002 | 25530 | 0.002 | 25498 | 0.001 | 25512 | 0.002
25564 | 0.003 | 25613 | 0.003 | 25648 | 0.002 | 25616 | 0.001 | 25629 | 0.002
25681 | 0.003 | 25731 | 0.003 | 25764 | 0.002 | 25733 | 0.001 | 25747 | 0.002
25799 | 0.003 | 25847 | 0.003 | 25882 | 0.002 | 25850 | 0.001 | 25865 | 0.002
25916 | 0.003 | 25965 | 0.003 | 26000 | 0.002 | 25968 | 0.001 | 25981 | 0.002
26034 | 0.003 | 26083 | 0.003 | 26117 | 0.002 | 26085 | 0.002 | 26099 | 0.002
26151 | 0.003 | 26200 | 0.002 | 26234 | 0.002 | 26203 | 0.001 | 26217 | 0.002
26269 | 0.003 | 26317 | 0.002 | 26352 | 0.002 | 26320 | 0.001 | 26334 | 0.002
26390 | 0.003 | 26435 | 0.003 | 26469 | 0.002 | 26437 | 0.001 | 26451 | 0.002
26507 | 0.003 | 26552 | 0.002 | 26587 | 0.002 | 26555 | 0.001 | 26569 | 0.002
26625 | 0.003 | 26670 | 0.002 | 26707 | 0.002 | 26673 | 0.001 | 26687 | 0.002
26753 | 0.003 | 26787 | 0.003 | 26825 | 0.002 | 26789 | 0.001 | 26808 | 0.002
26870 | 0.003 | 26905 | 0.003 | 26943 | 0.002 | 26907 | 0.001 | 26925 | 0.002
26987 | 0.003 | 27022 | 0.002 | 27060 | 0.002 | 27025 | 0.001 | 27042 | 0.002
27105 | 0.003 | 27140 | 0.002 | 27188 | 0.002 | 27146 | 0.001 | 27160 | 0.002
27223 | 0.003 | 27257 | 0.002 | 27304 | 0.002 | 27264 | 0.001 | 27287 | 0.002
27340 | 0.003 | 27374 | 0.002 | 27422 | 0.002 | 27380 | 0.001 | 27405 | 0.002
27457 | 0.003 | 27492 | 0.002 | 27540 | 0.002 | 27498 | 0.001 | 27523 | 0.002
27575 | 0.003 | 27610 | 0.002 | 27658 | 0.002 | 27626 | 0.001 | 27639 | 0.002
27693 | 0.003 | 27730 | 0.002 | 27774 | 0.002 | 27743 | 0.001 | 27757 | 0.002
27809 | 0.003 | 27847 | 0.002 | 27892 | 0.002 | 27860 | 0.001 | 27875 | 0.002
27927 | 0.003 | 27965 | 0.002 | 28010 | 0.002 | 27978 | 0.001 | 27992 | 0.002
28045 | 0.003 | 28083 | 0.002 | 28127 | 0.002 | 28095 | 0.001 | 28109 | 0.002
28162 | 0.003 | 28199 | 0.002 | 28244 | 0.002 | 28213 | 0.001 | 28227 | 0.002
28279 | 0.003 | 28327 | 0.002 | 28362 | 0.002 | 28331 | 0.001 | 28344 | 0.002
28397 | 0.003 | 28445 | 0.002 | 28479 | 0.002 | 28447 | 0.001 | 28462 | 0.002
28514 | 0.003 | 28562 | 0.002 | 28597 | 0.002 | 28565 | 0.001 | 28579 | 0.002
28632 | 0.003 | 28680 | 0.002 | 28715 | 0.002 | 28683 | 0.001 | 28696 | 0.002
28749 | 0.003 | 28797 | 0.002 | 28831 | 0.002 | 28800 | 0.001 | 28814 | 0.002
28866 | 0.003 | 28914 | 0.002 | 28949 | 0.002 | 28917 | 0.001 | 28932 | 0.002
28984 | 0.003 | 29032 | 0.002 | 29067 | 0.002 | 29035 | 0.001 | 29048 | 0.002
29102 | 0.003 | 29150 | 0.002 | 29184 | 0.002 | 29152 | 0.001 | 29166 | 0.002
29218 | 0.003 | 29267 | 0.002 | 29301 | 0.002 | 29270 | 0.002 | 29284 | 0.002
29336 | 0.003 | 29384 | 0.002 | 29419 | 0.002 | 29387 | 0.001 | 29402 | 0.002
29454 | 0.003 | 29502 | 0.002 | 29536 | 0.002 | 29504 | 0.001 | 29518 | 0.002
29571 | 0.003 | 29620 | 0.002 | 29654 | 0.002 | 29622 | 0.001 | 29636 | 0.002
29688 | 0.003 | 29736 | 0.002 | 29771 | 0.002 | 29740 | 0.001 | 29754 | 0.002
29806 | 0.003 | 29854 | 0.002 | 29888 | 0.002 | 29856 | 0.001 | 29871 | 0.002
29923 | 0.003 | 29972 | 0.002 | 30006 | 0.002 | 29974 | 0.001 | 29988 | 0.002
30041 | 0.003 | 30089 | 0.002 | 30124 | 0.002 | 30092 | 0.001 | 30106 | 0.002
30158 | 0.003 | 30206 | 0.003 | 30240 | 0.002 | 30209 | 0.001 | 30223 | 0.002
30276 | 0.003 | 30324 | 0.003 | 30358 | 0.002 | 30326 | 0.001 | 30341 | 0.002
30393 | 0.003 | 30441 | 0.002 | 30476 | 0.002 | 30444 | 0.001 | 30458 | 0.002
30511 | 0.003 | 30559 | 0.002 | 30593 | 0.002 | 30561 | 0.001 | 30575 | 0.002
30628 | 0.003 | 30676 | 0.002 | 30710 | 0.002 | 30679 | 0.001 | 30693 | 0.002
30745 | 0.003 | 30793 | 0.002 | 30828 | 0.002 | 30796 | 0.001 | 30844 | 0.002
30863 | 0.003 | 30911 | 0.002 | 30980 | 0.002 | 30914 | 0.001 | 30962 | 0.002
31013 | 0.003 | 31029 | 0.002 | 31096 | 0.002 | 31031 | 0.001 | 31080 | 0.002
31134 | 0.002 | 31145 | 0.002 | 31214 | 0.002 | 31149 | 0.001 | 31197 | 0.002
31251 | 0.003 | 31263 | 0.002 | 31335 | 0.002 | 31300 | 0.001 | 31314 | 0.002
31369 | 0.003 | 31381 | 0.002 | 31452 | 0.002 | 31417 | 0.001 | 31435 | 0.002
31486 | 0.003 | 31498 | 0.002 | 31569 | 0.002 | 31535 | 0.001 | 31553 | 0.002
31604 | 0.003 | 31616 | 0.002 | 31687 | 0.002 | 31652 | 0.001 | 31670 | 0.002
31721 | 0.003 | 31733 | 0.003 | 31805 | 0.002 | 31773 | 0.001 | 31788 | 0.002
31839 | 0.003 | 31850 | 0.002 | 31922 | 0.002 | 31890 | 0.001 | 31905 | 0.002
31956 | 0.003 | 31968 | 0.002 | 32039 | 0.002 | 32008 | 0.001 | 32022 | 0.002
32073 | 0.002 | 32086 | 0.002 | 32157 | 0.002 | 32125 | 0.001 | 32140 | 0.002
32191 | 0.003 | 32202 | 0.003 | 32274 | 0.002 | 32242 | 0.001 | 32258 | 0.002
32309 | 0.003 | 32354 | 0.002 | 32392 | 0.002 | 32360 | 0.001 | 32375 | 0.002
32425 | 0.002 | 32475 | 0.003 | 32509 | 0.002 | 32478 | 0.001 | 32492 | 0.002
32543 | 0.002 | 32591 | 0.002 | 32626 | 0.002 | 32595 | 0.001 | 32610 | 0.002
32661 | 0.002 | 32709 | 0.003 | 32744 | 0.002 | 32712 | 0.001 | 32728 | 0.002
32778 | 0.003 | 32827 | 0.003 | 32862 | 0.003 | 32830 | 0.001 | 32844 | 0.002
32896 | 0.003 | 32944 | 0.002 | 32978 | 0.003 | 32948 | 0.001 | 32962 | 0.002
33013 | 0.003 | 33061 | 0.003 | 33096 | 0.003 | 33064 | 0.001 | 33080 | 0.002
33130 | 0.002 | 33179 | 0.003 | 33214 | 0.002 | 33182 | 0.001 | 33197 | 0.002
33248 | 0.003 | 33296 | 0.002 | 33331 | 0.002 | 33300 | 0.001 | 33314 | 0.002
33366 | 0.003 | 33414 | 0.003 | 33448 | 0.002 | 33417 | 0.001 | 33432 | 0.002
33482 | 0.002 | 33531 | 0.003 | 33566 | 0.002 | 33534 | 0.001 | 33549 | 0.002
33600 | 0.002 | 33648 | 0.002 | 33683 | 0.003 | 33652 | 0.001 | 33667 | 0.002
33718 | 0.003 | 33766 | 0.003 | 33801 | 0.003 | 33769 | 0.001 | 33784 | 0.002
33835 | 0.002 | 33884 | 0.003 | 33918 | 0.002 | 33887 | 0.001 | 33901 | 0.002
33952 | 0.003 | 34000 | 0.003 | 34035 | 0.003 | 34004 | 0.001 | 34019 | 0.002
34070 | 0.003 | 34118 | 0.003 | 34153 | 0.003 | 34121 | 0.001 | 34137 | 0.002
34187 | 0.002 | 34236 | 0.003 | 34271 | 0.003 | 34239 | 0.001 | 34253 | 0.002
34305 | 0.002 | 34353 | 0.003 | 34387 | 0.002 | 34357 | 0.001 | 34371 | 0.002
34422 | 0.002 | 34470 | 0.003 | 34505 | 0.003 | 34473 | 0.001 | 34489 | 0.002
34539 | 0.003 | 34588 | 0.003 | 34623 | 0.003 | 34591 | 0.001 | 34606 | 0.002
34657 | 0.002 | 34705 | 0.003 | 34741 | 0.003 | 34709 | 0.001 | 34723 | 0.002
34775 | 0.002 | 34823 | 0.003 | 34857 | 0.003 | 34826 | 0.001 | 34841 | 0.002
34892 | 0.002 | 34940 | 0.003 | 34975 | 0.003 | 34944 | 0.001 | 34958 | 0.002
35009 | 0.002 | 35057 | 0.003 | 35093 | 0.003 | 35061 | 0.001 | 35076 | 0.002
35127 | 0.002 | 35175 | 0.003 | 35210 | 0.003 | 35178 | 0.002 | 35193 | 0.002
35245 | 0.002 | 35293 | 0.003 | 35328 | 0.003 | 35296 | 0.001 | 35310 | 0.002
35361 | 0.002 | 35410 | 0.003 | 35445 | 0.003 | 35414 | 0.001 | 35428 | 0.002
35479 | 0.002 | 35527 | 0.003 | 35562 | 0.003 | 35530 | 0.001 | 35546 | 0.002
35597 | 0.002 | 35645 | 0.003 | 35680 | 0.003 | 35648 | 0.002 | 35662 | 0.002
35717 | 0.002 | 35763 | 0.003 | 35798 | 0.003 | 35766 | 0.002 | 35780 | 0.002
35834 | 0.002 | 35879 | 0.003 | 35914 | 0.003 | 35884 | 0.002 | 35898 | 0.002
35952 | 0.002 | 35997 | 0.003 | 36035 | 0.003 | 36000 | 0.002 | 36016 | 0.002
36069 | 0.002 | 36115 | 0.003 | 36153 | 0.003 | 36118 | 0.002 | 36135 | 0.002
36187 | 0.002 | 36232 | 0.003 | 36271 | 0.003 | 36236 | 0.001 | 36253 | 0.002
36304 | 0.002 | 36349 | 0.003 | 36387 | 0.003 | 36353 | 0.001 | 36371 | 0.002
36422 | 0.002 | 36467 | 0.003 | 36505 | 0.003 | 36468 | 0.002 | 36488 | 0.002
36539 | 0.002 | 36584 | 0.003 | 36623 | 0.003 | 36585 | 0.001 | 36605 | 0.002
36657 | 0.002 | 36702 | 0.003 | 36740 | 0.003 | 36703 | 0.002 | 36723 | 0.002
36774 | 0.002 | 36819 | 0.003 | 36857 | 0.003 | 36820 | 0.002 | 36841 | 0.002
36891 | 0.002 | 36936 | 0.003 | 36975 | 0.003 | 36937 | 0.002 | 36958 | 0.002
37009 | 0.002 | 37057 | 0.003 | 37092 | 0.003 | 37055 | 0.002 | 37075 | 0.002
37127 | 0.002 | 37175 | 0.003 | 37210 | 0.003 | 37173 | 0.002 | 37193 | 0.002
37243 | 0.002 | 37293 | 0.003 | 37327 | 0.003 | 37289 | 0.002 | 37310 | 0.002
37361 | 0.002 | 37409 | 0.003 | 37444 | 0.003 | 37407 | 0.002 | 37428 | 0.002
37479 | 0.002 | 37527 | 0.003 | 37562 | 0.003 | 37525 | 0.002 | 37545 | 0.002
37596 | 0.002 | 37645 | 0.003 | 37680 | 0.003 | 37642 | 0.002 | 37662 | 0.002
37713 | 0.002 | 37762 | 0.003 | 37796 | 0.003 | 37760 | 0.002 | 37780 | 0.002
37831 | 0.002 | 37879 | 0.003 | 37914 | 0.003 | 37877 | 0.002 | 37898 | 0.002
37948 | 0.002 | 37997 | 0.003 | 38032 | 0.003 | 37994 | 0.002 | 38014 | 0.002
38066 | 0.002 | 38114 | 0.003 | 38149 | 0.003 | 38112 | 0.001 | 38132 | 0.002
38183 | 0.002 | 38232 | 0.003 | 38266 | 0.003 | 38230 | 0.001 | 38250 | 0.002
38300 | 0.002 | 38349 | 0.003 | 38384 | 0.003 | 38346 | 0.001 | 38367 | 0.002
38418 | 0.002 | 38466 | 0.003 | 38501 | 0.003 | 38464 | 0.001 | 38484 | 0.002
38536 | 0.003 | 38584 | 0.002 | 38619 | 0.003 | 38582 | 0.001 | 38602 | 0.002
38652 | 0.003 | 38702 | 0.003 | 38736 | 0.003 | 38700 | 0.001 | 38719 | 0.002
38770 | 0.002 | 38818 | 0.003 | 38854 | 0.003 | 38816 | 0.002 | 38837 | 0.002
38888 | 0.002 | 38936 | 0.003 | 38971 | 0.003 | 38934 | 0.002 | 38955 | 0.002
39005 | 0.002 | 39054 | 0.003 | 39089 | 0.003 | 39052 | 0.001 | 39071 | 0.002
39122 | 0.002 | 39171 | 0.003 | 39206 | 0.003 | 39169 | 0.001 | 39189 | 0.002
39240 | 0.003 | 39288 | 0.003 | 39323 | 0.003 | 39286 | 0.001 | 39307 | 0.002
39357 | 0.003 | 39406 | 0.003 | 39441 | 0.003 | 39404 | 0.002 | 39424 | 0.002
39475 | 0.003 | 39523 | 0.003 | 39559 | 0.003 | 39521 | 0.002 | 39541 | 0.002
39592 | 0.003 | 39641 | 0.003 | 39675 | 0.003 | 39639 | 0.001 | 39659 | 0.002
39709 | 0.003 | 39758 | 0.003 | 39793 | 0.003 | 39756 | 0.001 | 39776 | 0.002
39827 | 0.003 | 39876 | 0.003 | 39911 | 0.003 | 39873 | 0.001 | 39894 | 0.002
39945 | 0.003 | 39993 | 0.003 | 40028 | 0.003 | 39991 | 0.001 | 40011 | 0.002
40062 | 0.003 | 40111 | 0.003 | 40145 | 0.003 | 40109 | 0.001 | 40129 | 0.002
40179 | 0.003 | 40228 | 0.003 | 40263 | 0.003 | 40225 | 0.001 | 40246 | 0.002
40300 | 0.002 | 40345 | 0.003 | 40380 | 0.003 | 40343 | 0.002 | 40364 | 0.002
40418 | 0.003 | 40463 | 0.003 | 40498 | 0.003 | 40461 | 0.002 | 40481 | 0.002
40536 | 0.003 | 40581 | 0.003 | 40619 | 0.003 | 40578 | 0.001 | 40598 | 0.002
40652 | 0.003 | 40697 | 0.003 | 40736 | 0.003 | 40695 | 0.001 | 40719 | 0.002
40770 | 0.003 | 40815 | 0.003 | 40853 | 0.003 | 40813 | 0.001 | 40837 | 0.002
40888 | 0.003 | 40933 | 0.003 | 40971 | 0.003 | 40930 | 0.001 | 40954 | 0.002
41005 | 0.003 | 41050 | 0.003 | 41089 | 0.003 | 41051 | 0.001 | 41071 | 0.002
41122 | 0.003 | 41167 | 0.003 | 41206 | 0.003 | 41168 | 0.001 | 41189 | 0.002
41240 | 0.003 | 41285 | 0.003 | 41323 | 0.003 | 41286 | 0.001 | 41341 | 0.002
41357 | 0.003 | 41402 | 0.003 | 41441 | 0.003 | 41403 | 0.001 | 41458 | 0.002
41475 | 0.003 | 41520 | 0.003 | 41559 | 0.003 | 41521 | 0.001 | 41576 | 0.002
41592 | 0.003 | 41640 | 0.003 | 41709 | 0.003 | 41638 | 0.001 | 41694 | 0.002
41710 | 0.003 | 41758 | 0.003 | 41827 | 0.003 | 41756 | 0.001 | 41810 | 0.002
41860 | 0.003 | 41875 | 0.003 | 41945 | 0.003 | 41908 | 0.001 | 41928 | 0.002
41977 | 0.003 | 41993 | 0.003 | 42061 | 0.003 | 42024 | 0.001 | 42046 | 0.002
42095 | 0.003 | 42110 | 0.003 | 42179 | 0.003 | 42142 | 0.001 | 42163 | 0.002
42213 | 0.003 | 42227 | 0.003 | 42297 | 0.003 | 42260 | 0.001 | 42280 | 0.002
42330 | 0.003 | 42345 | 0.003 | 42415 | 0.003 | 42377 | 0.001 | 42398 | 0.002
42447 | 0.003 | 42463 | 0.003 | 42531 | 0.003 | 42494 | 0.001 | 42516 | 0.002
42565 | 0.003 | 42580 | 0.003 | 42649 | 0.003 | 42612 | 0.001 | 42633 | 0.002
42682 | 0.003 | 42697 | 0.003 | 42767 | 0.003 | 42729 | 0.001 | 42750 | 0.002
42800 | 0.003 | 42815 | 0.003 | 42884 | 0.003 | 42847 | 0.001 | 42868 | 0.002
42917 | 0.003 | 42933 | 0.003 | 43001 | 0.003 | 42964 | 0.001 | 42985 | 0.002
43035 | 0.003 | 43083 | 0.003 | 43119 | 0.003 | 43082 | 0.001 | 43103 | 0.002
43152 | 0.003 | 43201 | 0.003 | 43237 | 0.003 | 43199 | 0.001 | 43220 | 0.002
43270 | 0.003 | 43318 | 0.003 | 43354 | 0.003 | 43317 | 0.001 | 43337 | 0.002
43387 | 0.003 | 43436 | 0.002 | 43471 | 0.003 | 43435 | 0.001 | 43455 | 0.002
43504 | 0.003 | 43553 | 0.002 | 43589 | 0.003 | 43551 | 0.001 | 43573 | 0.002
43622 | 0.003 | 43670 | 0.003 | 43706 | 0.003 | 43669 | 0.001 | 43691 | 0.002
43740 | 0.003 | 43788 | 0.003 | 43824 | 0.003 | 43787 | 0.001 | 43807 | 0.002
43856 | 0.003 | 43906 | 0.003 | 43941 | 0.003 | 43905 | 0.001 | 43925 | 0.002
43974 | 0.003 | 44022 | 0.003 | 44058 | 0.003 | 44021 | 0.001 | 44043 | 0.002
44092 | 0.003 | 44140 | 0.003 | 44176 | 0.003 | 44139 | 0.001 | 44160 | 0.002
44210 | 0.003 | 44258 | 0.003 | 44294 | 0.003 | 44257 | 0.001 | 44277 | 0.002
44326 | 0.003 | 44376 | 0.003 | 44411 | 0.003 | 44374 | 0.001 | 44395 | 0.002
44444 | 0.003 | 44492 | 0.003 | 44528 | 0.003 | 44491 | 0.001 | 44513 | 0.002
44562 | 0.003 | 44610 | 0.003 | 44646 | 0.003 | 44609 | 0.001 | 44630 | 0.002
44679 | 0.003 | 44728 | 0.003 | 44764 | 0.003 | 44726 | 0.001 | 44747 | 0.002
44796 | 0.003 | 44845 | 0.003 | 44880 | 0.003 | 44844 | 0.001 | 44865 | 0.002
44914 | 0.003 | 44962 | 0.003 | 44998 | 0.003 | 44961 | 0.001 | 44982 | 0.002
45035 | 0.003 | 45080 | 0.002 | 45116 | 0.003 | 45079 | 0.001 | 45100 | 0.002
45152 | 0.003 | 45197 | 0.003 | 45233 | 0.003 | 45196 | 0.001 | 45217 | 0.002
45270 | 0.003 | 45315 | 0.003 | 45348 | 0.003 | 45314 | 0.001 | 45338 | 0.002
45387 | 0.003 | 45432 | 0.003 | 45466 | 0.003 | 45431 | 0.001 | 45455 | 0.002
45505 | 0.003 | 45550 | 0.003 | 45583 | 0.003 | 45548 | 0.001 | 45573 | 0.002
45622 | 0.003 | 45667 | 0.003 | 45701 | 0.003 | 45669 | 0.001 | 45690 | 0.002
45740 | 0.003 | 45785 | 0.003 | 45818 | 0.003 | 45787 | 0.001 | 45808 | 0.002
45857 | 0.003 | 45902 | 0.003 | 45935 | 0.003 | 45904 | 0.001 | 45925 | 0.002
45974 | 0.003 | 46019 | 0.003 | 46053 | 0.003 | 46021 | 0.001 | 46043 | 0.002
46092 | 0.003 | 46137 | 0.003 | 46171 | 0.003 | 46139 | 0.001 | 46160 | 0.002
46210 | 0.003 | 46255 | 0.003 | 46288 | 0.003 | 46257 | 0.001 | 46277 | 0.002
46326 | 0.003 | 46375 | 0.003 | 46405 | 0.003 | 46374 | 0.001 | 46395 | 0.002
46444 | 0.003 | 46492 | 0.003 | 46523 | 0.003 | 46491 | 0.001 | 46513 | 0.002
46562 | 0.003 | 46610 | 0.003 | 46641 | 0.003 | 46609 | 0.001 | 46630 | 0.002
46680 | 0.003 | 46728 | 0.003 | 46757 | 0.003 | 46727 | 0.001 | 46747 | 0.002
46796 | 0.003 | 46845 | 0.003 | 46875 | 0.003 | 46843 | 0.001 | 46865 | 0.002
46914 | 0.002 | 46962 | 0.003 | 46993 | 0.003 | 46961 | 0.001 | 46983 | 0.002
47032 | 0.002 | 47080 | 0.003 | 47110 | 0.003 | 47079 | 0.001 | 47099 | 0.002
47149 | 0.003 | 47198 | 0.003 | 47227 | 0.003 | 47197 | 0.001 | 47217 | 0.002
47266 | 0.003 | 47314 | 0.003 | 47345 | 0.003 | 47313 | 0.001 | 47335 | 0.002
47384 | 0.003 | 47432 | 0.003 | 47463 | 0.003 | 47431 | 0.001 | 47452 | 0.002
47502 | 0.003 | 47550 | 0.003 | 47580 | 0.003 | 47549 | 0.001 | 47569 | 0.002
47619 | 0.003 | 47668 | 0.003 | 47697 | 0.003 | 47666 | 0.001 | 47687 | 0.002
47736 | 0.002 | 47784 | 0.003 | 47815 | 0.003 | 47783 | 0.001 | 47805 | 0.002
47854 | 0.002 | 47902 | 0.003 | 47932 | 0.003 | 47901 | 0.001 | 47922 | 0.002
47971 | 0.003 | 48020 | 0.003 | 48050 | 0.003 | 48018 | 0.001 | 48039 | 0.002
48089 | 0.003 | 48137 | 0.003 | 48167 | 0.003 | 48136 | 0.001 | 48157 | 0.002
48206 | 0.003 | 48254 | 0.002 | 48285 | 0.003 | 48253 | 0.001 | 48274 | 0.002
48323 | 0.002 | 48372 | 0.003 | 48402 | 0.003 | 48371 | 0.001 | 48392 | 0.002
48441 | 0.003 | 48489 | 0.003 | 48520 | 0.003 | 48488 | 0.001 | 48509 | 0.002
48559 | 0.002 | 48607 | 0.003 | 48637 | 0.003 | 48606 | 0.001 | 48627 | 0.002
48676 | 0.002 | 48724 | 0.003 | 48754 | 0.003 | 48723 | 0.001 | 48744 | 0.002
48793 | 0.003 | 48842 | 0.003 | 48872 | 0.003 | 48840 | 0.001 | 48862 | 0.002
48911 | 0.003 | 48959 | 0.003 | 48990 | 0.003 | 48958 | 0.001 | 48979 | 0.002
49029 | 0.003 | 49077 | 0.003 | 49106 | 0.003 | 49076 | 0.001 | 49096 | 0.002
49145 | 0.003 | 49195 | 0.003 | 49224 | 0.003 | 49193 | 0.001 | 49214 | 0.002
49263 | 0.003 | 49311 | 0.003 | 49342 | 0.003 | 49310 | 0.001 | 49332 | 0.002
49381 | 0.003 | 49429 | 0.003 | 49460 | 0.003 | 49428 | 0.001 | 49448 | 0.002
49498 | 0.003 | 49547 | 0.003 | 49576 | 0.003 | 49546 | 0.001 | 49566 | 0.002
49618 | 0.003 | 49664 | 0.003 | 49694 | 0.003 | 49662 | 0.002 | 49684 | 0.002
49736 | 0.003 | 49781 | 0.003 | 49812 | 0.003 | 49780 | 0.001 | 49802 | 0.002
49854 | 0.003 | 49899 | 0.003 | 49933 | 0.003 | 49898 | 0.001 | 49918 | 0.002
49972 | 0.003 | 50017 | 0.003 | 50050 | 0.003 | 50015 | 0.001 | 50039 | 0.002
50088 | 0.003 | 50134 | 0.003 | 50167 | 0.003 | 50132 | 0.001 | 50157 | 0.002
50206 | 0.003 | 50251 | 0.003 | 50285 | 0.003 | 50250 | 0.001 | 50275 | 0.002
50324 | 0.003 | 50369 | 0.003 | 50402 | 0.003 | 50371 | 0.001 | 50392 | 0.002
50441 | 0.003 | 50486 | 0.003 | 50520 | 0.003 | 50489 | 0.001 | 50509 | 0.002
50558 | 0.003 | 50604 | 0.003 | 50637 | 0.003 | 50606 | 0.001 | 50627 | 0.002
50676 | 0.003 | 50721 | 0.003 | 50755 | 0.003 | 50723 | 0.002 | 50744 | 0.002
50794 | 0.003 | 50838 | 0.003 | 50872 | 0.003 | 50841 | 0.001 | 50862 | 0.002
50911 | 0.003 | 50959 | 0.003 | 50990 | 0.003 | 50958 | 0.001 | 50979 | 0.002
51028 | 0.003 | 51077 | 0.003 | 51107 | 0.003 | 51076 | 0.002 | 51097 | 0.002
51146 | 0.003 | 51194 | 0.003 | 51224 | 0.003 | 51193 | 0.002 | 51214 | 0.002
51263 | 0.003 | 51312 | 0.003 | 51342 | 0.003 | 51310 | 0.002 | 51332 | 0.002
51381 | 0.003 | 51429 | 0.003 | 51460 | 0.003 | 51428 | 0.002 | 51449 | 0.002
51499 | 0.003 | 51547 | 0.003 | 51577 | 0.003 | 51546 | 0.002 | 51566 | 0.002
51615 | 0.003 | 51664 | 0.003 | 51694 | 0.003 | 51663 | 0.002 | 51684 | 0.002
51733 | 0.003 | 51781 | 0.003 | 51812 | 0.003 | 51780 | 0.002 | 51802 | 0.002
51851 | 0.003 | 51899 | 0.003 | 51930 | 0.003 | 51898 | 0.002 | 51953 | 0.002
51969 | 0.003 | 52017 | 0.003 | 52046 | 0.003 | 52016 | 0.002 | 52071 | 0.002
52085 | 0.003 | 52133 | 0.003 | 52164 | 0.003 | 52132 | 0.002 | 52189 | 0.002
52203 | 0.003 | 52251 | 0.003 | 52282 | 0.003 | 52250 | 0.002 | 52305 | 0.002
52321 | 0.003 | 52369 | 0.003 | 52433 | 0.003 | 52368 | 0.002 | 52423 | 0.002
52438 | 0.003 | 52487 | 0.003 | 52550 | 0.003 | 52485 | 0.002 | 52541 | 0.002
52555 | 0.003 | 52603 | 0.003 | 52668 | 0.003 | 52636 | 0.002 | 52659 | 0.002
52706 | 0.003 | 52721 | 0.003 | 52786 | 0.003 | 52754 | 0.002 | 52775 | 0.002
52824 | 0.003 | 52839 | 0.003 | 52902 | 0.003 | 52872 | 0.002 | 52893 | 0.002
52940 | 0.003 | 52956 | 0.003 | 53020 | 0.003 | 52989 | 0.002 | 53011 | 0.002
53058 | 0.003 | 53073 | 0.003 | 53138 | 0.003 | 53106 | 0.001 | 53128 | 0.002
53176 | 0.003 | 53191 | 0.003 | 53255 | 0.003 | 53224 | 0.002 | 53245 | 0.002
53294 | 0.003 | 53308 | 0.003 | 53372 | 0.003 | 53342 | 0.002 | 53363 | 0.002
53410 | 0.003 | 53426 | 0.003 | 53490 | 0.002 | 53458 | 0.002 | 53480 | 0.002
53528 | 0.003 | 53543 | 0.003 | 53608 | 0.002 | 53576 | 0.002 | 53598 | 0.002
53646 | 0.003 | 53661 | 0.003 | 53725 | 0.003 | 53694 | 0.002 | 53715 | 0.002
53763 | 0.003 | 53778 | 0.003 | 53842 | 0.003 | 53811 | 0.002 | 53833 | 0.002
53880 | 0.003 | 53928 | 0.003 | 53960 | 0.003 | 53928 | 0.002 | 53950 | 0.002
53998 | 0.003 | 54046 | 0.003 | 54077 | 0.003 | 54046 | 0.002 | 54068 | 0.002
54115 | 0.003 | 54164 | 0.003 | 54195 | 0.003 | 54164 | 0.002 | 54185 | 0.002
54233 | 0.003 | 54282 | 0.003 | 54312 | 0.003 | 54281 | 0.002 | 54302 | 0.002
54348 | 0.003 | 54398 | 0.003 | 54429 | 0.003 | 54398 | 0.002 | 54420 | 0.002
54465 | 0.003 | 54516 | 0.003 | 54550 | 0.003 | 54516 | 0.002 | 54538 | 0.002
54583 | 0.003 | 54634 | 0.003 | 54668 | 0.003 | 54633 | 0.002 | 54659 | 0.002
54701 | 0.003 | 54751 | 0.003 | 54786 | 0.003 | 54751 | 0.002 | 54775 | 0.002
54817 | 0.003 | 54868 | 0.003 | 54903 | 0.003 | 54868 | 0.002 | 54893 | 0.002
54935 | 0.003 | 54986 | 0.003 | 55020 | 0.003 | 54989 | 0.002 | 55011 | 0.002
55053 | 0.003 | 55103 | 0.003 | 55138 | 0.003 | 55106 | 0.002 | 55129 | 0.002
55171 | 0.003 | 55221 | 0.003 | 55256 | 0.003 | 55224 | 0.002 | 55245 | 0.002
55287 | 0.003 | 55339 | 0.003 | 55372 | 0.003 | 55342 | 0.002 | 55363 | 0.002
55405 | 0.003 | 55456 | 0.003 | 55490 | 0.003 | 55459 | 0.002 | 55481 | 0.002
55523 | 0.003 | 55573 | 0.003 | 55608 | 0.003 | 55576 | 0.002 | 55599 | 0.002
55651 | 0.003 | 55694 | 0.003 | 55726 | 0.003 | 55694 | 0.002 | 55715 | 0.002
55769 | 0.003 | 55812 | 0.003 | 55842 | 0.003 | 55812 | 0.002 | 55833 | 0.002
55885 | 0.003 | 55929 | 0.003 | 55960 | 0.003 | 55928 | 0.002 | 55951 | 0.002
56003 | 0.003 | 56046 | 0.003 | 56078 | 0.003 | 56046 | 0.002 | 56068 | 0.002
56121 | 0.003 | 56164 | 0.003 | 56206 | 0.003 | 56164 | 0.002 | 56185 | 0.002
56239 | 0.003 | 56282 | 0.003 | 56323 | 0.003 | 56282 | 0.002 | 56313 | 0.002
56355 | 0.003 | 56398 | 0.003 | 56440 | 0.003 | 56398 | 0.002 | 56431 | 0.002
56473 | 0.003 | 56516 | 0.003 | 56558 | 0.003 | 56516 | 0.002 | 56548 | 0.002
56591 | 0.003 | 56634 | 0.003 | 56675 | 0.003 | 56644 | 0.002 | 56666 | 0.002
56708 | 0.003 | 56752 | 0.003 | 56793 | 0.003 | 56762 | 0.002 | 56783 | 0.002
56825 | 0.003 | 56868 | 0.003 | 56910 | 0.003 | 56879 | 0.002 | 56900 | 0.002
56943 | 0.003 | 56986 | 0.003 | 57027 | 0.003 | 56996 | 0.002 | 57018 | 0.002
57060 | 0.003 | 57104 | 0.003 | 57145 | 0.003 | 57114 | 0.002 | 57136 | 0.002
57178 | 0.003 | 57222 | 0.003 | 57263 | 0.003 | 57231 | 0.002 | 57252 | 0.002
57295 | 0.003 | 57349 | 0.003 | 57380 | 0.003 | 57349 | 0.002 | 57370 | 0.002
57413 | 0.003 | 57466 | 0.003 | 57497 | 0.003 | 57466 | 0.002 | 57488 | 0.002
57530 | 0.003 | 57584 | 0.003 | 57615 | 0.003 | 57583 | 0.002 | 57606 | 0.002
57648 | 0.003 | 57701 | 0.003 | 57733 | 0.003 | 57701 | 0.002 | 57722 | 0.002
57765 | 0.003 | 57819 | 0.003 | 57849 | 0.003 | 57819 | 0.002 | 57840 | 0.002
57882 | 0.003 | 57936 | 0.003 | 57967 | 0.003 | 57936 | 0.002 | 57958 | 0.002
58000 | 0.003 | 58054 | 0.003 | 58085 | 0.003 | 58053 | 0.002 | 58075 | 0.002
58118 | 0.003 | 58171 | 0.003 | 58202 | 0.003 | 58171 | 0.002 | 58192 | 0.002
58234 | 0.003 | 58289 | 0.003 | 58319 | 0.003 | 58289 | 0.002 | 58310 | 0.002
58352 | 0.003 | 58406 | 0.003 | 58437 | 0.003 | 58405 | 0.002 | 58428 | 0.002
58470 | 0.003 | 58523 | 0.003 | 58555 | 0.003 | 58523 | 0.002 | 58545 | 0.002
58588 | 0.003 | 58641 | 0.003 | 58672 | 0.002 | 58641 | 0.002 | 58662 | 0.002
58704 | 0.003 | 58759 | 0.003 | 58789 | 0.002 | 58759 | 0.002 | 58780 | 0.002
58822 | 0.003 | 58875 | 0.003 | 58907 | 0.003 | 58875 | 0.002 | 58897 | 0.002
58943 | 0.003 | 58993 | 0.003 | 59024 | 0.003 | 58993 | 0.002 | 59015 | 0.002
59061 | 0.003 | 59111 | 0.003 | 59142 | 0.003 | 59111 | 0.001 | 59132 | 0.002
59177 | 0.003 | 59229 | 0.003 | 59262 | 0.003 | 59228 | 0.002 | 59249 | 0.002
59295 | 0.003 | 59345 | 0.003 | 59380 | 0.003 | 59345 | 0.002 | 59370 | 0.002
59413 | 0.003 | 59463 | 0.003 | 59498 | 0.003 | 59463 | 0.002 | 59488 | 0.002
59531 | 0.003 | 59581 | 0.003 | 59615 | 0.003 | 59584 | 0.002 | 59606 | 0.002
59648 | 0.003 | 59698 | 0.003 | 59732 | 0.002 | 59701 | 0.002 | 59723 | 0.002
59765 | 0.003 | 59815 | 0.003 | 59850 | 0.003 | 59819 | 0.002 | 59840 | 0.002
59883 | 0.003 | 59933 | 0.003 | 59967 | 0.003 | 59936 | 0.002 | 59958 | 0.002
60000 | 0.003 | 60050 | 0.003 | 60085 | 0.003 | 60054 | 0.002 | 60076 | 0.002
60118 | 0.003 | 60168 | 0.003 | 60203 | 0.003 | 60171 | 0.002 | 60192 | 0.002
60235 | 0.002 | 60289 | 0.003 | 60319 | 0.003 | 60289 | 0.002 | 60310 | 0.002
60352 | 0.003 | 60406 | 0.003 | 60437 | 0.003 | 60406 | 0.002 | 60428 | 0.002
60470 | 0.003 | 60524 | 0.003 | 60555 | 0.003 | 60523 | 0.001 | 60545 | 0.002
60588 | 0.003 | 60641 | 0.003 | 60673 | 0.003 | 60641 | 0.002 | 60662 | 0.002
60705 | 0.003 | 60759 | 0.003 | 60789 | 0.003 | 60759 | 0.002 | 60780 | 0.002
60822 | 0.002 | 60876 | 0.003 | 60907 | 0.003 | 60876 | 0.001 | 60898 | 0.002
60940 | 0.003 | 60993 | 0.003 | 61025 | 0.003 | 60993 | 0.002
61058 | 0.003 | 61111 | 0.003 | 61142 | 0.003 | 61111 | 0.002
61174 | 0.003 | 61229 | 0.003 | 61259 | 0.003 | 61229 | 0.002
61292 | 0.003 | 61345 | 0.003 |  |  | 61345 | 0.002
61410 | 0.002 | 61463 | 0.003 |  |  | 61463 | 0.002
61528 | 0.003 |  |  |  |  | 61581 | 0.002
61644 | 0.003 |  |  |  |  | 61698 | 0.002
[6 empty] | 61815 | 0.002
[6 empty] | 61933 | 0.002
[6 empty] | 62051 | 0.002
[6 empty] | 62168 | 0.002
[6 empty] | 62285 | 0.002