Height Correction when using well-type HPGe detectors

Validation of methodology

Issue 1

GAU/RC/2032

Produced by				………………………	Date: ……………..

Prof I W Croudace

Height Correction when using well-type HPGe detectors

Introduction

The counting efficiency of samples is altered when the sample vial is not completely filled. When this occurs it is necessary to apply a height correction factor to account for this.

Procedure used for determining height correction factors

A QCYK8163 mixed gamma standard in an aqueous matrix for detectors 1, 2, 8 and 10 and a silica matrix for detectors 1, 2 and 8 was made to various heights and counted. Fitzpeaks gamma analysis software was used for analysis and calibration (energy, peak shape and efficiency) using STDANALYSIS.LIB.  The data for each height was compared against the data obtained for the full counting vial and the average value for the radionuclides was plotted graphically to give the height correction factor required.

Table 1: Data for Detector 1

| silica |  | 0.2162 g | 0.1555 g | 0.0786 g | 0.0174 g | Ratio to full height |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nuclide | keV | 45 mm | 30 mm | 15 mm | 4 mm | 4mm | 15mm | 30mm |
| Am-241 | 59.54 | 302 | 360 | 400 | 463 | 1.53 | 1.29 | 1.16 |
| Cd-109 | 88.03 | 1920 | 2280 | 2570 | 2950 | 1.54 | 1.29 | 1.15 |
| Co-57 | 122.1 | 79.5 | 95.8 | 106.5 | 126 | 1.58 | 1.32 | 1.18 |
| Hg-203 | 279.2 | 255 | 284 | 313 | 457 | 1.79 | 1.61 | 1.46 |
| Sn-113 | 391.7 | 323 | 385 | 426 | 485 | 1.50 | 1.26 | 1.14 |
| Sr-85 | 514 | 414 | 477 | 551 | 605 | 1.46 | 1.29 | 1.10 |
| Cs-137 | 661.6 | 332 | 385 | 434 | 492 | 1.48 | 1.28 | 1.13 |
| Mn-54 | 834.8 | 366 | 430 | 481 | 514 | 1.40 | 1.20 | 1.07 |
| Zn-65 | 1116 | 869 | 1007 | 1160 | 1260 | 1.45 | 1.25 | 1.09 |
| Co-60 | 1173, 1333 | 323 | 364 | 393 | 427 | 1.32 | 1.17 | 1.09 |
| Y-88 | 898, 1836 | 626 | 713 | 792 | 849 | 1.36 | 1.19 | 1.07 |
|  |  |  |  |  | Average | 1.48 | 1.28 | 1.14 |
| water |  | 0.212 | 0.212 | 0.212 | 0.212 |  |  |  |
| Nuclide | keV | 45 mm | 30 mm | 15 mm | 4 mm |  |  |  |
| Am-241 | 59.54 | 308 | 382 | 435 | 479 | 1.56 | 1.25 | 1.10 |
| Cd-109 | 88.03 | 1740 | 2150 | 2450 | 2700 | 1.55 | 1.26 | 1.10 |
| Co-57 | 122.1 | 67.9 | 83.4 | 96 | 105.3 | 1.55 | 1.26 | 1.10 |
| Hg-203 | 279.2 | 209 | 229 | 280 | 310 | 1.48 | 1.35 | 1.11 |
| Sn-113 | 391.7 | 257 | 313 | 357 | 397 | 1.54 | 1.27 | 1.11 |
| Sr-85 | 514 | 349 | 397 | 483 | 520 | 1.49 | 1.31 | 1.08 |
| Cs-137 | 661.6 | 291 | 349 | 392 | 437 | 1.50 | 1.25 | 1.11 |
| Mn-54 | 834.8 | 336 | 406 | 468 | 510 | 1.52 | 1.26 | 1.09 |
| Zn-65 | 1116 | 853 | 1003 | 1136 | 1258 | 1.47 | 1.25 | 1.11 |
| Co-60 | 1173, 1333 | 335 | 382 | 425 | 461 | 1.38 | 1.21 | 1.08 |
| Y-88 | 898, 1836 | 641 | 750 | 816 | 906 | 1.41 | 1.21 | 1.11 |
|  |  |  |  |  | Average | 1.49 | 1.26 | 1.10 |

Table 2: Data for Detectors 2 and 8

| silica |  | 0.2162 | 0.1534 | 0.0851 | 0.0177 | Ratio to full height |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nuclide | keV | 45 mm | 30 mm | 15 mm | 4 mm | 4mm | 15mm | 30mm | 45mm |
| Am-241 | 59.54 | 334 | 390 | 413 | 482 | 1.44 | 1.24 | 1.17 | 1.00 |
| Cd-109 | 88.03 | 2050 | 2420 | 2580 | 3020 | 1.47 | 1.25 | 1.17 | 1.00 |
| Co-57 | 122.1 | 73.3 | 87.2 | 94.8 | 109.8 | 1.50 | 1.26 | 1.16 | 1.00 |
| Hg-203 | 279.2 | 185 | 228 | 250 | 307 | 1.66 | 1.35 | 1.23 | 1.00 |
| Sn-113 | 391.7 | 266 | 319 | 344 | 412 | 1.55 | 1.29 | 1.20 | 1.00 |
| Sr-85 | 514 | 400 | 455 | 503 | 632 | 1.58 | 1.39 | 1.26 | 1.00 |
| Cs-137 | 661.6 | 349 | 416 | 453 | 504 | 1.44 | 1.21 | 1.11 | 1.00 |
| Mn-54 | 834.8 | 437 | 520 | 583 | 648 | 1.48 | 1.25 | 1.11 | 1.00 |
| Zn-65 | 1116 | 1130 | 1351 | 1477 | 1740 | 1.54 | 1.29 | 1.18 | 1.00 |
| Co-60 | 1173, 1333 | 363 | 412 | 441 | 485 | 1.34 | 1.18 | 1.10 | 1.00 |
| Y-88 | 898, 1836 | 657 | 742 | 807 | 959 | 1.46 | 1.29 | 1.19 | 1.00 |
|  |  |  |  |  | Average | 1.48 | 1.26 | 1.16 | 1.00 |
| water |  | 0.2175 | 0.2175 | 0.2175 | 0.2175 |  |  |  |  |
| Nuklide | keV | 45 mm | 30 mm | 15 mm | 4 mm |  |  |  |  |
| Am-241 | 59.54 | 318 | 385 | 426 | 459 | 1.44 | 1.19 | 1.08 | 1.00 |
| Cd-109 | 88.03 | 1940 | 2350 | 2630 | 2830 | 1.46 | 1.20 | 1.08 | 1.00 |
| Co-57 | 122.1 | 69.2 | 84 | 93.9 | 102.9 | 1.49 | 1.23 | 1.10 | 1.00 |
| Hg-203 | 279.2 | 181 | 227 | 258 | 275 | 1.52 | 1.21 | 1.07 | 1.00 |
| Sn-113 | 391.7 | 250 | 296 | 343 | 378 | 1.51 | 1.28 | 1.10 | 1.00 |
| Sr-85 | 514 | 349 | 439 | 483 | 549 | 1.57 | 1.25 | 1.14 | 1.00 |
| Cs-137 | 661.6 | 329 | 396 | 453 | 493 | 1.50 | 1.24 | 1.09 | 1.00 |
| Mn-54 | 834.8 | 410 | 495 | 572 | 631 | 1.54 | 1.27 | 1.10 | 1.00 |
| Zn-65 | 1116 | 1065 | 1274 | 1483 | 1608 | 1.51 | 1.26 | 1.08 | 1.00 |
| Co-60 | 1173, 1333 | 343 | 398 | 444 | 480 | 1.40 | 1.21 | 1.08 | 1.00 |
| Y-88 | 898, 1836 | 630 | 719 | 788 | 850 | 1.35 | 1.18 | 1.08 | 1.00 |
|  |  |  |  |  | Average | 1.48 | 1.23 | 1.09 | 1.00 |

Table 3: Data for Detector 10

| Water |  | 1.0740g |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nuclide | keV | 1mm | 10mm | 20mm | 30mm | 40mm | 45mm | 1mm | 10mm | 20mm | 30mm | 40mm | 45mm |
| Am-241 | 59.54 | 147.9 | 137.1 | 128.1 | 117.5 | 103.8 | 95.8 | 1.54 | 1.43 | 1.34 | 1.23 | 1.08 | 1.00 |
| Cd-109 | 88.03 | 811 | 745 | 693 | 629 | 566 | 516 | 1.57 | 1.44 | 1.34 | 1.22 | 1.10 | 1.00 |
| Co-57 | 122.1 | 31 | 28.3 | 26.8 | 24.1 | 21.3 | 20.3 | 1.53 | 1.39 | 1.32 | 1.19 | 1.05 | 1.00 |
| Sn-113 | 391.7 | 127 | 156 | 131 | 107 | 97 | 94 | 1.35 | 1.66 | 1.39 | 1.14 | 1.03 | 1.00 |
| Cs-137 | 661.6 | 143.4 | 136.2 | 125.7 | 115.6 | 102.4 | 93.9 | 1.53 | 1.45 | 1.34 | 1.23 | 1.09 | 1.00 |
| Mn-54 | 834.8 | 135.4 | 126 | 115.1 | 104.3 | 90 | 87.1 | 1.55 | 1.45 | 1.32 | 1.20 | 1.03 | 1.00 |
| Zn-65 | 1116 | 290 | 277 | 249 | 237 | 212 | 189 | 1.53 | 1.47 | 1.32 | 1.25 | 1.12 | 1.00 |
| Co-60 | 1173, 1333 | 152.3 | 147.5 | 137.8 | 127.8 | 117.6 | 112.5 | 1.35 | 1.31 | 1.22 | 1.14 | 1.05 | 1.00 |
| Y-88 | 898, 1836 | 190 | 200 | 160 | 167 |  | 151 |  |  |  |  |  |  |
|  |  |  |  |  |  |  | Average | 1.48 | 1.41 | 1.30 | 1.16 | 1.04 | 1.00 |

Det 02 Correction for Volume

silica - marix

y = -0.011x + 1.4817

R

2

= 0.956

0

0.5

1

1.5

2

0

5

10

15

20

25

30

35

40

45

50

height [mm]

Figure 1-5: Height Correction factor

Conclusions

Table 4: Summary of Height correction equations

| Det 1 | Silica | y = -0.0112x + 1.4849 |
| --- | --- | --- |
|  | Water | y = -0.0116x + 1.4858 |
| Det 2 or 8 | Silica | y = -0.011x + 1.4817 |
|  | Water | y = -0.0112x + 1.4629 |
| Det 10 | Water | y = -0.0114x + 1.5085 |

There are only small differences between the different matrices and detector (Table 4) which are negligible compared to the uncertainties associated with counting statistics. On this basis an average value of y =-0.01128x + 1.48476 is used for all detectors (Table 5).

Table 5: Average height Correction factors for Detectors 1, 2, 8 and 10.

| Height | height | Height | height |
| --- | --- | --- | --- |
| [mm] | Coeff. | [mm] | Coeff. |
| 1 | 0.68 | 24 | 0.82 |
| 2 | 0.68 | 25 | 0.83 |
| 3 | 0.69 | 26 | 0.84 |
| 4 | 0.69 | 27 | 0.85 |
| 5 | 0.70 | 28 | 0.86 |
| 6 | 0.71 | 29 | 0.86 |
| 7 | 0.71 | 30 | 0.87 |
| 8 | 0.72 | 31 | 0.88 |
| 9 | 0.72 | 32 | 0.89 |
| 10 | 0.73 | 33 | 0.90 |
| 11 | 0.73 | 34 | 0.91 |
| 12 | 0.74 | 35 | 0.92 |
| 13 | 0.75 | 36 | 0.93 |
| 14 | 0.75 | 37 | 0.94 |
| 15 | 0.76 | 38 | 0.95 |
| 16 | 0.77 | 39 | 0.96 |
| 17 | 0.77 | 40 | 0.97 |
| 18 | 0.78 | 41 | 0.98 |
| 19 | 0.79 | 42 | 0.99 |
| 20 | 0.79 | 43 | 1.00 |
| 21 | 0.80 | 44 | 1.01 |
| 22 | 0.81 | 45 | 1.02 |
| 23 | 0.82 |  |  |