## Validation
Spreadsheet summary
Document Reference Number | GAU/SSH/044
Spreadsheet title | Alpha spectrometry - efficiency check
Issue number | Issue 1
Issue date | 2014-09-11
Constructed by | P. Gaca
Updated by
Date validated | 2014-09-10
Validated by | P. Warwick

## Calculations
GAU/SSH/044
Alpha spectrometry - efficiency check
Issue 1
U-236 | (SOC/RAM/186) | A= | 7.429 | Bq/g | Ref. date: | 2013-06-28
Pu-242 | (SOC/RAM/176) | A= | 100 | Bq/g | Ref. date: | 2013-06-28
Am-241 | (SOC/RAM/410) | A= | 434 | Bq/g | Ref. date: | 2013-06-28
DATE OF CALIBRATION: |  | 2020-12-15
Code | SOC/RAM number | U-236 (SOC/RAM/186) [g] | Pu-242 (SOC/RAM/176) [g] | Am-241 (SOC/RAM/410) [g] | U-236 [Bq] | Pu-242 [Bq] | Am-241 [Bq]
STD 1 | SOC/RAM/1031 | 0.5335 | 0.0432 | 0.01 | 3.9634 | 4.3199 | 4.2884
STD 2 | SOC/RAM/1032 | 0.5344 | 0.0427 | 0.01 | 3.9701 | 4.2699 | 4.2884
STD 3 | SOC/RAM/1033 | 0.534 | 0.0427 | 0.0104 | 3.9671 | 4.2699 | 4.4599
STD 4 | SOC/RAM/1034 | 0.5347 | 0.0427 | 0.0102 | 3.9723 | 4.2699 | 4.3742
STD 5 | SOC/RAM/1035 | 0.5347 | 0.0428 | 0.0102 | 3.9723 | 4.2799 | 4.3742
STD 6 | SOC/RAM/1036 | 0.5345 | 0.0426 | 0.0103 | 3.9708 | 4.2599 | 4.417
STD 7 | SOC/RAM/1037 | 0.5357 | 0.0427 | 0.0103 | 3.9797 | 4.2699 | 4.417
STD 8 | SOC/RAM/1038 | 0.5356 | 0.0427 | 0.0105 | 3.979 | 4.2699 | 4.5028
No | Det No | STD used | Instrument | U-236 counts | Pu-242 counts | Am-241 counts | counting time (live) [s] | U-236 counting efficiency [%] | Pu-242 countin efficiency [%] | Am-241 counting efficiency [%] | Average efficiency [%] | U-236 peak position [channel] | Pu-242 peak position [channel] | Am-241peak position [channel] | U-236 FWHM | Pu-242 FWHM | Am-241 FWHM
1 | 1 | STD 1 | Octete 1 | 1408 | 1566 | 1788 | 1000 | 35.5253 | 36.2505 | 41.694 | 37.8233 | 153.8 | 194.98 | 255.27
2 | 2 | STD 2 | Octete 1 | 1544 | 1595 | 1809 | 1000 | 38.8911 | 37.3541 | 42.1837 | 39.4763 | 155.8 | 197.57 | 258.02
3 | 3 | STD 3 | Octete 1 | 1362 | 1524 | 1697 | 1000 | 34.3325 | 35.6914 | 38.05 | 36.0246 | 158.15 | 200.32 | 260.29
4 | 4 | STD 4 | Octete 1 | 1464 | 1527 | 1708 | 1000 | 36.8554 | 35.7616 | 39.0475 | 37.2215 | 156.48 | 198.02 | 258.61
5 | 5 | STD 5 | Octete 1 | 1418 | 1560 | 1690 | 1000 | 35.6973 | 36.4491 | 38.636 | 36.9275 | 160.66 | 202.47 | 263.63
6 | 6 | STD 6 | Octete 1 | 1448 | 1633 | 1823 | 1000 | 36.4662 | 38.3339 | 41.272 | 38.6907 | 155.05 | 196.47 | 256.62
7 | 7 | STD 7 | Octete 1 | 1436 | 1586 | 1717 | 1000 | 36.083 | 37.1434 | 38.8722 | 37.3662 | 158.53 | 200.21 | 261.11
8 | 8 | STD 8 | Octete 1 | 1426 | 1509 | 1628 | 1000 | 35.8384 | 35.3401 | 36.1552 | 35.7779 | 156.95 | 199.43 | 259.83
9 | 9 | STD 1 | Octete 2 | 1497 | 1555 | 1735 | 1000 | 37.7709 | 35.9959 | 40.4581 | 38.0749 | 152.34 | 194.65 | 254.89
10 | 10 | STD 2 | Octete 2 | 1599 | 1650 | 1799 | 1000 | 40.2765 | 38.6422 | 41.9505 | 40.2897 | 157.61 | 200.22 | 260.69
11 | 11 | STD 3 | Octete 2 | 1406 | 1507 | 1757 | 1000 | 35.4416 | 35.2932 | 39.3953 | 36.71 | 155.95 | 197.72 | 258.27
12 | 12 | STD 4 | Octete 2 | 1458 | 1632 | 1706 | 1000 | 36.7043 | 38.2207 | 39.0018 | 37.9756 | 156.2 | 198.23 | 258.18
13 | 13 | STD 5 | Octete 2 | 1506 | 1619 | 1731 | 1000 | 37.9127 | 37.8276 | 39.5733 | 38.4379 | 157.88 | 199.51 | 260.55
14 | 14 | STD 6 | Octete 2 | 1433 | 1579 | 1797 | 1000 | 36.0884 | 37.0662 | 40.6833 | 37.946 | 160.71 | 203.31 | 264.37
15 | 15 | STD 7 | Octete 2 | 1480 | 1599 | 1772 | 1000 | 37.1886 | 37.4478 | 40.1173 | 38.2513 | 155.89 | 197.85 | 259.29
16 | 16 | STD 8 | Octete 2 | 1457 | 1540 | 1616 | 1000 | 36.6175 | 36.0661 | 35.8887 | 36.1908 | 156.73 | 198.46 | 259.18
17 | 17 | STD 1 | Alpha Duo 1 | 1282 | 1411 | 1652 | 1000 | 32.3462 | 32.6625 | 38.5226 | 34.5104 | 319.04 | 375.96 | 459.37
18 | 18 | STD 2 | Alpha Duo 1 | 1301 | 1446 | 1603 | 1000 | 32.7703 | 33.8646 | 37.38 | 34.6716 | 317.03 | 374.54 | 456.64
19 | 19 | STD 3 | Alpha Duo 2 | 1262 | 1393 | 1650 | 1000 | 31.8118 | 32.6234 | 36.9961 | 33.8104 | 306.92 | 362.8 | 443.79
20 | 20 | STD 4 | Alpha Duo 2 | 1227 | 1412 | 1409 | 1000 | 30.889 | 33.0684 | 32.2119 | 32.0564 | 307.34 | 362.79 | 445.13
21 | A 1-1A | STD 1 | Alpha Analyst | 1612 | 1773 | 1879 | 1000 | 40.6725 | 41.0422 | 43.816 | 41.8436 | 285.43 | 345.46 | 433.63
22 | A 1-1B | STD 1 | Alpha Analyst | 1281 | 1419 | 1609 | 1000 | 32.321 | 32.8477 | 37.5199 | 34.2295 | 286.97 | 346.99 | 435.03
23 | A 1-2A | STD 2 | Alpha Analyst | 1440 | 1541 | 1669 | 1000 | 36.2715 | 36.0895 | 38.919 | 37.0933 | 284 | 344.16 | 432.09
24 | A 1-2B | STD 2 | Alpha Analyst | 1464 | 1604 | 1788 | 1000 | 36.876 | 37.5649 | 41.694 | 38.7116 | 276.56 | 335.82 | 423.09
25 | A 1-3A | STD 3 | Alpha Analyst | 1426 | 1560 | 1823 | 1000 | 35.9458 | 36.5345 | 40.8751 | 37.7851 | 290.59 | 351.25 | 440.18
26 | A 1-3B | STD 3 | Alpha Analyst | 1383 | 1497 | 1759 | 1000 | 34.8619 | 35.059 | 39.4401 | 36.4537 | 288.51 | 348.77 | 437.34
27 | A 1-4A | STD 4 | Alpha Analyst | 1667 | 1699 | 1815 | 1000 | 41.9658 | 39.7898 | 41.4937 | 41.0831 | 294.39 | 355.54 | 445.27
28 | A 1-4B | STD 4 | Alpha Analyst | 1184 | 1266 | 1370 | 1000 | 29.8065 | 29.6491 | 31.3203 | 30.2587 | 292.19 | 353.35 | 442.76
29 | A 1-5A | STD 5 | Alpha Analyst | 1529 | 1682 | 1843 | 1000 | 38.4917 | 39.2996 | 42.1338 | 39.975 | 297.89 | 358.96 | 448.75
30 | A 1-5B | STD 5 | Alpha Analyst | 1443 | 1675 | 1741 | 1000 | 36.3267 | 39.1361 | 39.8019 | 38.4216 | 285.54 | 345.94 | 434
31 | A 1-6A | STD 6 | Alpha Analyst | 1474 | 1698 | 1909 | 1000 | 37.121 | 39.8597 | 43.219 | 40.0666 | 320.79 | 381.29 | 470.27
32 | A 1-6B | STD 6 | Alpha Analyst | 1268 | 1383 | 1572 | 1000 | 31.9331 | 32.4652 | 35.5894 | 33.3293 | 286.85 | 347.22 | 435.22
Date: |  | 2020-12-18
Data enetered by: |  | mj

## Half lives and decay const.
GAU/SSH/044
Alpha spectrometry - efficiency check
Issue 1
| Nuclide | T 1/2 (years) | T 1/2 (s) | decay constatnt (s-1) | source
| U-236 | 23430000 | 739394568000000 | 9.3745e-16 | http://www.nucleide.org/DDEP_WG/Nuclides/U-236_tables.pdf
| Pu-242 | 373000 | 11770984800000 | 5.8886e-14 | http://www.nucleide.org/DDEP_WG/Nuclides/Pu-242_tables.pdf
| Am-241 | 432.6 | 13651817760 | 5.0773e-11 | http://www.nucleide.org/DDEP_WG/Nuclides/Am-241_tables.pdf
| Data enetered: |  |  |  |  | Date:
| Data checked: |  |  |  |  | Date:
Pswd: decay

## Summary table
GAU/SSH/044
Alpha spectrometry - efficiency check
Issue 1
Det No | Average efficiency [%]
OCTETE 1
1 | 37.8233
2 | 39.4763
3 | 36.0246
4 | 37.2215
5 | 36.9275
6 | 38.6907
7 | 37.3662
8 | 35.7779
OCTETE 2
9 | 38.0749
10 | 40.2897
11 | 36.71
12 | 37.9756
13 | 38.4379
14 | 37.946
15 | 38.2513
16 | 36.1908
ALPHA DUO 1
17 | 34.5104
18 | 34.6716
ALPHA DUO 2
19 | 33.8104
20 | 32.0564
ALPHA ANALYST
A 1-1A | 41.8436
A 1-1B | 34.2295
A 1-2A | 37.0933
A 1-2B | 38.7116
A 1-3A | 37.7851
A 1-3B | 36.4537
A 1-4A | 41.0831
A 1-4B | 30.2587
A 1-5A | 39.975
A 1-5B | 38.4216
A 1-6A | 40.0666
A 1-6B | 33.3293
Date: | 2020-12-18

## Sheet1
mean
1 | 35.5253 | 36.2505 | 41.694 | 37.8233 | 3.3717 | 6.1686
2 | 38.8911 | 37.3541 | 42.1837 | 39.4763 | 2.4674 | 3.2925
3 | 34.3325 | 35.6914 | 38.05 | 36.0246 | 1.881 | 3.7174
4 | 36.8554 | 35.7616 | 39.0475 | 37.2215 | 1.6733 | 2.1922
5 | 35.6973 | 36.4491 | 38.636 | 36.9275 | 1.5266 | 2.9387
6 | 36.4662 | 38.3339 | 41.272 | 38.6907 | 2.4227 | 2.9381
7 | 36.083 | 37.1434 | 38.8722 | 37.3662 | 1.4079 | 2.7892
8 | 35.8384 | 35.3401 | 36.1552 | 35.7779 | 0.41092 | 0.31679
9 | 37.7709 | 35.9959 | 40.4581 | 38.0749 | 2.2466 | 4.4622
10 | 40.2765 | 38.6422 | 41.9505 | 40.2897 | 1.6542 | 1.674
11 | 35.4416 | 35.2932 | 39.3953 | 36.71 | 2.3267 | 3.9536
12 | 36.7043 | 38.2207 | 39.0018 | 37.9756 | 1.1682 | -1.5164
13 | 37.9127 | 37.8276 | 39.5733 | 38.4379 | 0.98425 | 1.6606
14 | 36.0884 | 37.0662 | 40.6833 | 37.946 | 2.4205 | 4.5949
15 | 37.1886 | 37.4478 | 40.1173 | 38.2513 | 1.6213 | 2.6695
16 | 36.6175 | 36.0661 | 35.8887 | 36.1908 | 0.38006 | -0.7288
17 | 32.3462 | 32.6625 | 38.5226 | 34.5104 | 3.4782 | 6.1764
18 | 32.7703 | 33.8646 | 37.38 | 34.6716 | 2.4085 | 4.6097
19 | 31.8118 | 32.6234 | 36.9961 | 33.8104 | 2.7886 | 5.1844
20 | 30.889 | 33.0684 | 32.2119 | 32.0564 | 1.098 | 1.3229
A 1-1A | 40.6725 | 41.0422 | 43.816 | 41.8436 | 1.7181 | 2.7737
A 1-1B | 32.321 | 32.8477 | 37.5199 | 34.2295 | 2.8617 | 5.1989
A 1-2A | 36.2715 | 36.0895 | 38.919 | 37.0933 | 1.5837 | 2.6475
A 1-2B | 36.876 | 37.5649 | 41.694 | 38.7116 | 2.6056 | 4.8179
A 1-3A | 35.9458 | 36.5345 | 40.8751 | 37.7851 | 2.6922 | 4.9293
A 1-3B | 34.8619 | 35.059 | 39.4401 | 36.4537 | 2.5882 | 4.3811
A 1-4A | 41.9658 | 39.7898 | 41.4937 | 41.0831 | 1.1446 | 1.7039
A 1-4B | 29.8065 | 29.6491 | 31.3203 | 30.2587 | 0.92279 | 1.5138
A 1-5A | 38.4917 | 39.2996 | 42.1338 | 39.975 | 1.9127 | 3.6421
A 1-5B | 36.3267 | 39.1361 | 39.8019 | 38.4216 | 1.8445 | 3.4752
A 1-6A | 37.121 | 39.8597 | 43.219 | 40.0666 | 3.0542 | 6.098
A 1-6B | 31.9331 | 32.4652 | 35.5894 | 33.3293 | 1.9754 | 3.6563

## Sheet2
| U-236 | Pu-242 | Am-241 | U-236 | Pu-242 | Am-241 |  | U-236 | Pu-242 | Am-241
| 155 | 197 | 259 | 313 | 371 | 454 |  | 287 | 347 | 435
| 157 | 199 | 261 | 312 | 369 | 451 |  | 285 | 345 | 433
| 152 | 194 | 254 | 301 | 357 | 438 |  | 288 | 348 | 436
| 157 | 199 | 261 | 303 | 360 | 442 |  | 282 | 341 | 429
| 152 | 194 | 255 | 313 | 370 | 453 |  | 295 | 357 | 446
| 156 | 198 | 260 | 312 | 368 | 450 |  | 293 | 354 | 443
| 158 | 201 | 262 | 300 | 356 | 437 |  | 295 | 357 | 446
| 156 | 199 | 260 | 302 | 358 | 440 |  | 293 | 355 | 444
| 156 | 198 | 260 | 313.5 | 370.01 | 453.38 |  | 298 | 359 | 449
| 156 | 198 | 261 | 311.71 | 368.48 | 450.52 |  | 297 | 358 | 447
| 156 | 198 | 259 | 301.1 | 356.92 | 437.8 |  | 289 | 350 | 438
| 157 | 199 | 261 | 303.83 | 359.3 | 441.06 |  | 288 | 348 | 436
| 157 | 200 | 262 |  |  |  |  | 287 | 347 | 435
| 157 | 199 | 261 |  |  |  |  | 286 | 347 | 434
| 154 | 196 | 257 |  |  |  |  | 289 | 349 | 438
| 155 | 197 | 259 |  |  |  |  | 283 | 343 | 431
[8 empty] | 296 | 357 | 447
Mean | 155.6875 | 197.875 | 259.5 | 307.1783 | 363.6425 | 445.6467 |  | 290 | 350 | 439
std | 1.7405 | 1.9279 | 2.3381 | 5.6962 | 6.1635 | 6.8493 |  | 294 | 355 | 444
STD % | 1.1179 | 0.97428 | 0.901 | 1.8544 | 1.6949 | 1.5369 |  | 293 | 354 | 443
[8 empty] | 297 | 359 | 448
5% | 7.7844 | 9.8938 | 12.975 | 15.3589 | 18.1821 | 22.2823 |  | 294 | 355 | 444
10% | 15.5687 | 19.7875 | 25.95 | 30.7178 | 36.3642 | 44.5647 |  | 291 | 351 | 440
[8 empty] | 288 | 349 | 437
[8 empty] | 285.84 | 346.01 | 434.03
[8 empty] | 284.1 | 344.1 | 431.7
[8 empty] | 289.19 | 349.66 | 437.9
[8 empty] | 290.37 | 351.09 | 440.19
[8 empty] | 293.9 | 354.85 | 444.17
[8 empty] | 292.88 | 353.7 | 442.68
[8 empty] | 292.96 | 353.9 | 443.47
[8 empty] | 292.94 | 354.3 | 443.6
[8 empty] | 297.43 | 358.53 | 448.45
[8 empty] | 291.7 | 352.35 | 440.78
[8 empty] | 292.57 | 353.32 | 442.58
[8 empty] | 287.59 | 347.81 | 436.1
[7 empty] | Mean | 290.8186 | 351.5172 | 440.2125
[7 empty] | stdev | 4.2568 | 4.8023 | 5.3952
[7 empty] | STDev % | 1.4637 | 1.3662 | 1.2256
[7 empty] | 5% | 14.5409 | 17.5759 | 22.0106
[7 empty] | 10% | 29.0819 | 35.1517 | 44.0213