## Validation
Spreadsheet title | Sr-90 and Ac-228 Quantulus calibration using Cerenkov radiation measurement
Document Reference Number | GAU/SSH/056
Issue number | Issue 1
Issue date
Constructed by | P. Gaca
Updated by
Date validated
Validated by

## Calculation
GAU/SSH/056: | Sr-90 and Ac-228 Quantulus calibration using Cerenkov radiation measurement
Issue 1
Counter (select from the drop-down list): | [7 empty] | Quantulus 3 | [5 empty] | Quantulus 1
Quantulus 2
Calibration date: | [7 empty] | 2020-09-29 09:12 | [5 empty] | Quantulus 3
Quantulus 4
Radionuclide (select from the drop-down list): | [7 empty] | Sr-90 | [5 empty] | Quantulus 5
Quantulus 6
955/20
Calibration number:
Sr-90 decay constant = | 7.6266e-10 | s-1
Ac-228 decay constant = | 3.1307e-05 | s-1
STD used: | [7 empty] | SOC/RAM/1553
STD Ref. date (from certificate or dilution preparation sheet): | [7 empty] | 2018-08-17 17:46
STD activity concentration uncertainty from certificate (2ϭ) [%] | [7 empty] | 0.75
STD activity concentration at reference date: | [7 empty] | 106.3 | Bq/g | +/- | 0.79725 | Bq/g
Decay constant [s-1] | [7 empty] | 7.6266e-10
STD activity concentration at calibration date: | [7 empty] | 101.0168 | +/- | 0.75763 | Bq/g
Total STD activity concentration (Y-90 in equilibrium included) - FOR Sr-90 ONLY | [7 empty] | 202.0337 | +/- | 1.5153 | Bq/g
Code | STD mass [g] | STD activity [Bq] | +/- | counting time [min] | CPM | 2 S.D. | BKG | 2 S.D. | CCPM | 2 S.D. |  | Counting effciency | 2 S.D.
STD 1 | 0.1028 | 10.3845 | 0.077884 | 20 | 447.45 | 9.4599 | 0.86 | 0.41473 | 446.59 | 9.469 |  | 71.6755 | 1.612
STD 2 | 0.1036 | 10.4653 | 0.07849 | 20 | 456.66 | 9.5568 | 0.86 | 0.41473 | 455.8 | 9.5658 |  | 72.5888 | 1.6178
STD 3 | 0.103 | 10.4047 | 0.078036 | 20 | 455.08 | 9.5402 | 0.86 | 0.41473 | 454.22 | 9.5492 |  | 72.7585 | 1.6241
STD 4 | 0.1035 | 10.4552 | 0.078414 | 20 | 460.25 | 9.5943 | 0.86 | 0.41473 | 459.39 | 9.6032 |  | 73.2312 | 1.6264
STD 5 | 0.103 | 10.4047 | 0.078036 | 20 | 465.97 | 9.6537 | 0.86 | 0.41473 | 465.11 | 9.6626 |  | 74.5029 | 1.6456
STD 6 | 0.1031 | 10.4148 | 0.078111 | 20 | 455.54 | 9.5451 | 0.86 | 0.41473 | 454.68 | 9.5541 |  | 72.7616 | 1.6234
STD 7 | 0.1028 | 10.3845 | 0.077884 | 20 | 451.54 | 9.5031 | 0.86 | 0.41473 | 450.68 | 9.5121 |  | 72.3319 | 1.6202
STD 8 | 0.1036 | 10.4653 | 0.07849 | 20 | 454.23 | 9.5313 | 0.86 | 0.41473 | 453.37 | 9.5403 |  | 72.2018 | 1.613
STD 9 | 0.1029 | 10.3946 | 0.07796 | 20 | 461.16 | 9.6037 | 0.86 | 0.41473 | 460.3 | 9.6127 |  | 73.8041 | 1.6377
[9 empty] | Efficiency [%]: |  | 72.8729 | +/- | 2.3778
Entered by: |  |  |  |  | Checked by:
Date: |  |  |  |  | Date:

## Nuclear data
Sr-90 and Ac-228 Quantulus calibration using Cerenkov radiation measurement
Counters: |  | Radionuclides:
Quantulus 1 |  | Sr-90
Quantulus 2 |  | Ac-228
Quantulus 3
Quantulus 4
Quantulus 5
Quantulus 6
Nuclear data
Radionuclide | T 1/2 | unit | Source
Sr-90 | 28.8 | y | http://www.nucleide.org/DDEP_WG/Nuclides/Sr-90_tables.pdf
Y-90 | 2.6684 | d | http://www.nucleide.org/DDEP_WG/Nuclides/Y-90_tables.pdf
Ac-228 | 6.15 | h | http://www.nucleide.org/DDEP_WG/Nuclides/Ac-228_tables.pdf
Entered by |  |  |  | Checked by
Date |  |  |  | Date