- Scope

  - The procedure describes the determination of americium and plutonium in soils, sediments, sewage, swabs and building materials (e.g. concrete). Americium can be measured by gamma spectrometry if there are high activities present (Typical LOD for americium by gamma spectrometry is approximately 1 Bq/kg). For low activity samples, it is necessary to separate Am using chromatography followed by alpha spectrometry. The procedure uses anion chromatography to separate Pu isotopes from Am. A TRU column is then used to separate Am from Th and U. Further separation using anion exchange chromatography is required to remove rare earth elements from the Am fraction.

  - The sample origin will determine the initial treatment of the samples (i.e. acid leach or borate fusion). Solid samples will be dried and ignited first and the dry and ignited weights recorded. Swab samples are ignited without drying and the residue either acid digested or fused. Samples which are likely to contain refractory Pu (e.g. from weapons fallout; soils and riverine sediments) will require borate fusion to ensure complete dissolution of the Pu. Soluble Pu, such as that delivered by marine discharge, can be quantitatively leached using aqua regia (e.g. marine sediments).

  - A standard and blank sample (matrix-matched as far as possible) is run with each batch of samples to assess the overall accuracy of the procedure, and whether any cross- contamination has occurred.

- Method summary (for reports)

  - Borate Fusion

An aliquot of the ignited sample was spiked with 243Am and 242Pu as chemical yield monitors and fused at 1000°C with lithium metaborate flux. The resultant glass was dissolved and the solution filtered. The Am and Pu were isolated from the solution and one another by anion exchange chromatography. Americium and 238,239,240Pu activities were determined by alpha spectrometry.

  - Acid leach

An aliquot of the sample was spiked with 243Am and 242Pu as chemical yield monitors, and leached using aqua regia. The Am and Pu were isolated from the solution and one another by anion exchange chromatography. Americium and 238,239,240Pu activities were determined by alpha spectrometry.

- Interferences

All known interferences in the determination of Pu and Am isotopes are effectively separated when using the described methodology.

- Equipment and Reagents

  - Large (1cm in diameter) ion exchange columns

  - Lithium Metaborate flux

  - 0.2M PEG-2000 (40 g Polyethyleneglycol-2000 in 100 ml MQ water)

  - 4 % (w/v) ammonium oxalate (40 g ammonium oxalate in 1 litre MQ water)

  - 10 mg/ml Fe solution (4.8g FeCl3*9H2O dissolved in 10 ml 6M HCl + 90 ml M.Q. water)

  - Bromocresol green indicator

  - pH meter / indicator paper

  - Whatman filter papers: 9 cm GF/A; 9 cm 40; 15 cm 41

  - Filter funnels and vacuum filtration flasks

  - Borosilicate beakers: 50, 250 and 500 ml volume

  - Pt or stainless steel plate

  - Heating block / magnetic stirrer, plus stirrer bars

  - 50 ml centrifuge tubes

  - Ascorbic acid

  - 1 % (w/v) Na2SO4 solution (1 g Na2SO4 in 100 ml MQ water)

  - Eichrom 100-200 mesh anion-exchange resin Cl form

  - Eichrom TRU resin

  - Ammonia solution (S.G. 0.88)

  - Concentrated nitric acid (S.G. 1.42)

  - Concentrated hydrochloric acid (S.G. 1.18)

  - 1M nitric + 93% MeOH (4.3ml c. nitric acid in 70ml MQ water and add 930ml MeOH)

  - 2M nitric acid + NaNO2 (0.8g NaNO2 in 87ml MQ water and add 13ml c. nitric acid)

  - 0.1M hydrochloric acid + 0.5M NH4SCN + 80% MeOH (2ml c.HCl diluted to 200ml with MQ water, add 38g NH4SCN and then add 800ml MeOH)

  - 3M nitric acid (300 ml conc. HNO3 + 1300 ml MQ water)

  - 8M nitric acid (800 ml conc. HNO3 + 800 ml MQ water)

  - 0.5M hydrochloric acid (50 ml conc. HCl + 1150 ml MQ water)

  - 4M hydrochloric acid (400 ml conc. HCl + 800 ml MQ water)

  - 9M hydrochloric acid (900 ml conc. HCl + 300 ml MQ water)

  - Aqua Regia (3 parts conc HCl: 1 part conc. HNO3)

  - Ammonium iodide reagent (prepare fresh before use, do not store) – 9M HCl / 0.1M NH4I

(2.8g NH4I in 150 ml conc. HCl + 50 ml MQ water).

  - Traceable 242Pu (or 236Pu) and 243Am spikes

  - NaNO2 0.4 g/ml (3.2 g NaNO2 + 8 ml MQ water)

N.B. All reagents are Analytical grade unless otherwise stated.

- Methodology

  - Preparation of soils, sediments, sewage and building materials

    - Transfer approximately 10-15g of sample into a weighed, marked porcelain crucible. Transfer the crucible into an oven and dry at 110C for 4 hours. Note – at any stage of the analysis, especially during borate fusion and evaporation, avoid placing the samples with elevated Pu or Am activity concentrations next to the low-active ones. Use the data for Am-241 from the gamma spectrometric measurements as the guideline for Am and and Pu concentration assessment. If such data is not available, use any means necessary to avoid cross-contamination (e.g. by not overheating the samples on the hotplate during evaporation)

    - Remove the crucible from the oven and allow to it cool. Weigh the dry sample and transfer it into a sieve (1.41 mm mesh) and gently sieve. Retain the sieved fraction. Accurately weigh 10g of the sieved fraction back into the original crucible and transfer into a furnace at 450C. Leave the samples for 10 minutes and then increase the furnace temperature to 900C. Ignite the sample for 4 hours. Remove the sample from the furnace and allow it to cool. Record the mass of the crucible and ignited sample and calculate the dry / ignited mass ratio.

    - Transfer the sample to a labelled bag and retain.

    - Determine whether samples will be fused or leached using the decision path shown in Figure 1. Record the reasons for the choice in the laboratory notebook. Note - borate fusion should be used as the default procedure if insufficient information is available to accurately assess the form of Pu in the sample.

  - Preparation of swab samples

    - Transfer the swab into a clean beaker and label the beaker with a furnace-resistant marker pen. Place a borosilicate glass watch glass on top of the beaker.

    - Spike the swab with ~ 0.05 Bq of 242Pu (or 236Pu) and ~ 0.07 Bq of 243Am and record the mass of tracer added.

    - Ignite the swab by ramping the temperature from room temperature to 450C at approx. 5C/min and holding at 450C for approx. 8 hours (usually overnight). The exact temperature ramp rate is not critical to the ignition process. Avoid putting the sample into hot furnace (heated to above 200C) as it may cause rapid ignition of the sample resulting in the loss of the sample material.

    - Continue analysis as from Section 5.4.2. If swabs are known to contain intractable PuO2, transfer the ignited material into a crucible, add 5g lithium metaborate flux and continue with a borate fusion as described in Section 5.3.

Figure 1 : Decision pathway for selecting the most appropriate digestion procedure

  - Samples requiring Borate fusion

    - Accurately weigh ~ 2 g of ignited sample from the bag into a graphite crucible. Add approximately 5g of lithium metaborate flux, and stir the sample carefully. Spike the sample with ~ 0.05 Bq of 242Pu (or 236Pu) and ~ 0.07 Bq of 243Am and record the mass of tracer added. Transfer the sample to an oven set at 60C to dry the sample.

    - Place the crucible in a furnace at 1000C for approximately 15 minutes. For each sample, prepare a 250 ml beaker containing 200 ml of Milli-Q water. Place a clean platinum (or stainless steel) plate in the beaker. Remove the crucible from the furnace and ensure that the sample is completely mixed. Carefully pour the mixture into the water, ensuring that the fused mixture does not congregate as a single mass on the bottom of the beaker. If any of the mixture remains in the crucible, replace the crucible in the furnace and reheat. Repeat the transfer of the fused mixture into the beaker.

    - Remove the Pt plate, and carefully decant the water from the beaker. Add 40 ml of 8M HNO3 and 1ml of 0.2M PEG-2000.

    - Place a magnetic stirrer in the bottom of the beaker, cover the beaker with a watch glass and transfer to a magnetic stirrer/hotplate. Heat the mixture to 80C and stir the sample overnight to effect dissolution.

    - Add 1 ml NaNO2 (0.4g/ml) solution to each sample and continue heating for another 30 minutes.

    - Switch off the heating block and let the sample cool down. Filter the sample under suction through a 9.0cm GF/A filter paper supported on a Whatman 540 filter. Rinse the beaker with ~ 20 ml of MQ water and transfer the solution to the filter apparatus. Wash the filter and residue with MQ water. Go to section 5.5.

  - Samples requiring acid digest

    - Accurately weigh ~ 2 g of the ignited sample into a beaker, and spike with ~ 0.05 Bq of 242Pu (or 236Pu) and ~ 0.07 Bq of 243Am and record the mass of tracer added.

    - Add 20 ml fresh aqua regia to the sample, cover with a watch glass and place on a hot plate set at 90°C for 2 hours.

    - Remove from the hot plate, allow the sample to settle and cool, and decant the clear liquid into a centrifuge tube labelled with the sample i.d. Repeat the aqua regia attack detailed in section 5.4.2 if residue remains.

    - Remove the sample from the hot plate and allow to cool. Transfer all of the sample into the centrifuge tube from section 5.3.3 with MQ water washings, and centrifuge at 2500 rpm for 5 minutes. Transfer the supernatant to a 250 ml beaker, and go to section 5.5.

  - Iron hydroxide precipitation.

    - For low-Fe samples add 2 ml of 10 mg/ml Fe solution (e.g. to blank). Add 4 drops of bromocresol green to the filtrate and adjust the pH of the filtrate to ~ 4 using NH4OH (colour change to blue). Allow the solution to cool and then filter through a Whatman 41 filter. Discard the filtrate.

    - Dissolve the residue on the filter in 5 ml conc. HCl. If the volume of solution exceeds 5ml, add a further 5ml conc. HCl and sufficient water to give a final HCl concentration of 9M. Wash the filter with 15 ml 9M HCl. Collect all the solution. Add 4 drops of conc. HNO3 to the solution and proceed to section 5.6.

  - Sequential Separation of Am and Pu

    - Prepare an anion column (10 cm, 0.7 cm diameter, but see N.B.). Precondition the column with 10 ml 9M HCl. Place a clean beaker labelled with sample I.D. and “Am fraction” under the column. Load the solution onto the column.

    - Rinse the beaker with 5 ml 9M HCl and transfer the washings to the column, and then wash the column with 10 ml 9M HCl. Collect all the eluents in a beaker marked with the sample i.d. and “Am fraction” and place this fraction on a hot plate to reduce the volume to about 20 ml by evaporation. Reserve for Am purification.

    - Place a waste beaker under the column and wash the column with a further 30 ml of 9M HCl followed by 30 ml 8M HNO3. Wash the column again with 10 ml 9M HCl and discard all washings.

    - Place a clean beaker labelled with the sample i.d. and “Pu fraction” under the column and elute the Pu with 40 ml 9M HCl / NH4I (prepare fresh before use, do not store the unused solution). Add 5 ml conc. HNO3 to the eluent, place on a hotplate to evaporate to dryness prior to electrodeposition.

(N.B. If there is a lot of Fe present in the sample it may be necessary to increase the amount of anion resin used to prevent Fe breakthrough from the column. If Fe breakthrough does occur a second anion column will be required. Load the 1st column eluent directly onto the 2nd column, and collect the eluent (Am fraction). Then treat the 2nd column exactly the same as the first, combining the Pu eluents at the end of the chemical separation. If this is not done Pu recoveries will be reduced.)

  - Am purification

    - Add 1 ml of 10 mg/ml Fe solution to the Am fraction from section 5.6.2. Adjust the pH of the solution using NH4OH until an Fe(OH)3 precipitate forms. Allow the sample to cool and then centrifuge. Discard the supernatant.

    - Prepare a TRU resin column (5 cm high, 0.5 cm diameter). Precondition the column with 10 ml of 3M HNO3.

    - Dissolve the residue from section 5.7.1 in 10 ml 3M HNO3 plus 10 drops conc. HNO3, and add ~ 1 g ascorbic acid. Transfer the sample onto the TRU column and allow the sample to percolate into the column. Rinse the centrifuge tube with 5 ml of 3M HNO3, and add the rinsings to the column. Wash the column with 10ml of 3M nitric acid followed by 10ml of 2M HNO3/NaNO2. Wash the column again with 10ml 3M HNO3. Discard all eluents up to this stage.

    - Elute Am from the column using 30ml of 4M HCl and add 1 ml of 1% Na2SO4 solution to the eluent. Evaporate the eluent to dryness.

    - To separate Am from REE, prepare an anion column (2.5 x 0.7 cm) and precondition the column with 10ml of 1M HNO3/MeOH.  Dissolve the sample from 5.6.4 in 5 drops conc. HNO3 then add 1 ml MQ water. Add 10 ml 1M HNO3/MeOH and load onto the column with a 5 ml 1M HNO3/MeOH beaker rinse. Wash the column with 20ml of 0.1M HCl/NH4SCN/MeOH then with 15 ml of 1M HNO3/MeOH and discard the eluents.

    - Place a clean beaker labelled with the sample i.d. and “Am fraction” under the column and elute the Am with 20ml 0.5M HCl adding 1 ml of 1 % Na2SO4 to the eluent. Evaporate the sample to dryness prior to electrodeposition.

  - Electrodeposition of Am & Pu

    - Prepare the electrodeposition cells by placing a new stainless steel disc (cathode) labelled with the sample i.d and analyte into the base of the EDP cell. Place a new insert into the top of the EDP cell and screw on the top.  Check that the cell does not leak by adding distilled water to the cell. Discard water if no leaks are detected.

    - Dissolve the samples in 1 ml 1.2M HCl, then add 5 ml of 4 % ammonium oxalate and warm gently. Transfer the solution to the electrodeposition cell. Rinse the beaker using 5 ml MQ water, transferring the rinse solution to the cell, and attach the cathode connection (black wire).

    - Lower the Pt anode (red wire) into the solution, switch on the power supply and note whether bubbles are formed. Adjust the current to 400 mA and electrodeposit for 1.5 hours. Check current and sample level in the cell at regular intervals. Adjust the current accordingly and cover the cell with laboratory film.

    - After 1.5 hours add 5 drops of conc. NH4OH solution. Continue deposition for 30 seconds and then disconnect the power supply. Rinse the cell with water and remove the stainless steel disc. Wash the disc with water and dry on a hot plate. Determine the alpha activity on the disc by alpha spectrometry

- Cleaning Glassware

  - All non-volumetric glassware to be rinsed and cleaned in glasswasher. Glassware used for Fe(OH)3 precipitation should  be rinsed with 6M HCl followed by water prior to a glasswasher cleaning.

  - Plastic columns and filter funnels to be rinsed and soaked in 10 % HNO3 bath.

  - Watch glasses and magnetic stirrer bars to be rinsed and soaked in 2 % Decon90® solution followed by approx. 10% HCl wash and MQ water rinse.

- Safety

  - Always wear safety glasses, a labcoat and protective gloves during the operation.

  - Exercise extreme care in handling very hot apparatus. Ensure that equipment left to cool is clearly labelled as ‘hot’.

  - When pouring the fusion into water care should be taken to avoid splashing of the hot fusion mixture. Ensure that the beaker is placed on a heat-proof surface.

  - Take care whilst handling equipment in the acid and Decon90® baths.

- Instrument configuration and analyte measurement

When counting discs ensure they are placed on the top shelf within the alpha spectrometer chambers.

- QC

A blank sample (matrix matched as appropriate) and a control sample (either in-house standard, spiked matrix or certified reference material) with a known 241Am and Pu isotopic composition must be analysed with each batch of samples.

- Calculation spreadsheets

  - Calculate 241Am activity concentration using GAU/SSH/016 validated spreadsheet

  - Calculate 239,240Pu and 238Pu activity concentration using GAU/SSH/017 validated spreadsheet. If 242Pu activity concentration is to be determined save a copy of GAU/SSH/017 spreadsheet with the “242Pu analysis” added at the end of the file name. Enter 236Pu spike activity concentration and mass in the cells which relate to 242Pu tracer. Enter 242Pu measurement results into 239,240Pu cells. Sign the spreadsheet as “242Pu analysis” before passing it for cross-checking.

- Calculations

  - Alpha Spectrometry:

Bqg=Ciso−CBkgCspike−CBkg×A×1M

where	Ciso =	the total counts of the energy range of the isotope in question

CBkg =	the total background counts of that energy range

Cspike =	the total counts of the energy range of the experimental yield monitor

A = spike activity added

M = the mass of sample analysed (enter 1 to determine Bq/sample)

Counting uncertainty (Bqg)=2×activity×1σ100

Where      σ=Cspike+CbkgCspike−Cbkg×1002+Ciso+CbkgCiso−Cbkg×1002

The uncertainty is calculated at the 95% confidence interval and refers to counting statistical uncertainty only.

  - Limit of detection:

Limit of detection LD(Bqg)=2.86+4.78BkgT+1.36T×100E×100R×1M

Where 	BkgT = count time corrected background counts

T = total count time

E = detector efficiency

R = chemical recovery

M = sample mass

- Method uncertainties

Current method uncertainties are recorded in the method validation file

- Document modification log

Section 5.7.3	Incorrect reference to section 5.6.1 amended to 5.7.1