## Performance Summary
|  |  | 2018-06-21 | 2018-08-09
Parameter | Value | Units
Rate Bias Offset |  | deg/hr
Rate Bias Instability |  | deg/hr
Angle Random Walk
Rate Temperature Coefficient, c1
Shupe Coefficient, c2
(𝑑𝑒𝑔/ℎ𝑟)/(°𝐶/𝑚𝑖𝑛)
(𝑑𝑒𝑔/ℎ𝑟)/(°𝐶)
𝜇𝑑𝑒𝑔/√(ℎ𝑟.)

## Calc vs Measured
Constants
[5 empty] | Speed of Light |  | 299792458 | m/s
| Parameter | Value | Units |  | Index |  | 1.468
Constants
| Fibre Refractive Index | 1.468 |  | s | Light Source
| Light Source Mean Wavelength | 1547.21 | nm |  | Mean Wavelength |  | 1547.21 | nm
| Coil Mean Diameter | 0.068 | m |  |  |  | 1.5472e-06 | m
| Total Sensing Loop Length | 1001.493 | m
| Southampton Latitude | 50.9376 | deg |  | Coil
| Rotation Rate - Input Axis Vertical | 11.6469 | deg/hr. |  | Coil Length from input |  | 998 | m
| Equivalent Sagnac Phase - IA Vertical | 52.0897 | µrad |  | Coil Diameter | Max | 75.8 | mm
[6 empty] | Min | 60.2 | mm
Calculated Values | [5 empty] | Mean | 68 | mm
| Length-Diameter Product | 68.1015 | m2 |  |  |  | 0.068 | m
| Inverse Scale Factor | 1.084 | Hz |  | Coil Turns |  | 4671.6657
| Scale Factor | 0.9225 | sec |  | Total Loop Length |  | 1001.493 | m
| Eigenfrequency | 102.3138 | kHz |  | L*D |  | 68.1015 | m^2
[5 empty] | Loop Transit Time |  | 4.8869e-06 | s
Measured Values |  |  |  |  | Eigen Frequency |  | 102313.7754 | Hz
| Fractional Demodulation Word - IA Up | 0.00011318
| Equivalent Sagnac Phase - IA Up | 49.8576 | µrad |  |  |  | 102.3138 | kHz
| Fractional Demodulation Word - IA Down | -0.00013068 |  |  | SF |  | 0.9225 | s |  | J1 | 0.56682
| Equivalent Sagnac Phase - IA Down | -57.5702 | µrad |  | Inverse SF |  | 1.084 | Hz |  | J2 | 0.2497
| Rate Bias Offset | -0.83617 | deg/hr. | [6 empty] | Frac | 2.27
| Apparent Scale Factor | 0.95126 | sec |  | Inputs
| Apparent Eigenfrequency | 85.245 | kHz |  | ORC Latitude |  | 50.9376 | ° | Parameter |  | Value | Units
[5 empty] | Rate | IA Vertical | 11.6469 | °/hr | IA Up | Demod Word | 0.00011318 | unitless
[6 empty] | IA North | 9.4525 | °/hr |  | ∆φS | 49.8576 | µrad |  | 4.9858e-05 | 95520.6003
[5 empty] | ∆φS | IA Vertical | 0.0029845 | ° | IA Down | Demod Word | -0.00013068 | unitless |  | 5.2412e-05 | 10.8107
[7 empty] | 52.0897 | µrad |  | ∆φS | -57.5702 | µrad |  | 10.8107
[6 empty] | IA North | 0.0024222 | °/hr | True Rate |  | 11.6469 | °/hr
[7 empty] | 42.2754 | µrad |  |  | 56.4658 | µrad/s
[9 empty] | Apparent Scale Factor |  | 0.95126 | sec
[5 empty] | Detector |  |  |  | Rate Bias Offset |  | -4.0539 | µrad/s
[5 empty] | Power at Detector |  | -10 | dBm |  |  | -0.83617 | deg/hr
[7 empty] | 0.0001 | W
[7 empty] | 100 | µW |  |  | 52.4119
[5 empty] | Responsivitity |  | 1.05 | A/W |  |  | 10.8107
[5 empty] | TIA Gain (50Ω load) |  | 750 | V/A |  |  | -60.5197
[5 empty] | Expected Peak Voltage |  | 0.07875 | V |  |  | -12.4831
[7 empty] | 78.75 | mV
[5 empty] | DC Offset: Optical Power Off |  | <1mV | mV
[5 empty] | Avg Voltage: Modulation Off |  | 585.2 | mV
[5 empty] | Max Voltage: >π mod depth |  | 538 | mV
[5 empty] | Min Voltage: >π mod depth |  | 72 | mV
[5 empty] | Full Inteferogram Depth (max-min) |  | 466 | mV
[5 empty] | Peak to Peak Voltage: π/2 mod depth |  | 40 | mV
[5 empty] | Mod Amplitude Required to Null 2nd Harmonic |  | 3.368 | V
[5 empty] | J2 Root |  | 5.1356 | rad
[5 empty] | Desired Mod Depth, α (assume at eigen) |  | 0.5 | π rad
[5 empty] | Mod Amplitude for desired Mod Depth |  | 1.0301 | V
[5 empty] | 0th Bessel: J0(α) |  | 0.472
[5 empty] | 1st Bessel: J1(α) |  | 0.56682
[5 empty] | 2nd Bessel: J2(α) |  | 0.2497
[5 empty] | IA Vertical Fundamental (RMS): |  | -9.7291 | µV
[5 empty] | IA Vertical 2nd Harmonic (RMS): |  | -82.2796 | mV
2018-05-02 11:42
Detector Off, Power off Mean voltage =
<1mV
Detector On, Optical Power Off =
[10 empty] | Scope | mean | 8.8 | mV
[11 empty] | max | 10.6 | mV
[11 empty] | min | 7 | mV
Optical power on, Modulation Off =
[10 empty] | FLS | Setpoint | 6 | dBm
[11 empty] | Panel Reading | 5.97 | dBm
[10 empty] | Power Meter | Mean | -10.049 | dBm
[12 empty] | 98.8781 | µW
[10 empty] | Scope | mean | 47.25 | mV
[11 empty] | max | 49 | mV
[11 empty] | min | 45.4 | mV
Optical power on, Modulation Off =
[10 empty] | FLS | Setpoint | 9 | dBm
[11 empty] | Panel Reading | 8.99 | dBm
[10 empty] | Power Meter | Mean | -7.028 | dBm
[12 empty] | 198.244 | µW
[10 empty] | Scope | mean | 87.9 | mV
[11 empty] | max | 90.7 | mV
[11 empty] | min | 85.2 | mV
Optical power on, Modulation Off =
[10 empty] | FLS | Setpoint | 12 | dBm
[11 empty] | Panel Reading | 11.99 | dBm
[10 empty] | Power Meter | Mean | -3.9746 | dBm
[12 empty] | 400.4423 | µW
[10 empty] | Scope | mean | 167.2 | mV
[11 empty] | max | 172 | mV
[11 empty] | min | 162.5 | mV
Optical power on, Modulation Off =
[10 empty] | FLS | Setpoint | 15 | dBm
[11 empty] | Panel Reading | 15 | dBm
[10 empty] | Power Meter | Mean | -0.9213 | dBm
[12 empty] | 808.8537 | µW
[10 empty] | Scope | mean | 331 | mV
[11 empty] | max | 349 | mV
[11 empty] | min | 315 | mV
Optical power on, Modulation Off =
[10 empty] | FLS | Setpoint | 17.5 | dBm
[11 empty] | Panel Reading | 17.49 | dBm
[10 empty] | Power Meter | Mean | 1.6319 | dBm
[12 empty] | 1456.096 | µW
[10 empty] | Scope | mean | 585.2 | mV
[11 empty] | max | 604.2 | mV
[11 empty] | min | 566.2 | mV
Modulation On
[10 empty] | FG | Output Amp | 2.7 | Vp-p
[10 empty] | Detector | mean | 186 | mV
[11 empty] | High | 538 | mV
[11 empty] | Low | 72 | mV
[10 empty] | Modulation | Amplitude | 5.43 | Vp-p

## Setup
Constants
Speed of Light |  | 299792458 | m/s
Index |  | 1.468
Light Source
Mean Wavelength |  | 1547.21 | nm
|  | 1.5472e-06 | m
Coil
Coil Length from input |  | 1071 | m
Coil Diameter | Max | 104.46 | mm
| Min | 100.54 | mm
| Mean | 102.5 | mm
|  | 0.1025 | m
Coil Turns |  | 3325.9501
Total Loop Length |  | 1071 | m
L*D |  | 109.7775 | m^2
Loop Transit Time |  | 5.2444e-06 | s
Eigen Frequency |  | 95340.0073 | Hz
|  | 95.34 | kHz |  | Mod Depth | 0.5 | pi
SF |  | 1.487 | s |  | J1 | 0.56682
Inverse SF |  | 0.67248 | Hz |  | J2 | 0.2497
[5 empty] | Frac | 2.27
Inputs
ORC Latitude |  | 50.9376 | ° | Parameter |  | Value | Units
Rate | IA Vertical | 11.6469 | °/hr | IA Up | Demod Word | 0.00011318 | unitless
| IA North | 9.4525 | °/hr |  | ∆φS | 49.8576 | µrad |  | 4.9858e-05 | 95520.6003
∆φS | IA Vertical | 0.004811 | ° | IA Down | Demod Word | -0.00013068 | unitless |  | 5.2412e-05 | 10.8107
|  | 83.967 | µrad |  | ∆φS | -57.5702 | µrad |  | 10.8107
| IA North | 0.0039045 | °/hr | True Rate |  | 11.6469 | °/hr
|  | 68.1466 | µrad |  |  | 56.4658 | µrad/s
|  |  |  | Apparent Scale Factor |  | 0.95126 | sec
Detector |  |  |  | Rate Bias Offset |  | -4.0539 | µrad/s
Power at Detector |  | ? | dBm |  |  | -0.83617 | deg/hr
|  | #VALUE! | W
|  | #VALUE! | µW |  |  | 52.4119
Responsivitity |  | 1.05 | A/W |  |  | 10.8107
TIA Gain (50Ω load) |  | 750 | V/A |  |  | -60.5197
Expected Peak Voltage |  | #VALUE! | V |  |  | -12.4831
|  | #VALUE! | mV
DC Offset: Optical Power Off |  |  | mV
Avg Voltage: Modulation Off |  | 267 | mV | 267 | mV | <-- MEASURED DC POWER AT 17.5dBm & 10^4 gain setting
Max Voltage: >π mod depth |  |  | mV
Min Voltage: >π mod depth |  |  | mV
Full Inteferogram Depth (max-min) |  |  | mV
Peak to Peak Voltage: π/2 mod depth |  |  | mV
Mod Amplitude Required to Null 2nd Harmonic |  |  | V
J2 Root |  | 5.1356 | rad
Desired Mod Depth, α (assume at eigen) |  | 0.5 | π rad
Mod Amplitude for desired Mod Depth |  | 0 | V
0th Bessel: J0(α) |  | 0.472
1st Bessel: J1(α) |  | 0.56682
2nd Bessel: J2(α) |  | 0.2497
IA Vertical Fundamental (RMS): |  | 0 | µV
IA Vertical 2nd Harmonic (RMS): |  | 0 | mV
2018-05-02 11:42
Detector Off, Power off Mean voltage =
<1mV
Detector On, Optical Power Off =
[5 empty] | Scope | mean | 8.8 | mV
[6 empty] | max | 10.6 | mV
[6 empty] | min | 7 | mV
Optical power on, Modulation Off =
[5 empty] | FLS | Setpoint | 6 | dBm
[6 empty] | Panel Reading | 5.97 | dBm
[5 empty] | Power Meter | Mean | -10.049 | dBm
[7 empty] | 98.8781 | µW
[5 empty] | Scope | mean | 47.25 | mV
[6 empty] | max | 49 | mV
[6 empty] | min | 45.4 | mV
Optical power on, Modulation Off =
[5 empty] | FLS | Setpoint | 9 | dBm
[6 empty] | Panel Reading | 8.99 | dBm
[5 empty] | Power Meter | Mean | -7.028 | dBm
[7 empty] | 198.244 | µW
[5 empty] | Scope | mean | 87.9 | mV
[6 empty] | max | 90.7 | mV
[6 empty] | min | 85.2 | mV
Optical power on, Modulation Off =
[5 empty] | FLS | Setpoint | 12 | dBm
[6 empty] | Panel Reading | 11.99 | dBm
[5 empty] | Power Meter | Mean | -3.9746 | dBm
[7 empty] | 400.4423 | µW
[5 empty] | Scope | mean | 167.2 | mV
[6 empty] | max | 172 | mV
[6 empty] | min | 162.5 | mV
Optical power on, Modulation Off =
[5 empty] | FLS | Setpoint | 15 | dBm
[6 empty] | Panel Reading | 15 | dBm
[5 empty] | Power Meter | Mean | -0.9213 | dBm
[7 empty] | 808.8537 | µW
[5 empty] | Scope | mean | 331 | mV
[6 empty] | max | 349 | mV
[6 empty] | min | 315 | mV
Optical power on, Modulation Off =
[5 empty] | FLS | Setpoint | 17.5 | dBm
[6 empty] | Panel Reading | 17.49 | dBm
[5 empty] | Power Meter | Mean | 1.6319 | dBm
[7 empty] | 1456.096 | µW
[5 empty] | Scope | mean | 585.2 | mV
[6 empty] | max | 604.2 | mV
[6 empty] | min | 566.2 | mV
Modulation On
[5 empty] | FG | Output Amp | 2.7 | Vp-p
[5 empty] | Detector | mean | 186 | mV
[6 empty] | High | 538 | mV
[6 empty] | Low | 72 | mV
[5 empty] | Modulation | Amplitude | 5.43 | Vp-p

## Eigen & Mod Tuning
Eigenfrequency Tuning
Start Time
Modulation Off
ASE Light Source Setup
C-Band = | 17.5 | dBm
L-Band = | 0 | dBm
Measured Power at Circulator L3
Power Out (unmodulated) = | -12.9 | dBm
| 0.051286 | mW
| 51.2861 | µW
Gyro configuration | Optics plate fully closed
Coil IA Vertical UP in IGLOO
Temperature Acquisition | Not On
Detector config
Gain | 10^4 L
Coupling | DC
Exit filter | 10MHz
DC Voltage from Detector (DVM) = | 0.973 | V
Mod coupled to PM1, PM2, and Scope on 1Mohm channel impedance
FG Configuration
Found Coarse eigenfrequency
Amplitude = | 2.2 | Vp-p
Frequency = | 85.4 | kHz
Waveform | Square
Load | Hi Z | Ohm
Modulated Detector Voltage
Mean | 245 | mV
Max | 342 | mV
Min | 233 | mV
High | 332 | mV
Low | 241 | mV
Scope Modulation Amplitude
Amplitude = | 2.017 | Vp-p
Cheanged to SINE modulation | X
Modulated Detector Voltage
Mean | 358.5 | mV
Max | 499 | mV
Min | 233 | mV
High | 487.8 | mV
Low | 243 | mV
Amplitude = | 2.16 | V
Mod coupled to PM1, PM2, and Scope on 1Mohm channel impedance + Lock-In Amplifier 2nd Harmonic
2nd Harmonic On
Amplitude | 0 | V
1st harmonic Mod Amplitude Required to Null 2nd Harmonic | 7.1 | V
J2 Root | 5.1356 | rad
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
Mod Amplitude for desired Mod Depth | 2.1716 | V
102181.2
FG Configuration
Amplitude = | 2.172 | Vp-p
Frequency = | 85.4 | kHz
Waveform | Sine
Load | Hi Z | Ohm
Pretuning Avg 1st Harmonic Demod | 1.4666e-05 | V
2nd Harmonic On
Amplitude | 1.1 | V
First Eigen Frequency | 85245 | Hz
Removed Second Harmonic | X
Measured correct 1st harmonic phase - by twisting table while monitoring y-component of 1st harmonic rate
Phase at which no rotation is in y = | 174.332 | °
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 7.1 | V
2nd Root | 11.64 | V
3rd Root | 16.06 | V
J2 1st Root | 5.1356 | rad
J2 2nd Root | 8.4172 | rad
J2 3rd Root | 11.6198 | Rad
Mod Amplitude from 1st Root | 2.1716 | V
Mod Amplitude from 2nd Root | 2.1722 | V
Mod Amplitude from 3rd Root | 2.171 | V
Mod Amplitude for desired Mod Depth | 2.1716 | V
Set FG
FG Configuration
Amplitude = | 2.172 | Vp-p
Frequency = | 85245 | Hz
Waveform | Sine
Load | Hi Z | Ohm
Phase | 0 | °
2nd Harmonic On
2nd Harmonic Amplitude | 1.1 | V
REMEASURED EigenFrequency = | 85304 | Hz
SAVED AS | SETUP 2
2019-05-31 14:52 | REMEASURING EIGENFREQUENCY AND MODULATION DEPTH
Modulation on PM2 ONLY
Photodetector going straight to LIA
Coil in Igloo - IA vertical UP
LIA on 10Hz acquisition rate
Phase Mod 2
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 13.69 | V
2nd Root | 11.64 | V
3rd Root | 16.06 | V
J2 1st Root | 5.1356 | rad
Mod Amplitude from 1st Root | 4.1873 | V
Mod Amplitude for desired Mod Depth | 4.1873 | V |  | 2.0936
Phase modulator V pi | 8.3745 | Vp-p
Modulation on PM1 ONLY
Phase Mod 1
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 13.69 | V
2nd Root | 11.64 | V
3rd Root | 16.06 | V
J2 1st Root | 5.1356 | rad
Mod Amplitude from 1st Root | 4.1873 | V
Mod Amplitude for desired Mod Depth | 4.1873 | V
Phase modulator V pi | 8.3745 | Vp-p
Applied same modulation to Both PMs
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 7.104 | V
2nd Root | 11.64 | V
3rd Root | 16.07 | V
J2 1st Root | 5.1356 | rad
J2 2nd Root | 8.4172 | rad
J2 3rd Root | 11.6198 | Rad
Mod Amplitude from 1st Root | 2.1729 | V
Mod Amplitude from 2nd Root | 2.1722 | V
Mod Amplitude from 3rd Root | 2.1724 | V
Mod Amplitude for desired Mod Depth | 2.1725 | V
Repeated
LIA SIGNAL OUT set demodulator 4 (2nd harmonic) - AMPLITUDE SET TO ZERO
Signal OUT from LIA (2nd harmonic modulation) connected to PM2 & Oscilliscope
FG 1st harmonic modulation connected to PM1 & Oscilliscope
Phase Mod 1
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 14.22 | V
2nd Root | 11.64 | V
3rd Root | 16.06 | V
J2 1st Root | 5.1356 | rad
Mod Amplitude from 1st Root | 4.3494 | V
Mod Amplitude for desired Mod Depth | 4.3494 | V
Phase modulator V pi | 8.6987 | Vp-p
Disconnected PM1 and
Set FG Modulation Depth to = | 42652 | Hz
Now 2nd Harmonic modulation is near eigenfrequency
Adjusted OUTPUT AMPLITUDE in LIA panel to match 2nd harmonic signal magnitude on Oscilliscope to that of the direct FG modulation
LIA OUTPUT AMPLITUDE = | 2.2 | Vpk
Turned off 2nd harmonic modulation from LIA
FG Configuration
Amplitude = | 4.349 | Vp-p
Frequency = | 85304 | Hz
Waveform | Sine
Load | Hi Z | Ohm
Phase | 0 | °
PM1 connected to FG and oscilliscope
With these settings
LIA demod X = | 1.4382e-05 | V
Turned on 2nd harmonic modulation
Tuned to new eigenfrequency
2nd Harmonic Amplitude | 2.2 | V
SET 2nd Harmonic DEMODULATOR PHASE To 0
REMEASURED EigenFrequency = | 85304 | Hz
| 85279 | 0.99971 | 0.00029307 | 293.0695
APPARENT EIGENFREQUENCY CHANGES WITH Phase of second harmonic modulation signal!!
Y demod of 1st harmonic is highly variable with this phase as well
Still no rotation in y component of demod though.
REMEASURED EigenFrequency = | 85282 | Hz
for LIA 2ndf harmonic phase = | 75.341
2nd harmonic Magnitude = | 1.1 | on PM2
2020-02-04 16:42 | Gyro Restart after sitting in lab for 9 months
Eigenfrequency Tuning
Start Time | 2020-02-04 16:42
Modulation Off
ASE Light Source Setup
C-Band = | 17.5 | dBm
L-Band = | 0 | dBm
Measured Power at Circulator L3
Power Out (unmodulated) = | -13.73 | dBm
| 0.042364 | mW
| 42.3643 | µW
Applied same modulation to Both PMs
Desired Mod Depth, α (assume at eigen) | 0.5 | π rad
1st harmonic Mod Amplitude Required to Null 2nd Harmonic (1st Root) | 7.109 | V
2nd Root | 11.65 | V
3rd Root | 16.09 | V
J2 1st Root | 5.1356 | rad
J2 2nd Root | 8.4172 | rad
J2 3rd Root | 11.6198 | Rad
Mod Amplitude from 1st Root | 2.1744 | V
Mod Amplitude from 2nd Root | 2.1741 | V
Mod Amplitude from 3rd Root | 2.1751 | V
Mod Amplitude for desired Mod Depth | 2.1745 | V
2020-02-07 19:52 | POWER SWEEP PROFILE
measured power on circulator port 3 for these powers | Power Meter Reading - Circulator Port 3 (dBm)
ASE Output Power (front panel - dBm) |  | Unmodulated |  | Modulated at 2.175 Vp-p
Light Source |  | Detector
(dBm) | (mW) | (dBm) | (µW)
17.5 | 56.2341 | -13.93 | 40.4576 | -15.26
16 | 39.8107 | -15.474 | 28.3531 | -16.8
14.5 | 28.1838 | -17 | 19.9526
13 | 19.9526 | -18.54 | 13.9959
11.5 | 14.1254 | -20.07 | 9.8401
10 | 10 | -21.63 | 6.8707
8.5 | 7.0795 | -23.17 | 4.8195
7 | 5.0119 | -24.69 | 3.3963
5.5 | 3.5481 | -26.25 | 2.3714
4 | 2.5119 | -27.81 | 1.6558
2.5 | 1.7783 | -29.31 | 1.1722
1 | 1.2589 | -30.7 | 0.85114
Note on B126: Taranta A.A.: Truly REMEASURED - noit just copied from above
Note on B154: Taranta A.A.: Truly REMEASURED - noit just copied from above

## Power Sweep
|  |  |  | 13 Hz Mean | 11.6469 | <--True Rate
| Detector Power |  | ARW | IA Vertical Reported Rate | True Bias Offset
| dBm | µW | µdeg/rt-hr | deg/hr | deg/hr
1 | -13.93 | 40.4576 | 2471 | 11.596 | -0.050904
2 | -15.474 | 28.3531 | 2295 | 11.7509 | 0.104
3 | -17 | 19.9526 | 2196 | 11.8865 | 0.2396
4 | -18.54 | 13.9959 | 2241 | 11.9845 | 0.3376
5 | -20.07 | 9.8401 | 2297 | 12.0447 | 0.3978
6 | -21.63 | 6.8707 | 2609 | 12.0705 | 0.4236
7 | -23.17 | 4.8195 | 2902 | 12.1003 | 0.4534
8 | -24.69 | 3.3963 | 3467 | 12.1176 | 0.4707
9 | -26.25 | 2.3714 | 4104 | 12.0983 | 0.4514
10 | -27.81 | 1.6558 | 5448 | 12.1146 | 0.4677
11 | -29.31 | 1.1722 | 7202 | 12.0851 | 0.4382
12 | -30.7 | 0.85114 | 9953 | 12.045 | 0.3981
13 | -13.93 | 40.4576 | 2427 | 11.6293 | -0.017604
14 | -15.474 | 28.3531 | 2212 | 11.7831 | 0.1362
15 | -17 | 19.9526 | 2264 | 11.916 | 0.2691
16 | -18.54 | 13.9959 | 2245 | 11.9872 | 0.3403
17 | -20.07 | 9.8401 | 2398 | 12.0485 | 0.4016
18 | -21.63 | 6.8707 | 2541 | 12.0785 | 0.4316
19 | -23.17 | 4.8195 | 3065 | 12.101 | 0.4541
20 | -24.69 | 3.3963 | 3547 | 12.1345 | 0.4876
21 | -26.25 | 2.3714 | 4280 | 12.1323 | 0.4854
22 | -27.81 | 1.6558 | 5641 | 12.1212 | 0.4743
23 | -29.31 | 1.1722 | 7419 | 12.0941 | 0.4472
24 | -30.7 | 0.85114 | 9781 | 12.0623 | 0.4154

## Run Log
| Basic Config | [7 empty] | LIA Software |  | Thermal Data |  | Gyro Data |  |  |  | 1st Harmonic Demod (1) |  |  |  | 2nd Harmonic Demod (4) |  |  |  | Reference PLL (Dem 2) |  |  |  | LIA Settings |  |  |  |  | Modulation |  |  |  | Pt100 RTDs |  |  |  | Thermistors |  | Detector | [5 empty] | FLS |  | Power at Circulator Port 3
| Start Date / Time | Pre-Test Comment | Stop Date/Time | Post-Test Comment | Gryo Config | IA | Thermal Program | Heater Control | LabOne Version | LIA Config File | Fileset | Rate (Hz) | MFLI Acquisition Flags | Session Folder | Stream Folder | Downloaded Fileset | 3dB BW (Hz) | Order | Sinc | Rate Phase (°) | 3dB BW (Hz) | Order | Sinc | Phase (°) | 3dB BW (Hz) | Order | Sinc | Signal | Data Rate (Hz) | Range | AC | Float | 50Ω | Amp (Vpp) | Frequency (kHz) | Load | Signal | 2A (Blue) | 2B (Red) | 2C (Wht) | 2D (Y) | T1 | T2 | Type | Knob | H/L | BW | Coupling | Term | C-Band (dBm) | L-Band (dBm) | Unmodulated (dBm) | Modulated (dBm) | Data Folder on Network
A | 2019-05-26 19:09 | Post-Tuning | 2019-05-26 19:16 | Test Halted Due to Bad Run of Temp Logger | Coil Mounted in Igloo, Routed through foam umbilical | Up | N/A | Uncontrolled |  | 1Hz_MCIFOG1 | 20190526_190451_Temperature_Log.txt | 1 |  |  |  | data_dev3336_20190526_200704 | 80.13m | 2 | On | 174.332 | 80.13m | 2 | On | 75.341 |  |  |  |  | 817.4m | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
B | 2019-05-26 19:27 | Post-Tuning Repeat | 2019-05-28 14:18 | Opened Optics enclosure after end of run | Coil Mounted in Igloo, Routed through foam umbilical | Up | N/A | Uncontrolled |  | 1Hz_MCIFOG1 | 20190526_192513_Temperature_Log.txt | 1 |  |  |  | data_dev3336_20190528_144155 | 80.13m | 2 | On | 174.332 | 80.13m | 2 | On | 75.341 |  |  |  |  | 817.4m | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
C | 2019-05-28 18:22 | Overnight Bias | 2019-05-29 11:58 | Normal Completion | Both enclosures opened ~2hrs prior). Coil Mounted in Igloo, Routed through foam umbilical | Up | N/A | Uncontrolled |  | 1Hz_MCIFOG1 | 20190528_181938_Temperature_Log.txt | 1 |  |  |  | data_dev3336_20190529_124716 | 80.13m | 2 | On | 174.332 | 80.13m | 2 | On | 75.341 |  |  |  |  | 817.4m | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
D | 2019-05-29 12:20 | 7 kHz Noise Run | 2019-05-29 12:31 | Normal Completion | Coil Mounted in Igloo, Routed through foam umbilical | Up | N/A | Uncontrolled |  | 7kHz_MCIFOG1 | N/A | N/A |  |  |  | data_dev3336_20190529_131914 | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 |  |  |  |  | 6.696k | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
E | 2019-05-29 16:21 | Shut off: Oscilliscope & Zaber Stages. Opened and Reclosed Optics plate + Igloo, completely loosened coil mount screws and removed lead tape at coil mount. Igloo top weighted down with thin copper plate. | 2019-05-30 08:16 | Stopped Test Remotely | Coil Mounted in Igloo, Routed through foam umbilical, Igloo top secured with thin copper plate, coil hub resting on 150mm inner plate - not screwed in | Up | N/A | Uncontrolled |  | 1Hz_MCIFOG1 | 20190529_161905_Temperature_Log.txt | 1 |  |  |  | data_dev3336_20190530_090809 | 80.13m | 2 | On | 174.332 | 80.13m | 2 | On | 75.341 |  |  |  |  | 817.4m | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
F | 2019-05-30 11:55 | Heading Swing: 12:39:41 - Opened coil enclosure, inverted IA and reclosed coil enclosure with cover mounted on top again. 12:46:35 - Done opening enclosure - testing continues | 2019-05-30 13:34 | Normal Completion | Coil Mounted in Igloo, Routed through foam umbilical, Igloo top secured with thin copper plate, coil hub resting on 150mm inner plate - not screwed in | Up to Down | N/A | Uncontrolled |  | 10Hz_MCIFOG1 | 20190530_115251_Temperature_Log.txt | 2 |  |  |  | data_dev3336_20190530_142629 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
G | 2019-05-30 13:45 | Heading Swing at 3π/4 mod depth: 15:20:28 - Opened coil enclosure, inverted IA and reclosed coil enclosure with cover mounted on top again. | 2019-05-31 08:25 | Normal Completion | Coil Mounted in Igloo, Routed through foam umbilical, Igloo top secured with thin copper plate, coil hub resting on 150mm inner plate - not screwed in | Down to Up | N/A | Uncontrolled |  | 10Hz_MCIFOG1 | 20190530_134251_Temperature_Log.txt | 2 |  |  |  | data_dev3336_20190531_091728 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 3.257 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo Air | Coil | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
H | 2019-05-31 18:13 | FS Coil T-dot | 2019-06-03 07:15 | Normal Completion. Left Heater Setpoint at 30°C with coil inside blue enclosure | Coil DISMOUNTED - Freestanding inside Blue thermal enclosure. Coil and coil enclosure thermistors now monitoring air inside blue enclosure. Coil thermistor NOT touching surface of coil. Both Phase modulators connected to small BNCs with tee connected to barrel connected to FG. No Connections to Oscilliscope and Oscilliscope off. | Up | T-Dot | T-Dot |  | 1Hz_MCIFOG1 | 20190531_180858_Temperature_Log.txt | 1 |  |  |  | data_dev3336_20190603_080740 | 80.13m | 2 | On | 174.332 | 80.13m | 2 | On | 75.341 |  |  |  |  | 817.4m | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
I | 2019-06-03 07:34 | FS Coil 10Hz noise at 30°C 03/06/2019 08:28:19 - Changed LIA to GROUNDED | 2019-06-03 09:50 | Completed Remotely | Same config as above (FS coil on hotplate in blue enclosure) | UP | 30°C Setpoint | 30°C Static Setpoint |  | 10Hz_MCIFOG1 | 20190603_073200_Temperature_Log.txt | 2 |  |  |  | data_dev3336_20190603_104045 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On& OFF | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
J | 2019-06-03 11:28 | FS Coil 10Hz noise with Coil drifting back from 30°C to room temperature - Heater Shut off | 2019-06-04 10:09 | Completed Remotely | Same config as above (FS coil on hotplate in blue enclosure) | Up | N/A | Off - Drifting down from 30°C |  | 10Hz_MCIFOG1 | 20190603_112642_Temperature_Log.txt | 2 |  |  |  | data_dev3336_20190604_110121 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
K | 2019-06-04 11:56 | FS Coil 10Hz Bias at 3pi/4 mod depth. FLS interlock broken twice in the 15 min. prior to test start. | 2019-06-05 08:59 | Normal Completion | Same config as above (FS coil on hotplate in blue enclosure) | Up | N/A | Off |  | 10Hz_MCIFOG1 | 20190604_115455_Temperature_Log.txt | 2 |  |  |  | data_dev3336_20190605_095130 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 3.257 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
L | 2019-06-05 09:11 | No Light Test - High Speed at Full LIA Scale at 3pi/4 modulation | 2019-06-05 09:20 | Normal Completion | Same config as above (FS coil on hotplate in blue enclosure) | Up | N/A | Off |  | 7kHz_MCIFOG1 | 20190605_091029_Temperature_Log.txt | 2 |  | session_20190526_172112_00 | stream_011 | 1st: data_dev3336_20190605_101150 2nd: data_dev3336_20190605_103914 | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 |  |  |  |  | 6.696k | 0.3 | On | On | Off | 3.257 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 0 | 0
M | 2019-06-05 09:25 | No Light Test - High Speed at Full LIA Scale at pi/2 modulation | 2019-06-05 09:36 | Normal Completion | Same config as above (FS coil on hotplate in blue enclosure) | Up | N/A | Off |  | 7kHz_MCIFOG1 | 20190605_092404_Temperature_Log.txt | 2 |  | session_20190526_172112_00 | stream_012 | data_dev3336_20190605_102848 | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 |  |  |  |  | 6.696k | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 0 | 0
N | 2019-06-05 10:02 | No Light Test - 10 Hz at Full LIA Scale at pi/2 modulation | 2019-06-06 11:01 | Completed Remotely | Same config as above (FS coil on hotplate in blue enclosure) | Up | N/A | Off |  | 10Hz_MCIFOG1_new | 20190605_100032_Temperature_Log.txt | 2 | None | session_20190526_172112_00 | stream_013 | data_dev3336_20190606_115340 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Blue Enclosure Air | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 0 | 0
O | 2019-06-06 16:34 | FS Coil in optics enclosure - 10Hz Overnight at pi/2 | 2019-06-07 09:23 | Completed Remotely | FS Coil seated on inner routing platform INSIDE optics enclosure. Zaber stages on, and acquiring data on ARF in Blue box | Up | N/A | Off |  | CONFIG FILE OUT OF DATE | N/A | N/A | None | session_20190606_172040_02 | stream_000 | data_dev3336_20190607_101638 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 |  |  |  |  | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
P | 2019-06-07 16:32 | FS Coil in optics enclosure - 10Hz Weekend Bias at pi/2 | 2019-06-10 10:09 | Completed Remotely | FS Coil seated on inner routing platform INSIDE optics enclosure. Running General Photonics Thermal Cycling concurrently: PSGA-101 powered on, Zaber stages on, Thermal cycling & acquiring data on ARF in Blue box. Both OSAs and Walics rack OFF. | Up | N/A | N/A |  | 10Hz_MCIFOG1_v3 | N/A | N/A | None | session_20190607_171623_04 | stream_000 | data_dev3336_20190610_110212 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
Q | 2019-06-10 18:33 | FS Coil in Optics enclosure - 10 Hz. Music was on at beginning of test No Change from Prior configuration, Just Hit record. Interlock broken and Light source restarted in first 30 sec of test data acquisiton. | 2019-06-12 09:47 | Normal Completion | FS Coil seated on inner routing platform INSIDE optics enclosure. Running General Photonics Thermal Cycling concurrently: PSGA-101 powered on, Zaber stages on, Thermal cycling & acquiring data on ARF in Blue box. Both OSAs and Walics rack OFF. | Up | N/A | N/A |  | Not loaded - same as above | N/A | N/A | None | session_20190607_171623_04 | stream _001 | data_dev3336_20190612_103948 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
R | 2019-06-12 11:05 | Started run and attempted to start acquisition of thermal data with PTC10 logger while already reserved by Pol Test - DID NOT WORK |  | Bad Test - INCOMPLETE | Same config as above: FS Coil seated on inner routing platform INSIDE optics enclosure. Running General Photonics Thermal Cycling concurrently: PSGA-101 powered on, Zaber stages on, Thermal cycling & acquiring data on ARF in Blue box. Both OSAs and Walics rack OFF. | Up | N/A | N/A | LabOne® Release Version 18.12.59443 | 3Hz_MCIFOG1.xml | 20190612_110214_Temperature_Log.txt | 1 | BAD TEST (ONLY 5 min. long) DATA NOT DOWNLOADED |  |  |  | 300.1m | 2 | On | 174.332 | 300.1m | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 3.27 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Igloo air (Open to Lab) | Igloo air (Open to Lab) | Plate (no gyro inside) | blue enclosure air (no Gyro inside) | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
S | 2019-06-12 11:31 | FS Coil in Optics Enclosure at 3Hz. Checked Data with interim download: data_dev3336_20190613_121728 on 13/06/2019 11:28:04 | 24/06/2019 ~10:30 | Normal Completion | Same config as above except ZABER power off: FS Coil seated on inner routing platform INSIDE optics enclosure. Running General Photonics Thermal Cycling concurrently: PSGA-101 powered on, Zaber stages OFF, Thermal cycling & acquiring data on ARF in Blue box. Both OSAs and Walics rack OFF. | Up | N/A | N/A | LabOne® Release Version 18.12.59443 | 3Hz_MCIFOG1.xml | N/A | N/A | None | session_20190612_113602_01 | stream_001 | data_dev3336_20190624_112110 | 300.1m | 2 | On | 174.332 | 300.1m | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 3.27 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
T | 2020-02-04 17:29 | Restart overnight test. | 2020-02-05 08:01 | Normal Completion | Restart overnight bias after long hiatus. 9 months on table but without testing. Coil resting on top of optics plate, isolated inside black enclosure. | Up | N/A | N/A | LabOne® Release Version 18.12.59443 | 10Hz_MCIFOG1_v3 | 20200204_172355_Temperature_Log.txt | 1 | None | session_20200204_184736_00 | stream_000 | data_dev3336_20200205_100414 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.175 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Lab | Optics Plate Exit | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0 | -13.74 | -15.08
U | 2020-02-05 14:58 | ARW vs. optical power test. Used labview program to change optical power output of manlight source every 2 hours. Power setting was C-band only starting at +17.5dBm and going down to +1 dBm in 1.5dB steps and then starting back from the beginning. Set to 17.3 dBm manually at very beginning to confirm communication from GPIB line (watched value change to 17.5) | 2020-02-07 15:20 | Completed Remotely. Ending power on ASE Source was set to 17.5dBm. Shut off ASE Source Control software at 07/02/2020 15:59:59 | Gyro isolated inside optics enclosure. FS Coil sitting on upper plate. Optical power changed by labview program "Manlight Controller Main". | Up | N/A | N/A | LabOne® Release Version 18.12.59443 | 10Hz_MCIFOG1_v3 | 20200205_144810_Temperature_Log.txt | 1 | None | session_20200205_164834_01 | stream_000 | data_dev3336_20200207_173706 | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.175 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Lab | Optics Plate Exit | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0 | -13.79 | -15.11
V | 2020-02-07 20:06 | Set to 16dBm FLS power for weekend run. 1 Hz rate at lowered LIA Range (0.1V instead of 0.3V) | 2020-02-10 09:52 | Completed Remotely | Gyro isolated inside optics enclosure. FS Coil sitting on upper plate. Optical power of FLS set to 16 dBm instead of normal 17.5 dBm | Up | N/A | N/A | LabOne® Release Version 18.12.59443 | 1Hz_MCIFOG1 | 20200207_200102_Temperature_Log.txt | 1 | None | session_20200207_220058_02 | stream_000 | data_dev3336_20200210_115529 | 80.13m | 2 |  | 174.332 | 80.13m | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 817.4m | 0.1 | On | On | Off | 2.175 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Lab | Optics Plate Exit | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 16 | 0 | -15.48 | -16.81
W | 2020-02-11 10:46 | T-Dot testing at various reduced ramp rates and ranges. FS coil contacting plate surface inside thermal enclosure | 2020-02-14 08:28 | Paused program on PTC10 rather than completely stopping | FS Coil isolated inside thermal enclosure. Leads routed, hanging in air from optics enclosure to thermal enclosure (~14 in. of bifilar lead). | Up | "FOG_M1" (30-60 cycling at various ramps) | T-Dot | LabOne® Release Version 18.12.59443 | 3Hz_MCIFOG1 | 20200211_104311_Temperature_Log.txt | 3 | None | session_20200211_123512_04 | stream_000 | data_dev3336_20200214_103157 | 300.1m | 2 | On | 174.332 | 300.1m | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 3.27 | 0.3 | On | On | Off | 2.175 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Copper Plate | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0 | -13.29 | -14.62
X | 2020-02-14 08:46 | Continuation of previous T-Dot Test, started with Coil at 60°C in midst of thermal program. Had to perform with a new program, so started acquisition during coil thermal downramp | 2020-02-17 09:07 | Completed Remotely | FS Coil isolated inside thermal enclosure. Leads routed, hanging in air from optics enclosure to thermal enclosure (~14 in. of bifilar lead). | Up | "FOG_M2" (30-60 cycling at various ramps) | T-Dot | LabOne® Release Version 18.12.59443 | 3Hz_MCIFOG1 | 20200214_083505_Temperature_Log.txt | 3 | None | session_20200211_123512_04 | stream_001 | data_dev3336_20200217_111021 | 300.1m | 2 | On | 174.332 | 300.1m | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 3.27 | 0.3 | On | On | Off | 2.175 | 85.304 | Hi-Z | Both PMs w/ unterminated BNC Tee | Optics Plate | Lab | Copper Plate | Air near Coil | Plate | blue enclosure air | Femto OE-300-IN-01 | 10^4 | L | 10 MHz | DC | 50Ω @LIA | 17.5 | 0
Y | 2020-08-14 13:56 | No Light, No Gyro, No IOC, hooked up on SWIMA Setup for detector and LIA noise characterization. No Light, PD Powered on, LIA at full-scale | 2020-08-14 14:01 | Completed Normally | 7kHz No Light - PD ON - 0.3V LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | Stream_003 | LOADED DIRECTLY VIA NEW SW: \\filestore.soton.ac.uk\users\aat1f15\mydocuments\Zurich Instruments\LabOne\WebServer | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 6.696k | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 NO LIGHT | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 7kHz No Light - PD ON - 0.3V LIA Scale
Z | 2020-08-14 14:15 | Same as above, but LIA autoscaled | ~14/08/2020 14:20:17 | Completed Normally | 7kHz No Light - PD ON - 1mV LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | Stream_004 |  | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 6.696k | 1mV | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 NO LIGHT | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 7kHz No Light - PD ON - 1mV LIA Scale
AA | 14/08/2020 ~14:25 | Same as above but detector powered off, and LIA Rescaled to 0.3V again | 14/08/2020 ~14:30 | Completed Normally | 7kHz No Light - PD OFF - 0.3V LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | Stream_006 |  | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 6.696k | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 No Light, Powered OFF | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 7kHz No Light - PD OFF - 0.3V LIA Scale
AB | 14/08/2020 ~14:33 | Same as above but LIA Autoscaled | 14/08/2020 ~14:38 | Completed Normally | 7 kHz No Light - PD OFF - 1mV LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | Stream_007 |  | 599.9 | 2 | On | 174.332 | 599.9 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 6.696k | 1mV | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 No Light, Powered OFF | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 7kHz No Light - PD OFF - 1mV LIA Scale
AC | 14/08/2020 ~14:55 | Overnight 10 Hz Detector on, LIA Rescaled to 0.3V again | 2020-08-15 12:22 | Completed Remotely | 10 Hz No Light - PD ON - 0.3V LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | Stream_009 |  | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 NO LIGHT | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 10Hz Overnight No Light - PD ON - 0.3V LIA Scale
AD | 2020-08-15 12:32 | Overnight 10 Hz Detector on, LIA Rescaled to 1mV | 18/08/2020 ~1400 | Completed Normally | 10 Hz No Light - PD ON - 1mV LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | Signal Input Overload (OVI) | session_20200814_130717_00 | stream_010 | Saving direct to filestore (As above) | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 1mV | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 NO LIGHT | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200814 10Hz Overnight No Light - PD ON - 1mV LIA Scale
AE | 2020-08-18 14:43 | Overnight 10 Hz Detector OFF, LIA at 1mV | 2020-08-19 11:33 | Completed Remotely | 10 Hz No Light - PD OFF - 1mV LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | Signal Input Overload (OVI) | session_20200814_130717_00 | stream_011 | Saving direct to filestore (As above) | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 1mV | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 No Light, Powered OFF | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200818 10Hz Overnight No Light - PD OFF - 1 mV LIA Scale
AF | 2020-08-19 11:36 | Overnight 10 Hz Detector OFF, LIA at 0.3V | 25/09/2020 ~14:10 | Completed Normally | 10 Hz No Light - PD OFF - 0.3V LIA Scale | N/A | N/A | N/A | LabOne® Release Version 20.01.1818 | N/A | N/A | N/A | None | session_20200814_130717_00 | stream_012 | Saving direct to filestore (As above) | 1 | 2 | On | 174.332 | 1 | 2 | On | 75.341 | 29.43 | 4 | Off | Aux In 1 | 13.08 | 0.3 | On | On | Off | 2.172 | 85.304 | Hi-Z | Not in IOC - in SWIMA TLS | N/A | N/A | N/A | N/A | N/A | N/A | Femto OE-300-IN-01 No Light, Powered OFF | 104 (104 / 105) | L | 10 MHz | DC | 50Ω @LIA | N/A | N/A | N/A | N/A | 20200818 10Hz Whole Month - No Light - PD OFF - 300 mV LIA Scale
AG
AH
AI
AJ
AK
AL
AM
AN
AO
AP
AQ
AR
AS
AT
AU
AV
AW
AX
AY
AZ
BA
BB
Note on AI9: Taranta A.A.: 3pi/4
Note on AG11: Taranta A.A.: Changed partway through run
Note on Q14: Taranta A.A.: Downloaded Data again after noticing additional stream file in local (LIA) folder
Note on K16: Taranta A.A.: Original 10Hz_MCIFOG1 config file saved over with 7kHz settings -> renamed to corrupted and created this new 10Hz file
Note on AI22: Taranta A.A.: Feb 4, 2020 Newly measured from 2nd harmonic null points
Note on AI23: Taranta A.A.: Feb 4, 2020 Newly measured from 2nd harmonic null points
Note on AI24: Taranta A.A.: Feb 4, 2020 Newly measured from 2nd harmonic null points
Note on AY24: Taranta A.A.: Changed to minimum ARW power observed during power sweep
Note on AI25: Taranta A.A.: Feb 4, 2020 Newly measured from 2nd harmonic null points
Note on AI26: Taranta A.A.: Feb 4, 2020 Newly measured from 2nd harmonic null points

## Test Proc
|  | TEST | [5 empty] | = N/A |  |  | X | = COMPLETE
Step | Description | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | Y | Z | AA | AB | AC | AD | AE | AF | AG | AH | AI | AJ | AK | AL | AM | AN | AO | AP | AQ | AR | AS | AT | AU | AV | AW | AX | AY | AZ | BA | BB
0 | Shut off ALL INSTRUMENTS AND WATER | bar: TEST A to = COMPLETE AL
1 | Open MFLI | bar: TEST A to = COMPLETE AF
2 | Record LabOne Version (Config->WebServer->About) | bar: TEST A to = COMPLETE AF
3 | In Config check "preference" and load Appropriate file, e.g. "10Hz" | bar: TEST A to = COMPLETE AF
4 | Confirm Demod = ~15µV, Int = ~85mV | bar: TEST A to = COMPLETE AL
5 | Complete Run Log with current settings | bar: TEST A to = COMPLETE AF
6 | In plotter set window length to appropriate time: 2 Hrs. for 1sec data 30 min. for 3Hz & 10 Hz data 5 min. for High Speed data For new version - set save interval | bar: TEST A to = COMPLETE AF
7 | Start PTC Temp Logger | bar: TEST A to = COMPLETE AL
8 | Set VISA to COM6 | bar: TEST A to = COMPLETE AL
9 | Hit Run Button | bar: TEST A to = COMPLETE AL
11 | Select Channels on PTC10 | bar: TEST A to = COMPLETE AL
12 | On PTC10 Change Heater Out Ramp to 0°/s | bar: TEST A to = COMPLETE AL
13 | On PTC10 Load Thermal Program. "Program" --> "LOAD" --> (Select Program) | bar: TEST A to = COMPLETE AL
10 | Copy Log Filename to Run Log | bar: TEST A to = COMPLETE AL
14 | Overlay PTC Logger above MFLI | bar: TEST A to = COMPLETE AL
15 | Clear MFLI FLAGS | bar: TEST A to = COMPLETE AF
16 | Hit Start Acquisition in PTC logger | bar: TEST A to = COMPLETE AL
17 | In MFLI Stop Plotter (New for New Software) | bar: TEST A to = COMPLETE AL
17 | Hit Record on MFLI Control panel | bar: TEST A to = COMPLETE AF
18 | Record Start Time in Run Log | bar: TEST A to = COMPLETE AF
19 | Save Test Log | bar: TEST A to = COMPLETE AF
20 | On PTC10 Hit play button | bar: TEST A to = COMPLETE AL
21 | On PTC10 Enable Heater Outputs | bar: TEST A to = COMPLETE AL
22 | Simultaneously touch air thermistor and rotate table to register data streams | bar: TEST A to = COMPLETE AL
23 | In MFLI Stop Plotter | bar: TEST A to = COMPLETE AL
24 | Confirm Files are being streamed & copy link | bar: TEST A to = COMPLETE AF
24 | WAIT FOR TEST TO COMPLETE | bar: TEST A to = COMPLETE AF
25 | Simultaneously touch air thermistor and rotate table to register data streams | bar: TEST A to = COMPLETE AL
26 | In PTC Temp Logger - Stop Acquisition | bar: TEST A to = COMPLETE AL
27 | In PTC10 Disable Heater Outputs | bar: TEST A to = COMPLETE AL
28 | In PTC10 Hit play button to STOP program | bar: TEST A to = COMPLETE AL
29 | Change Heater Out Ramp to 0.008°/s | bar: TEST A to = COMPLETE AL
30 | In MFLI Controller hit record button to stop acq | bar: TEST A to = COMPLETE AF
31 | Record End Time in Run Log | bar: TEST A to = COMPLETE AF
32 | Record MFLI Flags in Run Log | bar: TEST A to = COMPLETE AF
33 | Download stream folder | bar: TEST A to = COMPLETE AL
34 | Record MFLI Session and Stream folders on local machine (LIA) | bar: TEST A to = COMPLETE AL
35 | Record downloaded folder name in Run Log & add hyperlink | bar: TEST A to = COMPLETE AE
36 | Copy Temperature Data to Network | bar: TEST A to = COMPLETE AL
37 | Copy Gyro Data to Network (or confirm already there) | bar: TEST A to = COMPLETE AE
38 | Restart any instruments and water supply | bar: TEST A to = COMPLETE AL
39 | Save Test Log | bar: TEST A to = COMPLETE AE
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
Note on Q3: Taranta A.A.: All instruments ON except general photonics
Note on R3: Taranta A.A.: All instruments on except OSAs and Walics rack. PSGA-101 ON with fan facing gyro
Note on S3: Taranta A.A.: All instruments on except OSAs and Walics rack. PSGA-101 ON with fan facing gyro
Note on T3: Taranta A.A.: All instruments on except OSAs and Walics rack. PSGA-101 ON with fan facing gyro
Note on U3: Taranta A.A.: All instruments on except OSAs and Walics rack. PSGA-101 ON with fan facing gyro
Note on Z3: Taranta A.A.: CONTINUATION OF PRIOR TEST
Note on T4: Taranta A.A.: New LabOne Version Installed
Note on N7: Taranta A.A.: Light Source Off
Note on O7: Taranta A.A.: Light Source Off
Note on P7: Taranta A.A.: Light Source Off
Note on T13: Taranta A.A.: Channels did not seem to be interrupted on PTC10, despite simultaneously running Pol Test V1.5 and acquiring GP Temperautre sweep data
Note on K14: Taranta A.A.: Set to 30°C
Note on T23: Taranta A.A.: Data Acquisition failed on temperature logger because of attempting to run PTC10 logger while simultaneously running pol test V1.5 and
Note on W27: Taranta A.A.: Started Power change labview program to write new power values to the Manlight
Note on C29: Taranta A.A.: Bad Test - Temp Logger Failed to Start
Note on J33: Taranta A.A.: Left at 30°C
Note on Y33: Taranta A.A.: Paused Program
Note on U36: Taranta A.A.: Delayed recording of end time by a few min.

## Action Log
Date / Time | Action
2019-06-06 15:30 | Removed coil from blue thermal enclosure, opened optics enclosure, and inserted freestanding coil into optics enclosure. No Fibre outside optics enclosure now.
2019-06-06 16:19 | While Screwing down enclosure monitored ASE throughput power of gyro TOTAL THROUGHPUT AT 17.5dBm C-Band (MODULATION OFF) = -14.077dBm
2019-06-12 10:49 | Installed updated version of MFLI software onto machine and computer: LabOne 18.12 MFLI Update (LabOneMF-18.12.59443.tar, 81.4 MB) --> INSTALLED ON MFLI LabOne 18.12 for 64bit Windows (LabOne64-18.12.59443.msi, 264.3 MB) --> INSTALLED ON THIS COMPUTER
2019-06-24 10:37 | Cleared all stream folders from MFLI after observing empty datasets during long run at 3Hz
2020-02-04 16:31 | Restarted gyro with coil inside black enclosure - Started Testing at max power
2020-02-05 08:19 | Tested GPIB control of ASE source to do optical power vs ARW sweep of instrument. Created new labview software to control. Started ARW vs power run.
2020-02-05 14:43 | Created new VI to run ASE source independently. During this time, ASE power was frequently modified.
2020-02-11 09:31 | Set Power back to 17.5 dBm after testing at 16dBm over weekend
2020-02-11 10:04 | Moved Coil to inside thermal enclosure. Set Yellow Pt100 to be in enclosure air next to coil and white Pt100 taped to copper plate. Closed optics enclosure and routed fibers to thermal enclosure.
2020-08-14 11:58 | No Light Testing on LIA / PD in SWIMA configuration
2020-08-14 13:09 | Installed updated version of MFLI software onto S^2 machine and computer (UOS-215974): LabOne 20.01 MFLI Update (LabOneMF-20.01.1818.tar, 86.5 MB) --> INSTALLED ON MFLI LabOne 20.01 for 64bit Windows (LabOne64-20.01.1818.msi, 317.4 MB) --> INSTALLED ON THIS COMPUTER
0