Type of manuscript: Original Research Article

Safety and Performance of the Ultrathin Sirolimus-eluting Coronary Stent in an All-comers Patient Population: The S-FLEX UK-II Registry

Short Title: S-FLEX UK-II registry – One-year clinical outcomes

V J Karthikeyan1 (VJ.Karthikeyan@wwl.nhs.uk), Abdul Mozid2 (abdul.mozid@nhs.net), Suneil Aggarwal3 (suneil.aggarwal@lhch.nhs.uk), Abhishek Kumar1 (Abhishek.Kumar@wwl.nhs.uk), David Hildick Smith4 (david.Hildick-smith@bsuh.nhs.uk), Richard Anderson5 (richard.anderson@wales.nhs.uk), Satheesh Nair6 (satheesh.nair@wales.nhs.uk), Neil Ruparelia7 (Neil.Ruparelia@royalberkshire.nhs.uk) , Nick Curzen8 (Nick.Curzen@uhs.nhs.uk), Murugu Veerasamy2 (mveerasamy@nhs.net), Sandra Elsheikh3 (sandra.elsheikh@lhch.nhs.uk), Azfar Zaman9 Azfar.Zaman@nhs.net

1Department of Cardiology, Wigan Hospital, UK

2Department of Cardiology, Leeds General Infirmary, UK.

3Department of Cardiology, Liverpool Heart and Chest Hospital, Liverpool, UK.

4Department of Cardiology, Brighton and Sussex University Hospital, Brighton, UK.

5Department of Cardiology, University Hospital of Wales, Cardiff, UK.

6Department of Cardiology, Glan Clwyd Hospital, Rhyl, UK.

7Department of Cardiology, Royal Berkshire Hospital, Reading, UK.

8Department of Cardiology, Southampton General Hospital, Southampton, UK.

9Department of Cardiology, Freeman Hospital and Newcastle University, Newcastle upon Tyne, UK.

Name of Corresponding author:

Prof. Azfar Zaman

Freeman Hospital

Freeman Road, High Heaton

Newcastle Upon Tyne

NE7 7DN, UK

Contact number: +44 191 9330209

Email ID: azfar.zaman@nhs.net

ABSTRACT

Objective: We evaluated the clinical safety and performance of the ultrathin strut biodegradable polymer-coated Supraflex Cruz (Sahajanand Medical Technologies Ltd., Surat, India) sirolimus-eluting stents (SES) in an all-comers patient population requiring coronary stent implantation.

Study Design: The study was a prospective, observational, multicenter, single-arm registry.

Study Settings: The study was conducted at 19 NHS Hospitals across the UK, from March 2020 to September 2021.

Study Participants: A total of 1904 patients with systematic coronary artery disease (age ≥18 years) who underwent percutaneous coronary intervention with at least one Supraflex Cruz SES were enrolled.

Primary and Secondary Outcomes Measure: The primary endpoint was target lesion failure (TLF), a composite of cardiac death, target vessel myocardial infarction (TV-MI) and clinically indicated target lesion revascularization (CI-TLR), at 12 months. Safety endpoints were stent thrombosis, all cause death, and any MI. Pre-specified subgroups analysis included patients with diabetes mellitus, bifurcation lesion, type B2/C lesion defined as per ACC/AHA lesion classification, and long coronary lesions (>20 mm).

Results: A total of 2973 Supraflex Cruz SES were implanted in 1835 patients (mean age: 65.20±11.03 years). Of these, 404 patients had diabetes mellitus (491 lesions), 271 had bifurcation lesions (293 lesions), 1541 had type B2/C lesions (1832 lesions) and 985 had long coronary lesions (>20 mm, 1139 lesions). Among the overall population, device success was achieved in 98.2% of lesions. TLF occurred in 12 (0.7%) patients (0.3% cardiac death, 0.2% TV-MI, 0.2% CI-TLR) at 30 days and in 43 (2.3%) patients (0.8% cardiac death, 0.8% TV-MI, 0.9% CI-TLR) at 12 months follow-up. The rate of definite stent thrombosis was 0.3% in the overall population at 12 months. The incidence of TLF and stent thrombosis was 6.2% and 1% respectively in the diabetic, 1.8% and none in bifurcation lesion, 2.5% and 0.3% in type B2/C lesion, and 2.7% and 0.3% in long coronary lesions (>20 mm) subgroups at 12 months follow-up, respectively.

Conclusion: The S-FLEX UK-II registry confirms the clinical safety and performance of the ultrathin Supraflex Cruz SES in an all-comers population with complex coronary artery disease, demonstrating low rates of TLF and stent thrombosis.

Keywords: coronary artery disease; diabetes mellitus; drug-eluting stents, coronary stents

STRENGTHS AND LIMITATIONS OF THIS STUDY

- The strength of this study was its real-world setting that included an all-comers population.

- The inclusion of a high proportion of high-risk patients, such as those with diabetes, bifurcation lesions, complex B2/C lesions, and long coronary lesions (>20 mm), enhances relevance to current practice of complex coronary intervention.

- Revascularization strategy and post-procedure management were operator guided - further enhancing the real-world design and purpose of this study.

- The study was a non-randomized single-arm registry and lacked a direct comparator.

- The follow-up was conducted for up to 12 months; data beyond this time point was not collected.

INTRODUCTION

Successive improvements in the design characteristics of drug-eluting stents (DES) have led to reduced rates of restenosis and stent thrombosis in patients with symptomatic coronary artery disease (CAD). The latest generation DES have demonstrated safety and effectiveness reflected in good clinical outcomes. Nevertheless, there remains a potential link between occurrence of late or very late stent thrombosis and a combination of thick stent struts and polymer-mediated inflammatory reaction together with delayed healing and incomplete re-endothelialization. Consequently, iterations of DES, featuring biodegradable polymer coatings and ultrathin struts, have emerged to mitigate this potential drawback of second-generation DES with the goal of further improving clinical outcomes.1 2 The Supraflex Cruz (Sahajanand Medical Technologies Ltd., Surat, India), represents the next generation of Supraflex stent and is a biodegradable polymer-coated sirolimus-eluting stent (SES) engineered on an ultrathin cobalt-chromium platform with open-cell design and dual valley-to-valley connection between strut rings with alternate long dual Z-link (LDZ) to enhance deliverability in complex and challenging coronary lesions.

Optical coherence tomography (OCT) imaging studies have revealed, rapid/uniform vascular healing with equivalent neointimal thickness with Supraflex Cruz SES compared to the Xience stent.3 4 However, there is limited real-world evidence on Supraflex Cruz SES, with available data primarily focused only on the South Asian population.5 6 Therefore, the S-FLEX UK-II registry was designed to evaluate the clinical safety and performance of the Supraflex Cruz SES in a real-world, all-comers population with CAD in UK. The registry prespecified high-risk subgroups for analysis — patients with diabetes mellitus, patients with bifurcation lesions, patients with type B2/C lesions and patients with long coronary lesions (>20 mm).

METHODS

Study design and population

The S-FLEX UK-II registry was a prospective, observational, single-arm, multi-center registry designed to evaluate the clinical safety and performance of Supraflex Cruz SES in an all-comers patient population. A total of 1904 patients (age ≥18 years) who underwent percutaneous coronary intervention (PCI) with at least one Supraflex Cruz SES at 19 NHS Hospitals across UK, from March 2020 to September 2021 were enrolled. The study protocol was approved by the Southwest - Cornwall & Plymouth Research Ethics Committee (NHS Health Research Authority, REC reference number: 18/SW/0130) and the study was conducted in accordance with good clinical practice and Declaration of Helsinki.7 Written informed consent was obtained either from the patient or from the patient’s legal representative.

The inclusion criteria were age ≥18 years, symptomatic CAD including patients with acute coronary syndrome (ACS) or chronic coronary syndrome (CCS), and clinical indication for PCI, signed informed consent for participation in registry, and willing to participate in all follow-up assessments as per protocol. The exclusion criteria were patients with cardiogenic shock, known hypersensitivity or contraindication to aspirin, clopidogrel, heparin or any other anticoagulation / antiplatelet therapy required for PCI, cobalt chromium, sirolimus or contrast media, planned surgery within 12 months of PCI, pregnancy, life expectancy of less than one year and participating in a clinical study of another drug or medical device.

Patient and public involvement

Patients and/or the public were not involved in the design, conduct, reporting, or dissemination plans of this study

Study device

The Supraflex Cruz stent is built upon ultrathin L-605 cobalt chromium alloy (60 µm). It has an open-cell design and dual valley-to-valley connection between strut rings with alternate LDZ orientation. The stent has a conformal coating of sirolimus drug (1.4 μg/mm2) in a blend with biodegradable polymers (4-6 μm). The multilayer drug polymer coating comprises of a combination of hydrophobic [PLLA: poly-L-lactic acid, and PLCL: poly (L-Lactide-co-ε-Caprolac-tone)] polymers blended with sirolimus in the middle and innermost layer, and the outer drug-free hydrophilic polymer (PVP: polyvinylpyrrolidone) coating. Nearly 80% of drug is released within one month with the remaining drug programmed to release at a slow rate over three months. The polymers are gradually degraded and excreted in the form of biologically inert molecules within 10 to 12 months. Supraflex Cruz SES is available in a wide range of sizes — diameter 2.00 mm to 4.50 mm and length 8 mm to 48 mm.

Interventional procedure and follow-up

Lesion severity was assessed by visual estimation after invasive coronary angiography and revascularisation and postoperative management done in accordance with physician practice. Dual antiplatelet therapy was recommended for at least 12 months in patients with acute coronary syndrome (ACS) and six months in patients with stable CAD. All patients were followed-up via clinic visit or telephonic communication at 30 days and 12 months after PCI. Angiographic data were collected for patients readmitted with clinical symptoms of ischemia. For those who underwent angiographic re-evaluation within 12 months of stent implantation, image analysis was conducted to assess whether disease progression had occurred in the previously stented segment, including peri-stent areas 2 mm adjacent to the stent, or if it was limited to arterial segments remote from the stented region.

In our study, all events were investigator reported without any event adjudication committee. However, to ensure continuous protocol compliance and accurate data reporting, periodic site monitoring was carried out by Psephos Biomedica (Sussex Innovation Centre, University of Sussex, Brighton BN1 9SB, UK).

Study endpoints and definition

The primary endpoint was the rate of target lesion failure (TLF) defined as a composite of cardiac death, target vessel myocardial infarction (TV-MI) and clinically indicated target lesion revascularization (CI-TLR) by percutaneous or surgical methods at 12 months. The secondary safety endpoints included overall stent thrombosis (definite/probable), all-cause death (cardiac, vascular, and non-cardiovascular), and myocardial infarction at 12 months. The secondary efficacy endpoints included repeat revascularization (target lesion, target vessel or non-target vessel revascularization) at 12 months. We also recorded target vessel failure (TVF) which was a composite of cardiac death, TV-MI and clinically indicated target vessel revascularization (TVR) at 12 months as a secondary endpoint.

Device success was defined as successful delivery and deployment of the study device (Supraflex Cruz SES) at the designated target lesions, along with successful retrieval of the stent delivery system and achievement of final residual stenosis of <30% on visual estimation. Procedure success was defined as achievement of device success without any adverse cardiac event occurring during hospital-stay. Cardiac death was defined as death due to any cardiac mechanisms (fatal arrhythmia, sudden death, MI, or low cardiac output heart failure), any unwitnessed death, death of unknown cause or procedural related death. Target vessel MI was defined as target vessel Q-wave or non-Q-wave MI with the evidence of myocardial necrosis in the vascular territory of previously treated target vessel. MI was adjudged according to the fourth universal definition.8 The TLR was defined as any repeat revascularization by percutaneous or surgical means of any part of the target vessel including the target lesion (including 5 mm proximal and distal to the stent). Clinically indicated TLR was defined as any repeat revascularization for in-lesion diameter stenosis of >50% by quantitative coronary angiography (QCA) and presence of ischemic signs and symptoms attributable to the target lesion, or any repeat revascularization done for in-lesion diameter stenosis of >70% and absence of ischemic signs and symptoms, or presence of severe ischemic signs and symptoms attributable to the target lesions in the absence of QCA data and if in-lesion diameter stenosis is ≤50%.9 TVR was defined as any repeat revascularization by percutaneous or surgical means of any segment of the target vessel including the target lesion.9 Stent thrombosis was defined according to Academic Research Consortium-2 (ARC-2) definition.9

Statistical analysis

Continuous variables are expressed as mean ± standard deviation, while categorical variables are presented as frequency (percentage). Cumulative rate of events up to 12 months were estimated by using the Kaplan Meier method. Additionally, a subgroup analysis was undertaken, wherein clinical outcomes were evaluated for patients with diabetes mellitus (patient was a known diabetic and on pharmacological treatment or document HbA1c  >7% even if not on pharmacological treatment), bifurcation coronary lesions, type B2/C lesions (defined according to ACC/AHA lesion classification10), and long coronary lesions (target lesion length  of >20 mm). The data were analyzed using R Statistical Software version 4.3.1.

RESULTS

Baseline characteristics

Of 1904 enrolled patients, a total of 1835 patients were included for the analysis (Supplementary Figure 1). The mean age was 65.20±11.03 years, >75% were male, and 22% (n=404) had diabetes mellitus of which 20.5% (n=81) were dependent on insulin. A total of 410 (22.3%) patients had history of revascularization either via PCI or coronary artery bypass grafting (CABG). A majority of patients (68.7%, n=1257/1830) presented with ACS and 43.4% (n=796) patients were diagnosed with multivessel CAD. Moderate left ventricular dysfunction (ejection fraction <60% and >35%) was reported in 36.2% (n=660) patients and severe left ventricular dysfunction (ejection fraction ≤35%) in 4.7% (n=85) patients. The baseline clinical characteristics of the entire cohort and four prespecified subgroups are outlined in Table 1.

Lesion and procedural characteristics

In the overall cohort of 1835 patients, 2230 lesions were treated with implantation of 2973 Supraflex Cruz SES. The detailed lesion and procedural characteristics are summarized in Table 2 and Table 3, respectively. Of all treated lesions, 88.6% lesions were type B2/C, 13.1% were bifurcation lesions, 55.5% (1139/2054) were long diffuse lesions (>20 mm), and 13.2% were ostial lesions (within 3 mm of vessel origin).  A total of 30% (612 lesions) lesions were moderately or heavily calcified, thrombus was visible in 22.7% (471 lesions), and 9% (185 lesions) were chronic total occlusions (≥3-month-old occlusion). Overall, the mean reference vessel diameter was 3.29±1.01 mm and represented mean diameter stenosis of more than 90%. During the procedure, predilatation was performed in 2043 (91.6%) lesions and post dilatation in 1879 (84.3%) treated lesions.

Device success was achieved in 98.2% (2189/2230) lesions with procedural success in 97.7% (1792/1835) patients. The details of antiplatelet medication prescribed at discharge and its adherence at 30 days and 12 months follow-up are outlined in Supplementary Table 1.

Study Outcomes

Clinical outcomes at 30 days and 12 months follow-up are presented in Table 4. At 12 months follow-up, the cumulative incidence of TLF, the primary endpoint, was 2.3% (n=43) comprising 0.8% (n=14) cardiac death, 0.8% (n=14) TV-MI, and 0.8% (n=15) CI-TLR. At 12 months, 34 (1.9%) patients underwent TVR, and 43 (2.3%) patients underwent non-TVR. The overall rate of definite/probable stent thrombosis was 0.3% at 12 months. In prespecified subgroups, the incidence of TLF were 6.2% and 1% in diabetes mellitus subgroup, 1.8% and none in bifurcation subgroup, 2.5% and 0.3% in type B2/C lesion subgroup, and 2.7% and 0.3% in long lesions (>20 mm) at 12 months follow-up, respectively. The time-to-event Kaplan Meier curves up to 12 months follow-up are presented in Figure 1 and Supplementary Figure 2 A-D.

DISCUSSION

The findings of this multicentre, single arm, prospective registry demonstrate safety and efficacy of the latest generation Supraflex Cruz ultrathin stent in an all-comers UK population. Our findings revealed a low incidence of the primary endpoint of TLF at 2.3% and the safety endpoint of definite or probable stent thrombosis of 0.3% at 12 months follow-up.

Technical iterations of the latest generation DES such as ultrathin struts with improved design and biodegradable polymer coating have improved clinical outcomes and reduced stent related complications.  The suggested mechanism for the observed improvements are the reduction in risk of acute/chronic inflammation and vessel wall injury, accelerated re-endothelialization and strut coverage, and reduced neointimal proliferation and thrombogenicity.11 12

The findings in this current registry are in line with the S-FLEX UK registry (466 patients) which reported a low TLF rate at 2.4% with Supraflex SES, a predecessor of Supraflex Cruz SES.13 Two real-world registries from India reported higher TLF rates of 3.8% and 5.75% with the Supraflex family of stents at 12 months follow-up.5 6 However, it is important to consider that these varying outcomes may be influenced by variations among the patient populations being studied with a higher proportion of patients with diabetes mellitus seen in the Indian population. Additionally, the relatively low TLF rate observed in the present registry might be attributed to a higher frequency of lesion predilatation (91.2%) before stenting and higher rate of post-dilatation (84.3%).

Incomplete strut coverage and inadequate, delayed endothelization are strongly associated with the incidence of late stent thrombosis. The strut thickness is the primary determining factor for the duration of strut coverage post stent implantation, with thinner struts potentially facilitating endothelial coverage and reducing risk of platelet aggregation.14 15 The SiBi and TAXCO studies assessed strut apposition and tissue endothelialization of the ultrathin Supraflex Cruz SES using OCT.3 4 The SiBi study reported an average of 91.26% strut coverage at 35.3±5 days 4 and the TAXCO study revealed 97.6% stent coverage at six months with Supraflex Cruz post-PCI, a rate comparable to that of the Xience everolimus-eluting stent.3 These findings collectively suggest low inflammation and early optimal healing response with Supraflex Cruz SES and give mechanistic insight to the safety endpoint seen both in our registry and in those reported by others.5 6 16 17

The device success rate was high and may be attributable to design changes in the newer generation Supraflex Cruz SES. It is an ultrathin cobalt-chromium platform with unique alternate LDZ links from valley-to-valley which has been developed with an aim to treat complex coronary lesion in high risks patients. The delivery mechanism has a softer balloon for stent retention and retrieval after deployment, and a redesigned proximal shaft to enhance crossability compared to its predecessor.

The results of the current S-FLEX UK-II registry confirm class leading safety and efficacy with low rates of TLF and stent thrombosis in patients with high-risk clinical and anatomical characteristics. Further evidence of safety and efficacy of this platform was reported in the Cruz-HBR registry demonstrating superior outcomes in PCI compared to the Biolimus-drug coated stent for high-risk bleeding (HBR) patients.18 In the recently published FIRE trial (2023), it was observed that in elderly patients (≥75 years) with MI and multivessel CAD, adopting a physiology-guided complete revascularization using Supraflex Cruz SES was associated with a significantly reduced risk when compared to culprit-only revascularization.19 The findings from the current S-FLEX UK-II registry add further evidence for the safety and effectiveness of the ultrathin Supraflex Cruz SES in an all-comers population in UK.

Study limitations: The study was non-randomized and lacked a direct comparator. Also, the coronary lesions were evaluated through visual estimation and no angiographic core lab assessment and intra-coronary imaging modalities were used.  The follow-up was conducted for up to 12 months with data beyond this time point not collected.	All events were recorded directly from patients and their medical notes with no adjudication.

Conclusion

The S-FLEX UK-II registry provides further evidence on the clinical safety and performance of Supraflex Cruz SES in all-comers patients with CAD in UK. The results reinforce both safety and performance of the Supraflex Cruz SES, with clinically low rates of TLF and stent thrombosis at 12 months, even among complex subgroups, with diabetes mellitus, bifurcation lesions, type B2/C lesions, and long coronary lesions (>20 mm).

Contributors: The authors were involved in recruiting patients, authenticity of the data, dealing with queries and adjudicating any serious adverse events. All authors (as mentioned above) were involved in drafting the manuscript, responsible for the final approval of the manuscript.

Funding: This study was funded by Sahajanand Medical Technologies Ltd., Surat, India. The funding source had no role in study design, trial conduct, or the decision to publish.

Competing interests: None declared.

Patient and public involvement: Patients and/or the public were not involved in the design, conduct, reporting, or dissemination plans of this study.

Patient consent for publication: Not required.

Ethics approval: The study was approved by the Southwest - Cornwall & Plymouth Research Ethics Committee (NHS Health Research Authority, REC reference number: 18/SW/0130)

Data availability statement: Data are available upon reasonable request.

REFERENCES

1. Nakamura M, Kadota K, Nakagawa Y, et al. Ultrathin, biodegradable-polymer sirolimus-eluting stent vs thin, durable-polymer everolimus-eluting stent. Cardiovasc Interv 2022;15(13):1324-34.

2. Monteleone PP. Ultrathin Stents in the" Real World". Cardiovasc Revasc Med 2021;30:18-19.

3. Abhyankar A, Abizaid A, Chamié D, et al. Comparison of neointimal coverage between ultrathin biodegradable polymer‐coated sirolimus‐eluting stents and durable polymer‐coated everolimus‐eluting stents: 6 months optical coherence tomography follow‐up from the TAXCO study. Catheter Cardiovasc Interv 2021;97(3):423-30.

4. Abhyankar A, Abizaid A, Chamié D, et al. Healing and early stent coverage after ultrathin strut biodegradable polymer‐coated sirolimus‐eluting stent implantation: SiBi optical coherence tomography study. Catheter Cardiovasc Interv 2021;98(7):1335-42.

5. Pothineni RB, Vijan V, Potdar A, et al. Clinical outcomes of ultrathin biodegradable polymer-coated sirolimus-eluting stents in an all-comer population: One-year results from the T-FLEX registry including high-risk subgroups. Anatol J Cardiol 2021;25(10):706.

6. Ajmera P, Pothineni R, Chawla KK, et al. Twelve months clinical outcomes of ultrathin strut sirolimus-eluting stent in real-world Indian patients with coronary artery disease. Am J Cardiovasc Dis 2022;12(5):262.

7. Emanuel EJ. Reconsidering the declaration of Helsinki. The Lancet 2013;381(9877):1532-33.

8. Thygesen K, Alpert JS, Jaffe AS, et al. Fourth universal definition of myocardial infarction (2018). Circulation 2018;138(20):e618-e51.

9. Garcia-Garcia HM, McFadden EP, Farb A, et al. Standardized end point definitions for coronary intervention trials: the academic research consortium-2 consensus document. Circulation 2018;137(24):2635-50.

10. Alraies MC, Darmoch F, Tummala R, et al. Diagnosis and management challenges of in-stent restenosis in coronary arteries. World J Cardiol 2017;9(8):640.

11. Madhavan MV, Kirtane AJ, Redfors B, et al. Stent-related adverse events> 1 year after percutaneous coronary intervention. Journal of the American College of Cardiology 2020;75(6):590-604.

12. Gherasie F-A, Valentin C, Busnatu S-S. Is There an Advantage of Ultrathin-Strut Drug-Eluting Stents over Second-and Third-Generation Drug-Eluting Stents? J Pers Med 2023;13(5):753.

13. Choudhury A, Garg S, Smith J, et al. Prospective evaluation of an ultrathin strut biodegradable polymer-coated sirolimus-eluting stent: 12 months’ results from the S-FLEX UK registry. BMJ open 2019;9(10):e026578.

14. Leone A, Simonetti F, Avvedimento M, et al. Ultrathin Struts Drug-Eluting Stents: A State-of-the-Art Review. J Pers Med 2022;12(9):1378.

15. Gao C, Kogame N, Modolo R, et al. The ultra-thin strut sirolimus-eluting coronary stent: SUPRAFLEX. Future Cardiol 2020;17(2):227-37.

16. Suwannasom P, Athiksakul S, Thonghong T, et al. Clinical outcomes of an ultrathin-strut sirolimus-eluting stent in all-comers population: Thailand Orsiro registry. BMC Cardiovasc Disord 2021;21(1):1-8.

17. Bartorelli AL, Versaci F, Briguori C, et al. The BIOFLOW-III Italian Satellite Registry: 18-month results of the Orsiro stent in an all-comer high-risk population. J Cardiovasc Med 2019;20(7):464-70.

18. CRUZ-HBR: a prospective, multi-center, allcomer registry on 1203 patients including 466 with high-bleeding risk. EuroPCR; 2022; Paris, France.

19. Biscaglia S, Guiducci V, Escaned J, et al. Complete or culprit-only PCI in older patients with myocardial infarction. New England Journal of Medicine 2023;389(10):889-98.

FIGURE LEGENDS

Figure 1: Kaplan Meier curve showing target lesion failure (TLF) up to 12 months follow-up in overall population

SUPPLEMENTARY FILES

Supplementary Table 1: Details of antiplatelet medication at discharge, at 30 days and 12 months follow-up

Supplementary Figure 1: Study flow chart

Supplementary Figure 2: Kaplan Meier curves for target lesion failure (TLF) up to 12 months follow-up in (A) Diabetes mellitus subgroup, (B) Bifurcation lesion subgroup, (C) Type B2/C lesion subgroup, and (D) Long lesion (>20 mm) subgroup

TABLE LEGENDS

Table 1: Baseline clinical characteristics

Table 2: Lesion characteristics

Table 3: Procedural characteristics

Table 4: Clinical outcomes at 30 days and 12 months follow-up

TABLES

Table 1: Baseline clinical characteristics

| Characteristics |  | Total (n = 1835) | Diabetes (n = 404) | Bifurcation (n = 271) | Lesion B2/C (n = 1541) | Long lesion (>20 mm) (n = 985) |
| --- | --- | --- | --- | --- | --- | --- |
| Age, years |  | 65.20±11.03 | 65.96±10.49 | 65.59±10.66 | 65.21±11.13 | 65.15±10.86 |
| Gender | Male | 1406 (76.6%) | 312 (77.2%) | 216 (79.7%) | 1189 (77.2%) | 773 (78.5%) |
|  | Female | 429 (23.4%) | 92 (22.8%) | 55 (20.3%) | 352 (22.8%) | 212 (21.5%) |
| Height cm |  | 171.27±10.37 (1782/1835) | 171.10±9.20 (395/404) | 171.63±9.66 (267/271) | 171.39±10.31 (1495/1541) | 171.60±10.06 (952/985) |
| Weight, kg |  | 86.00±17.50 (1803/1835) | 90.88±18.03 (398/404) | 85.71±17.68 (269/271) | 86.09±17.66 (1514/1541) | 86.54±18.02 (966/985) |
| Body mass index, kg/m2 |  | 29.41±6.15 (1779/1835) | 31.07±5.91 (395/404) | 29.14±5.54 (267/271) | 29.40±5.94 (1492/1541) | 29.48±6.26 (950/985) |
| Coexisting illness and addiction |  |  |  |  |  |  |
| Diabetes mellitus |  | 404 (22.0%) | 404 (100.0%) | 62 (22.9%) | 351 (22.8%) | 228 (23.1%) |
| Insulin requiring |  | 81/394 (20.5%) | 81/394 (20.5%) | 10/61 (16.3%) | 72/341 (21.1%) | 45/222 (20.27%) |
| Non-Insulin requiring |  | 313/394 (79.4%) | 313/394 (79.4%) | 51/61 (83.6%) | 269/341 (78.8%) | 177/222 (79.7%) |
| Hypertension |  | 939 (51.2%) | 275 (68.1%) | 138 (50.9%) | 788 (51.1%) | 496 (50.4%) |
| Renal insufficiency |  | 54 (2.9%) | 23 (5.7%) | 12 (4.4%) | 46 (3.0%) | 29 (2.9%) |
| Family history of CAD |  | 705 (38.4%) | 136 (33.7%) | 111 (41.0%) | 575 (37.3%) | 361 (36.6%) |
| Peripheral vascular disease |  | 85 (4.6%) | 33 (8.2%) | 20 (7.4%) | 71 (4.6%) | 53 (5.4%) |
| Smoker |  | 388 (21.1%) | 76 (18.8%) | 44 (16.2%) | 331 (21.5%) | 211 (21.4%) |
| Hypercholesterolemia |  | 865 (47.1%) | 232 (57.4%) | 142 (52.4%) | 719 (46.7%) | 441 (44.8%) |
| Congestive heart failure |  | 45 (2.5%) | 16 (4.0%) | 9 (3.3%) | 33 (2.1%) | 21 (2.1%) |
| Transient ischemic attack |  | 47 (2.6%) | 14 (3.5%) | 7 (2.6%) | 38 (2.5%) | 24 (2.4%) |
| Previous stroke |  | 39 (2.1%) | 17 (4.2%) | 4 (1.5%) | 35 (2.3%) | 21 (2.1%) |
| Previous MI |  | 270 (14.7%) | 88 (21.8%) | 41 (15.1%) | 229 (14.9%) | 145 (14.7%) |
| Previous PCI |  | 312 (17.0%) | 94 (23.3%) | 49 (18.1%) | 266 (17.3%) | 165 (16.8%) |
| Previous CABG |  | 98 (5.3%) | 28 (6.9%) | 15 (5.5%) | 94 (6.1%) | 60 (6.1%) |
| Clinical presentation |  |  |  |  |  |  |
| NSTEMI |  | 730/1830 (40.0%) | 185/403 (45.9%) | 102 (37.6%) | 583/1538 (37.91%) | 377/983 (38.4%) |
| STEMI |  | 417/1830 (22.8%) | 63/403 (15.6%) | 44 (16.2%) | 376/1538 (24.45%) | 229/983 (23.3%) |
| Unstable angina |  | 110/1830 (6.0%) | 25/403 (6.2%) | 24 (8.9%) | 92/1538 (6.0%) | 53/983 (5.4%) |
| Silent ischemia |  | 10/1830 (0.5%) | 2/403 (0.5%) | 0 (0.0%) | 8/1538 (0.5%) | 6/983 (0.6%) |
| Stable angina |  | 563/1830 (30.8%) | 128/403 (31.8%) | 101 (37.3%) | 479/1538 (31.1%) | 318/983 (32.3%) |
| Extent of CAD |  |  |  |  |  |  |
| Single vessel disease |  | 1039 (56.6%) | 230 (56.9%) | 119 (43.9%) | 852 (55.3%) | 511 (51.9%) |
| Multiple vessel disease |  | 796 (43.4%) | 174 (43.1%) | 152 (56.1%) | 689 (44.7%) | 474 (48.1%) |
| Left ventricular ejection fraction |  |  |  |  |  |  |
| Good (≥ 60%) |  | 479/1821 (26.3%) | 106/401 (26.4%) | 85 (31.4%) | 406/1530 (26.5%) | 261/981 (26.6%) |
| Moderate (>35% to <60%) |  | 660/1821 (36.2%) | 152/401 (37.9%) | 97 (35.8%) | 542/1530 (35.4%) | 356/981 (36.3%) |
| Severe (≤35%) |  | 85/1821 (4.7%) | 29/401 (7.2%) | 16 (5.9%) | 79/1530 (5.2%) | 45/981 (4.6%) |
| Unknown |  | 597/1821 (32.8%) | 114/401 (28.4%) | 73 (26.9%) | 503/1530 (32.9%) | 319/981 (32.5%) |

Data are mean (SD), n (%), or n/n (%) in case of missing data.

CAD: coronary artery disease; MI: myocardial infarction; PCI: percutaneous coronary intervention; CABG: coronary artery bypass grafting; STEMI: ST-elevation myocardial infarction; NSTEMI: non ST-elevation myocardial infarction.

Table 2: Lesion characteristics

| Characteristics | Total (2230 Lesions) | Diabetes (491 Lesions) | Bifurcation (293 Lesions) | Lesion B2/C (1832 Lesions) | Long lesion (>20 mm) (1139 Lesions) |
| --- | --- | --- | --- | --- | --- |
| Target vessel |  |  |  |  |  |
| Left anterior descending artery | 935 (41.9%) | 188 (38.3%) | 174 (59.4%) | 756 (41.3%) | 456 (40.0%) |
| Right coronary artery | 769 (34.5%) | 185 (37.7%) | 39 (13.3%) | 651 (35.5%) | 437 (38.4%) |
| Left circumflex artery | 445 (20.0%) | 96 (19.6%) | 43 (14.7%) | 347 (18.9%) | 199 (17.5%) |
| Left main artery | 60 (2.7%) | 13 (2.6%) | 36 (12.3%) | 57 (3.1%) | 36 (3.2%) |
| Saphenous vein grafts | 21 (0.9%) | 9 (1.8%) | 1 (0.3%) | 21 (1.1%) | 11 (1.0%) |
| Lesion characteristics |  |  |  |  |  |
| AHA/ACC lesion classification |  |  |  |  |  |
| Type A | 41/2067 (2.0%) | 6/462 (1.3%) | 0 (0.0%) | _ | 0 (0.0%) |
| Type B1 | 194/2067 (9.4%) | 34/462 (7.4%) | 11/287 (3.8%) | _ | 0 (0.0%) |
| Type B2 | 560/2067 (27.1%) | 128/462 (27.7%) | 76/287 (26.5%) | 560 (30.6%) | 0 (0.0%) |
| Type C | 1272/2067 (61.5%) | 294/462 (63.6%) | 200/287 (69.7%) | 1272 (69.4%) | 1139 (100.0%) |
| Length |  |  |  |  |  |
| discrete (<10 mm) | 220/2054 (10.7%) | 39/460 (8.5%) | 24/285 (8.4%) | 117/1824 (6.4%) | - |
| tubular (10 - 20 mm) | 695/2054 (33.8%) | 156/460 (33.9%) | 77/285 (27.0%) | 568/1824 (31.1%) | - |
| diffuse (>20 mm) | 1139/2054 (55.5%) | 265/460 (57.6%) | 184/285 (64.6%) | 1139/1824 (62.4%) | 1139 (100.0%) |
| Eccentricity |  |  |  |  |  |
| concentric | 991/1950 (50.8%) | 220/441 (49.9%) | 141/282 (50.0%) | 784/1734 (45.2%) | 490/1095 (44.7%) |
| eccentric | 959/1950 (49.2%) | 221/441 (50.1%) | 141/282 (50.0%) | 950/1734 (54.8%) | 605/1095 (55.3%) |
| Accessibility |  |  |  |  |  |
| readily accessible tortuosity of proximal segment | 1534/2019 (76.0%) | 341/450 (75.8%) | 222/284 (78.2%) | 1316/1792 (73.4%) | 819/1117 (73.3%) |
| moderate tortuosity of proximal segment | 434/2019 (21.5%) | 101/450 (22.4%) | 59/284 (20.8%) | 425/1792 (23.7%) | 270/1117 (24.2%) |
| excessive | 51/2019 (2.5%) | 8/450 (1.8%) | 3/284 (1.1%) | 51/1792 (2.8%) | 28/1117 (2.5%) |
| Lesion angulation |  |  |  |  |  |
| none (<45º) | 1485/1978 (75.1%) | 326/444 (73.4%) | 186/283 (65.7%) | 1263/1753 (72.0%) | 778/1109 (70.2%) |
| moderate (≥45º & <90º) | 450/1978 (22.8%) | 109/444 (24.5%) | 92/283 (32.5%) | 447/1753 (25.5%) | 299/1109 (27.0%) |
| location in severe bend point, (≥ 90º) | 43/1978 (2.2%) | 9/444 (2.0%) | 5/283 (1.8%) | 43/1753 (2.5%) | 32/1109 (2.9%) |
| Lesion contour |  |  |  |  |  |
| smooth | 1004/2020 (49.7%) | 204/454 (44.9%) | 132/284 (46.5%) | 787/1793 (43.9%) | 460/1117 (41.2%) |
| irregular | 1016/2020 (50.3%) | 250/454 (55.1%) | 152/284 (53.5%) | 1006/1793 (56.1%) | 657/1117 (58.8%) |
| Bifurcation or side branch lesions |  |  |  |  |  |
| no major branch involvement | 1719/2058 (83.5%) | 378/460 (82.2%) | 43/287 (15.0%) | 1492/1821 (81.9%) | 910/1136 (80.1%) |
| bifurcation lesions requiring double guidewire | 317/2058 (15.4%) | 76/460 (16.5%) | 236/287 (82.2%) | 307/1821 (16.9%) | 211/1136 (18.6%) |
| inability to protect major side branches | 22/2058 (1.1%) | 6/460 (1.3%) | 8/287 (2.8%) | 22/1821 (1.2%) | 15/1136 (1.3%) |
| Degenerated bypass grafts with unstable lesions | 30/1914 (1.6%) | 6/417 (1.4%) | 5/254 (2.0%) | 30/1663 (1.8%) | 14/1035 (1.4%) |
| Lesion complexity |  |  |  |  |  |
| Lesion requiring overlapping stents | 540 (24.2%) | 110 (22.4%) | 99 (33.8%) | 493 (26.9%) | 424 (37.2%) |
| Re-stenotic lesion | 52 (2.3%) | 16 (3.3%) | 6 (2.0%) | 48 (2.6%) | 29 (2.5%) |
| Bifurcation | 293 (13.1%) | 64 (13.0%) | 293 (100.0%) | 276 (15.1%) | 184 (16.2%) |
| Moderate to heavy calcification | 612/2038 (30%) | 163/461 (35.3%) | 92/284 (32.4%) | 606/1807 (33.5%) | 431/1138 (37.9%) |
| Ostial (within 3 mm of vessel origin) | 269/2035 (13.2%) | 55/457 (12.0%) | 68/284 (23.9%) | 263/1805 (14.6%) | 164/1136 (14.4%) |
| Thrombus | 471/2074 (22.7%) | 90/464 (19.4%) | 45/287 (15.7%) | 464/1827 (25.4%) | 272 (23.9%) |
| Total Occlusion (<3 months) | 351/2062 (17.0%) | 65/463 (14.0%) | 33/285 (11.6%) | 345/1825 (18.9%) | 192/1138 (16.9%) |
| Chronic Total Occlusion (≥3 months) | 185/2062 (9.0%) | 44/463 (9.5%) | 29/285 (10.2%) | 185/1825 (10.1%) | 143/1138 (12.6%) |

Data are n (%), or n/n (%) in case of missing data

Table 3: Procedural characteristics

| Characteristics | Total (2230 Lesions) | Diabetes (491 Lesions) | Bifurcation (293 Lesions) | Lesion B2/C (1832 Lesions) | Long lesion (>20 mm) (1139 Lesions) |
| --- | --- | --- | --- | --- | --- |
| Reference vessel diameter, mm | 3.29±1.01 (2004/2230) | 3.27±0.55 (452/491) | 3.28±0.57 (282/293) | 3.29±0.93 (1778/1832) | 3.30±1.09 (1129/1139) |
| Diameter stenosis, % | 90.18±10.76 (2041/2230) | 89.78±11.57 (459/491) | 88.14±12.28 (285/293) | 90.56±10.53 (1812/1832) | 91.30±9.91 (1134/1139) |
| Pre dilatation | 2043 (91.6%) | 450 (91.6%) | 279 (95.2%) | 1706 (93.1%) | 1091 (95.8%) |
| Maximum inflation pressure pre stent, atm | 13.69±3.43 (1930/2230) | 13.70±3.77 (424/491) | 14.14±3.25 (267/293) | 13.97±3.08 (1605/1832) | 14.15±3.04 (1024/1139) |
| Post dilatation | 1879 (84.3%) | 416 (84.7%) | 273 (93.2%) | 1560 (85.2%) | 1012 (88.8%) |
| Maximum inflation pressure post stent, atm | 16.95±3.78 (1804/2230) | 16.49±3.80 (394/491) | 17.06±2.95 (262/293) | 17.11±3.38 (1491/1832) | 17.22±3.11 (973/1139) |
| Stents | n = 2973 Stents | n = 657 Stents | n = 432 Stents | n = 2554 Stents | n = 1754 Stents |
| No. of stents per patient, mm | 1.62±0.93 | 1.63±0.87 | 1.59±0.85 | 1.66±0.95 | 1.78±0.98 |
| No. of stents per lesion, mm | 1.43±0.73 | 1.44±0.72 | 1.55±0.81 | 1.48±0.76 | 1.62±0.82 |
| Total stent length per patient, mm | 44.14±30.16 | 44.81±28.74 | 44.21±27.30 | 46.14±30.99 | 54.19±31.46 |
| Total stent length per vessel, mm | 38.87±24.62 | 39.70±24.04 | 44.30±27.69 | 40.55±25.23 | 47.68±26.25 |
| Mean Stent Length, mm | 27.25±10.64 | 27.56±10.92 | 27.73±10.22 | 27.84±10.71 | 30.43±10.70 |
| Mean Stent Diameter, mm | 3.16±0.55 | 3.15±0.54 | 3.13±0.54 | 3.17±0.55 | 3.16±0.55 |
| Device Success | 2189/2230 (98.2%) | 470 (95.7%) | 275 (93.9%) | 1792 (97.8%) | 1109 (97.4%) |
| Procedural Success | 1792/1835 (97.7%) | 393 (97.3%) | 267 (98.5%) | 1503 (97.5%) | 964 (97.9%) |

Data are mean (SD), n (%), or n/n (%) in case of missing data.

Table 4: Clinical outcomes at 30 days and 12 months follow-up

|  | 30 days follow-up |  |  |  |  | 12 months follow-up |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Total (n = 1835) | Diabetes (n = 404) | Bifurcation (n = 271) | Lesion B2/C (n = 1541) | Long lesion (>20 mm) (n = 985) | Total (n = 1835) | Diabetes (n = 404) | Bifurcation (n = 271) | Lesion B2/C (n = 1541) | Long lesion (>20 mm) (n = 985) |
| TLF | 12 (0.7%) | 6 (1.5%) | 1 (0.4%) | 10 (0.6%) | 6 (0.6%) | 43 (2.3%) | 25 (6.2%) | 5 (1.8%) | 38 (2.5%) | 27 (2.7%) |
| Target vessel failure | 12 (0.7%) | 6 (1.5%) | 1 (0.4%) | 10 (0.6%) | 6 (0.6%) | 48 (2.6%) | 26 (6.4%) | 7 (2.6%) | 43 (2.8%) | 30 (3.0%) |
| Death from any cause | 6 (0.3%) | 2 (0.5%) | 1 (0.4%) | 6 (0.4%) | 4 (0.4%) | 34 (1.9%) | 16 (4.0%) | 4 (1.5%) | 29 (1.9%) | 21 (2.1%) |
| Cardiac death | 5 (0.3%) | 2 (0.5%) | 1 (0.4%) | 5 (0.3%) | 4 (0.4%) | 14 (0.8%) | 9 (2.2%) | 1 (0.4%) | 13 (0.8%) | 10 (1.0%) |
| All MI | 10 (0.5%) | 4 (1.0%) | 0 (0.0%) | 8 (0.5%) | 4 (0.4%) | 31 (1.7%) | 13 (3.2%) | 3 (1.1%) | 27 (1.8%) | 18 (1.8%) |
| TV-MI | 4 (0.2%) | 3 (0.7%) | 0 (0.0%) | 3 (0.2%) | 1 (0.1%) | 14 (0.8%) | 9 (2.2%) | 0 (0.0%) | 12 (0.8%) | 8 (0.8%) |
| Non-TV MI | 6 (0.3%) | 1 (0.2%) | 0 (0.0%) | 5 (0.3%) | 3 (0.3%) | 17 (0.9%) | 4 (1.0%) | 3 (1.1%) | 15 (1.0%) | 10 (1.0%) |
| TLR | 4 (0.2%) | 2 (0.5%) | 0 (0.0%) | 3 (0.2%) | 1 (0.1%) | 25 (1.4%) | 11 (2.7%) | 6 (2.2%) | 23 (1.5%) | 15 (1.5%) |
| CI-TLR | 3 (0.2%) | 1 (0.2%) | 0 (0.0%) | 2 (0.1%) | 1 (0.1%) | 15 (0.8%) | 7 (1.7%) | 4 (1.5%) | 13 (0.8%) | 9 (0.9%) |
| Non-CI TLR | 1 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 0 (0.0%) | 10 (0.5%) | 4 (1.0%) | 2 (0.7%) | 10 (0.6%) | 6 (0.6%) |
| TVR | 4 (0.2%) | 2 (0.5%) | 0 (0.0%) | 3 (0.2%) | 1 (0.1%) | 34 (1.9%) | 13 (3.2%) | 8 (3.0%) | 32 (2.1%) | 20 (2.0%) |
| CI-TVR | 3 (0.2%) | 1 (0.2%) | 0 (0.0%) | 2 (0.1%) | 1 (0.1%) | 20 (1.1%) | 8 (2.0%) | 6 (2.2%) | 18 (1.2%) | 12 (1.2%) |
| Non-CI TVR | 1 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 0 (0.0%) | 14 (0.8%) | 5 (1.2%) | 2 (0.7%) | 14 (0.9%) | 8 (0.8%) |
| Non-Target vessel revascularization | 8 (0.4%) | 2 (0.5%) | 0 (0.0%) | 7 (0.5%) | 4 (0.4%) | 43 (2.3%) | 11 (2.7%) | 2 (0.7%) | 40 (2.6%) | 31 (3.1%) |
| Post procedure occlusion | 1 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 0 (0.0%) | 3 (0.2%) | 3 (0.7%) | 0 (0.0%) | 3 (0.2%) | 2 (0.2%) |
| Bleeding and vascular event | 24 (1.3%) | 6 (1.5%) | 3 (1.1%) | 22 (1.4%) | 14 (1.4%) | 43 (2.3%) | 14 (3.5%) | 5 (1.8%) | 36 (2.3%) | 23 (2.3%) |
| Any stent thrombosis | 2 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 1 (0.1%) | 6 (0.3%) | 4 (1.0%) | 0 (0.0%) | 4 (0.3%) | 3 (0.3%) |
| Definite stent thrombosis | 2 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 1 (0.1%) | 6 (0.3%) | 4 (1.0%) | 0 (0.0%) | 4 (0.3%) | 3 (0.3%) |
| Acute (0–1 days) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) |
| Subacute (2–30 days) | 2 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 1 (0.1%) | 2 (0.1%) | 1 (0.2%) | 0 (0.0%) | 1 (0.1%) | 1 (0.1%) |
| Late (31–360 days) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 4 (0.2%) | 3 (0.7%) | 0 (0.0%) | 3 (0.2%) | 2 (0.2%) |
| Probable stent thrombosis | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) | 0 (0.0%) |

Data expressed as frequency (percentages).

TLF: target lesion failure; MI: myocardial infarction; TV-MI: target vessel myocardial infarction; TLR: target lesion revascularization; CI-TLR: clinically indicated target lesion revascularization; TVR: target vessel revascularization; CI-TLR: clinically indicated target vessel revascularization.