## Methods
Name of Lab | Lab9
Type of Lysate received (e.g. Trizol / RLT) | RLT
RQ-PCR machine | Light Cycler 480
RQ-PCR analysis software | Light Cycler 480 SW 1.5
Method of RNA quantification e.g.Nanodrop, Qubit (if applicable) | DeNovix Spectrophotometer
Real time PCR protocol e.g. EAC, Emig, name of commercial kit | EAC
Real time PCR reaction volume | 20µl
Volume of Plasmid standard added to RQ-PCR reactions | 5µl
Volume of cDNA added to RQ-PCR reactions | 5µl
Please provide your laboratiories standard RQ-PCR conditions (i.e. master mix, reaction componets and PCR cycling conditions), if different from those specified in the study protocol. You may provide these in a separate document if necessary.
Please provide the sequences of your laboratories ABL and BCR-ABL primers/probes, if different from the standard EAC sequences. | ABL F
ABL R
ABL Probe
BCR-ABL F
BCR-ABL R
BCR-ABL Probe

## Overview
| Lysate 1 |  | Lysate 2 |  | Lysate 3
| Run 1 | Run 2 | Run 3 | Run 4 | Run 5 | Run 6
| Primer set 1 | Primer set 2 | Primer set 3 | Primer set 2 | Lab BCR-ABL1 | Lab ABL1
| Plasmid X 1 | X 1 | X 1 | X 1 | X 1 | X 1
| Plasmid X 2 | X 2 | X 2 | X 2 | X 2 | X 2
| Plasmid X 3 | X 3 | X 3 | X 3 | X 3 | X 3
| Plasmid X 4 | X 4 | X 4 | X 4 | X 4 | X 4
| Plasmid X 5 | X 5 | X 5 | X 5 | X 5 | X 5
| Plasmid X 6 | X 6 | X 6 | X 6 | X 6 | X 6
| Plasmid Y 1 | Y 1 | Y 1 | Y 1 | Y 1 | Y 1
| Plasmid Y 2 | Y 2 | Y 2 | Y 2 | Y 2 | Y 2
| Plasmid Y 3 | Y 3 | Y 3 | Y 3 | Y 3 | Y 3
| Plasmid Y 4 | Y 4 | Y 4 | Y 4 | Y 4 | Y 4
| Plasmid Y 5 | Y 5 | Y 5 | Y 5 | Y 5 | Y 5
| Plasmid Y 6 | Y 6 | Y 6 | Y 6 | Y 6 | Y 6
| Lysate A | A | A | A | A | A
| Lysate B | B | B | B | B | B
| Lysate C | C | C | C | C | C
| Lysate D | D | D | D | D | D
| Lysate E | E | E | E | E | E
| Lysate F | F | F | F | F | F
| Lysate G | G | G | G | G | G
| Lysate H | H | H | H | H | H
| Lysate I | I | I | I | I | I
| Lysate J | J | J | J | J | J
| NTC | NTC | NTC | NTC | NTC | NTC
Standard curve to use
23 points in duplicate = 46 | 46 | 46 | 46 | 46 | 46 | 46 | 276
276 PCRs
23 points in triplicate = 69 | 69 | 69 | 69 | 69 | 69 | 69 | 414
414 PCRs

## RNA - Optional
| RNA volume (µL) | RNA concentration (ng/µL) | RNA 260/280 | RNA 260/230 | Total RNA (µg)
Lysate A 1 | 42 | 109 | 2.1 | 0.52 | 4.578
Lysate A 2 | 42 | 97 | 1.97 | 0.96 | 4.074
Lysate A 3 | 42 | 102 | 2.08 | 0.26 | 4.284
Lysate B 1 | 42 | 103 | 2.05 | 1.84 | 4.326
Lysate B 2 | 42 | 103 | 2.03 | 1.09 | 4.326
Lysate B 3 | 42 | 107 | 2.04 | 0.48 | 4.494
Lysate C 1 | 42 | 73 | 2.01 | 2.11 | 3.066
Lysate C 2 | 42 | 93 | 2.01 | 0.99 | 3.906
Lysate C 3 | 42 | 105 | 2.02 | 1.05 | 4.41
Lysate D 1 | 42 | 116 | 2.01 | 1.08 | 4.872
Lysate D 2 | 42 | 101 | 1.98 | 1.97 | 4.242
Lysate D 3 | 42 | 104 | 2.11 | 0.46 | 4.368
Lysate E 1 | 42 | 92 | 2.04 | 0.35 | 3.864
Lysate E 2 | 42 | 101 | 1.94 | 0.43 | 4.242
Lysate E 3 | 42 | 108 | 1.95 | 1.13 | 4.536
Lysate F 1 | 42 | 82 | 2.06 | 1.59 | 3.444
Lysate F 2 | 42 | 99 | 1.99 | 0.3 | 4.158
Lysate F 3 | 42 | 105 | 2.04 | 1.44 | 4.41
Lysate G 1 | 42 | 88 | 1.99 | 0.51 | 3.696
Lysate G 2 | 42 | 96 | 2.05 | 0.39 | 4.032
Lysate G 3 | 42 | 95 | 1.99 | 0.28 | 3.99
Lysate H 1 | 42 | 98 | 2.03 | 0.34 | 4.116
Lysate H 2 | 42 | 109 | 2.03 | 0.43 | 4.578
Lysate H 3 | 42 | 103 | 2.17 | 0.21 | 4.326
Lysate I 1 | 42 | 85 | 1.98 | 0.26 | 3.57
Lysate I 2 | 42 | 93 | 1.92 | 0.35 | 3.906
Lysate I 3 | 42 | 102 | 2.01 | 0.46 | 4.284
Lysate J 1 | 42 | 103 | 2 | 1.26 | 4.326
Lysate J 2 | 42 | 103 | 1.99 | 0.44 | 4.326
Lysate J 3 | 42 | 92 | 2.01 | 0.32 | 3.864

## Run 1
Run 1 - Primer Set 1
Standard Curve | Plasmid X
Gradient | -3.362
Intercept | 40.23
Ct threshold
Samples | Ct value | Copy number | µL Plasmid X added
Plasmid X 1 | 33.45 | 100 | 5
Plasmid X 1 | 33.62 | 100 | 5
Plasmid X 1
Plasmid X 2 | 29.98 | 1000 | 5
Plasmid X 2 | 30.04 | 1000 | 5
Plasmid X 2
Plasmid X 3 | 26.65 | 10000 | 5
Plasmid X 3 | 26.67 | 10000 | 5
Plasmid X 3
Plasmid X 4 | 23.88 | 100000 | 5
Plasmid X 4 | 23.27 | 100000 | 5
Plasmid X 4
Plasmid X 5 | 19.91 | 1000000 | 5
Plasmid X 5 | 19.84 | 1000000 | 5
Plasmid X 5
Plasmid X 6 | 16.58 | 10000000 | 5
Plasmid X 6 | 16.62 | 10000000 | 5
Plasmid X 6
Plasmid Y 1 | 0 | 0
Plasmid Y 1 | 0 | 0
Plasmid Y 1
Plasmid Y 2 | 0 | 0
Plasmid Y 2 | 0 | 0
Plasmid Y 2
Plasmid Y 3 | 0 | 0
Plasmid Y 3 | 0 | 0
Plasmid Y 3
Plasmid Y 4 | 0 | 0
Plasmid Y 4 | 0 | 0
Plasmid Y 4
Plasmid Y 5 | 42.03 | 0.278
Plasmid Y 5 | 41.77 | 0.332
Plasmid Y 5
Plasmid Y 6 | 38.77 | 2.58
Plasmid Y 6 | 39.11 | 2.04
Plasmid Y 6
Lysate A | 27.2 | 7120
Lysate A | 27.19 | 7170
Lysate A
Lysate B | 29.7 | 1290
Lysate B | 29.73 | 1260
Lysate B
Lysate C | 32.27 | 222
Lysate C | 32.2 | 233
Lysate C
Lysate D | 34.08 | 64.2
Lysate D | 33.88 | 73.6
Lysate D
Lysate E | 36.9 | 9.29
Lysate E | 36.63 | 11.2
Lysate E
Lysate F | 0 | 0
Lysate F | 0 | 0
Lysate F
Lysate G | 0 | 0
Lysate G | 0 | 0
Lysate G
Lysate H | 0 | 0
Lysate H | 0 | 0
Lysate H
Lysate I | 0 | 0
Lysate I | 0 | 0
Lysate I
Lysate J | 0 | 0
Lysate J | 0 | 0
Lysate J

## Run 2
Run 2 - Primer Set 2
Standard Curve | Plasmid X
Gradient | -3.389
Intercept | 40.75
Ct threshold
Samples | Ct value | Copy number | µL Plasmid X added
Plasmid X 1 | 34.12 | 100 | 5
Plasmid X 1 | 33.79 | 100 | 5
Plasmid X 1
Plasmid X 2 | 30.44 | 1000 | 5
Plasmid X 2 | 30.34 | 1000 | 5
Plasmid X 2
Plasmid X 3 | 26.87 | 10000 | 5
Plasmid X 3 | 26.83 | 10000 | 5
Plasmid X 3
Plasmid X 4 | 23.52 | 100000 | 5
Plasmid X 4 | 23.51 | 100000 | 5
Plasmid X 4
Plasmid X 5 | 20.05 | 1000000 | 5
Plasmid X 5 | 20.06 | 1000000 | 5
Plasmid X 5
Plasmid X 6 | 16.69 | 10000000 | 5
Plasmid X 6 | 16.72 | 10000000 | 5
Plasmid X 6
Plasmid Y 1 | 33.74 | 115
Plasmid Y 1 | 33.66 | 121
Plasmid Y 1
Plasmid Y 2 | 30.55 | 869
Plasmid Y 2 | 30.59 | 847
Plasmid Y 2
Plasmid Y 3 | 27.17 | 8110
Plasmid Y 3 | 27.21 | 7920
Plasmid Y 3
Plasmid Y 4 | 23.78 | 81300
Plasmid Y 4 | 23.78 | 81300
Plasmid Y 4
Plasmid Y 5 | 20.29 | 868000
Plasmid Y 5 | 20.32 | 854000
Plasmid Y 5
Plasmid Y 6 | 16.85 | 9020000
Plasmid Y 6 | 16.81 | 9230000
Plasmid Y 6
Lysate A | 22.7 | 170000
Lysate A | 22.66 | 174000
Lysate A
Lysate B | 22.7 | 169000
Lysate B | 22.57 | 185000
Lysate B
Lysate C | 22.94 | 144000
Lysate C | 22.95 | 143000
Lysate C
Lysate D | 22.66 | 174000
Lysate D | 22.6 | 181000
Lysate D
Lysate E | 22.29 | 224000
Lysate E | 22.44 | 202000
Lysate E
Lysate F | 22.8 | 158000
Lysate F | 22.89 | 149000
Lysate F
Lysate G | 22.62 | 179000
Lysate G | 22.78 | 160000
Lysate G
Lysate H | 22.76 | 163000
Lysate H | 22.78 | 161000
Lysate H
Lysate I | 22.92 | 146000
Lysate I | 22.85 | 153000
Lysate I
Lysate J | 22.82 | 156000
Lysate J | 22.71 | 168000
Lysate J

## Run 3
Run 3 - Primer Set 3
Standard Curve | Plasmid Y
Gradient | -3.363
Intercept | 40.05
Ct threshold
Samples | Ct value | Copy number
Plasmid X 1 | 0 | 0
Plasmid X 1 | 0 | 0
Plasmid X 1
Plasmid X 2 | 0 | 0
Plasmid X 2 | 0 | 0
Plasmid X 2
Plasmid X 3 | 0 | 0
Plasmid X 3 | 0 | 0
Plasmid X 3
Plasmid X 4 | 0 | 0
Plasmid X 4 | 0 | 0
Plasmid X 4
Plasmid X 5 | 0 | 0
Plasmid X 5 | 0 | 0
Plasmid X 5
Plasmid X 6 | 0 | 0
Plasmid X 6 | 0 | 0
Plasmid X 6 |  |  | µL Plasmid Y added
Plasmid Y 1 | 33.35 | 100 | 5
Plasmid Y 1 | 33.29 | 100 | 5
Plasmid Y 1
Plasmid Y 2 | 30.11 | 1000 | 5
Plasmid Y 2 | 30.15 | 1000 | 5
Plasmid Y 2
Plasmid Y 3 | 26.83 | 10000 | 5
Plasmid Y 3 | 26.84 | 10000 | 5
Plasmid Y 3
Plasmid Y 4 | 23.53 | 100000 | 5
Plasmid Y 4 | 23.47 | 100000 | 5
Plasmid Y 4
Plasmid Y 5 | 20.11 | 1000000 | 5
Plasmid Y 5 | 20.1 | 1000000 | 5
Plasmid Y 5
Plasmid Y 6 | 16.76 | 10000000 | 5
Plasmid Y 6 | 16.75 | 10000000 | 5
Plasmid Y 6
Lysate A | 0 | 0
Lysate A | 0 | 0
Lysate A
Lysate B | 0 | 0
Lysate B | 0 | 0
Lysate B
Lysate C | 0 | 0
Lysate C | 0 | 0
Lysate C
Lysate D | 0 | 0
Lysate D | 0 | 0
Lysate D
Lysate E | 0 | 0
Lysate E | 0 | 0
Lysate E
Lysate F | 26.57 | 12100
Lysate F | 26.46 | 13000
Lysate F
Lysate G | 28.7 | 2760
Lysate G | 28.71 | 2740
Lysate G
Lysate H | 30.93 | 567
Lysate H | 30.95 | 559
Lysate H
Lysate I | 33.38 | 95.5
Lysate I | 33.8 | 70
Lysate I
Lysate J | 36.04 | 13
Lysate J | 35.35 | 22
Lysate J

## Run 4
Run 4 - Primer Set 2
Standard Curve | Plasmid Y
Gradient | -3.475
Intercept | 40.61
Ct threshold
Samples | Ct value | Copy number
Plasmid X 1 | 33.45 | 120
Plasmid X 1 | 33.73 | 97.9
Plasmid X 1
Plasmid X 2 | 30.3 | 1120
Plasmid X 2 | 30.26 | 1140
Plasmid X 2
Plasmid X 3 | 26.77 | 12700
Plasmid X 3 | 26.76 | 12800
Plasmid X 3
Plasmid X 4 | 23.35 | 126000
Plasmid X 4 | 23.34 | 127000
Plasmid X 4
Plasmid X 5 | 19.92 | 1220000
Plasmid X 5 | 19.96 | 1180000
Plasmid X 5
Plasmid X 6 | 16.63 | 10800000
Plasmid X 6 | 16.62 | 10900000
Plasmid X 6 |  |  | µL Plasmid Y added
Plasmid Y 1 | 33.6 | 100 | 5
Plasmid Y 1 | 33.78 | 100 | 5
Plasmid Y 1
Plasmid Y 2 | 30.47 | 1000 | 5
Plasmid Y 2 | 30.49 | 1000 | 5
Plasmid Y 2
Plasmid Y 3 | 27.14 | 10000 | 5
Plasmid Y 3 | 27.04 | 10000 | 5
Plasmid Y 3
Plasmid Y 4 | 23.72 | 100000 | 5
Plasmid Y 4 | 23.67 | 100000 | 5
Plasmid Y 4
Plasmid Y 5 | 20.23 | 1000000 | 5
Plasmid Y 5 | 20.2 | 1000000 | 5
Plasmid Y 5
Plasmid Y 6 | 16.74 | 10000000 | 5
Plasmid Y 6 | 16.75 | 10000000 | 5
Plasmid Y 6
Lysate A | 22.57 | 211000
Lysate A | 22.52 | 218000
Lysate A
Lysate B | 22.7 | 193000
Lysate B | 22.76 | 186000
Lysate B
Lysate C | 22.65 | 200000
Lysate C | 22.63 | 203000
Lysate C
Lysate D | 22.54 | 214000
Lysate D | 22.57 | 210000
Lysate D
Lysate E | 22.48 | 223000
Lysate E | 22.56 | 213000
Lysate E
Lysate F | 22.58 | 210000
Lysate F | 22.6 | 207000
Lysate F
Lysate G | 23.02 | 157000
Lysate G | 22.78 | 183000
Lysate G
Lysate H | 22.65 | 200000
Lysate H | 22.72 | 191000
Lysate H
Lysate I | 22.74 | 189000
Lysate I | 22.75 | 187000
Lysate I
Lysate J | 22.65 | 200000
Lysate J | 22.69 | 195000
Lysate J

## Run 5
Run 5 - Lab BCR-ABL1 Primers
Standard Curve | Plasmid Y
Gradient | -3.449
Intercept | 40.39
Ct threshold
Samples | Ct value | Copy number
Plasmid X 1 | 32.71 | 164
Plasmid X 1 | 33.23 | 116
Plasmid X 1
Plasmid X 2 | 29.51 | 1390
Plasmid X 2 | 29.58 | 1330
Plasmid X 2
Plasmid X 3 | 26.23 | 12400
Plasmid X 3 | 26.27 | 12100
Plasmid X 3
Plasmid X 4 | 22.87 | 117000
Plasmid X 4 | 22.9 | 115000
Plasmid X 4
Plasmid X 5 | 19.59 | 1050000
Plasmid X 5 | 19.58 | 1060000
Plasmid X 5
Plasmid X 6 | 16.1 | 10700000
Plasmid X 6 | 16.14 | 10500000
Plasmid X 6 |  |  | µL Plasmid Y added
Plasmid Y 1 | 33.53 | 100 | 5
Plasmid Y 1 | 33.48 | 100 | 5
Plasmid Y 1
Plasmid Y 2 | 29.93 | 1000 | 5
Plasmid Y 2 | 30.04 | 1000 | 5
Plasmid Y 2
Plasmid Y 3 | 26.49 | 10000 | 5
Plasmid Y 3 | 26.6 | 10000 | 5
Plasmid Y 3
Plasmid Y 4 | 23.14 | 100000 | 5
Plasmid Y 4 | 23.16 | 100000 | 5
Plasmid Y 4
Plasmid Y 5 | 19.64 | 1000000 | 5
Plasmid Y 5 | 19.72 | 1000000 | 5
Plasmid Y 5
Plasmid Y 6 | 16.13 | 10000000 | 5
Plasmid Y 6 | 16.22 | 10000000 | 5
Plasmid Y 6
Lysate A | 26.99 | 7460
Lysate A | 26.93 | 7800
Lysate A
Lysate B | 29.3 | 1600
Lysate B | 29.32 | 1580
Lysate B
Lysate C | 31.63 | 337
Lysate C | 31.62 | 339
Lysate C
Lysate D | 33.69 | 85.2
Lysate D | 33.93 | 72.4
Lysate D
Lysate E | 36.99 | 9.4
Lysate E | 36.12 | 16.9
Lysate E
Lysate F | 26.01 | 14400
Lysate F | 26.15 | 13100
Lysate F
Lysate G | 28.18 | 3370
Lysate G | 28.25 | 3220
Lysate G
Lysate H | 30.59 | 677
Lysate H | 30.78 | 594
Lysate H
Lysate I | 33.01 | 134
Lysate I | 32.84 | 151
Lysate I
Lysate J | 35.07 | 33.9
Lysate J | 36.04 | 17.8
Lysate J

## Run 6
Run 6 - Lab ABL1 Primers
Standard Curve | Plasmid Y
Gradient | -3.46
Intercept | 40.65
Ct threshold
Samples | Ct value | Copy number
Plasmid X 1 | 33.57 | 117
Plasmid X 1 | 33.65 | 110
Plasmid X 1
Plasmid X 2 | 30.47 | 1070
Plasmid X 2 | 30.56 | 1010
Plasmid X 2
Plasmid X 3 | 26.93 | 12600
Plasmid X 3 | 26.98 | 12100
Plasmid X 3
Plasmid X 4 | 23.53 | 124000
Plasmid X 4 | 23.56 | 122000
Plasmid X 4
Plasmid X 5 | 20.15 | 1180000
Plasmid X 5 | 20.1 | 1220000
Plasmid X 5
Plasmid X 6 | 16.85 | 10600000
Plasmid X 6 | 16.86 | 10500000
Plasmid X 6 |  |  | µL Plasmid Y added
Plasmid Y 1 | 33.98 | 100 | 5
Plasmid Y 1 | 33.57 | 100 | 5
Plasmid Y 1
Plasmid Y 2 | 30.51 | 1000 | 5
Plasmid Y 2 | 30.65 | 1000 | 5
Plasmid Y 2
Plasmid Y 3 | 27.25 | 10000 | 5
Plasmid Y 3 | 27.24 | 10000 | 5
Plasmid Y 3
Plasmid Y 4 | 23.79 | 100000 | 5
Plasmid Y 4 | 23.84 | 100000 | 5
Plasmid Y 4
Plasmid Y 5 | 20.48 | 1000000 | 5
Plasmid Y 5 | 20.48 | 1000000 | 5
Plasmid Y 5
Plasmid Y 6 | 16.87 | 10000000 | 5
Plasmid Y 6 | 16.92 | 10000000 | 5
Plasmid Y 6
Lysate A | 22.52 | 243000
Lysate A | 22.58 | 234000
Lysate A
Lysate B | 22.52 | 243000
Lysate B | 22.57 | 235000
Lysate B
Lysate C | 22.57 | 235000
Lysate C | 22.56 | 236000
Lysate C
Lysate D | 22.52 | 242000
Lysate D | 22.53 | 241000
Lysate D
Lysate E | 22.66 | 222000
Lysate E | 22.64 | 225000
Lysate E
Lysate F | 22.83 | 198000
Lysate F | 22.86 | 194000
Lysate F
Lysate G | 22.75 | 209000
Lysate G | 22.68 | 219000
Lysate G
Lysate H | 22.69 | 217000
Lysate H | 22.71 | 215000
Lysate H
Lysate I | 22.71 | 214000
Lysate I | 22.66 | 221000
Lysate I
Lysate J | 22.75 | 209000
Lysate J | 22.72 | 214000
Lysate J